## QC3_Summary
Instructions
1. Fill out the table below (columns D through M) based on the most up-to-date test for each detector
> The leak rate should be determined as the slope of the pressure versus time plot.
> For Y = Ax + B this is A; for Y = exp(A x + B) this is A.
2. If a test would supercede the below information, update the table to reflect only the most up-to-date information
3. Follow the instructions on "QC3_Curves" tab that correspond with the test you are entering in this tab
4. Save the excel spreadsheet
[8 empty] | Setup |  |  |  |  | Setup & Chamber |  |  |  |  | Chamber
| GE1/1 | Note | Test Date | Avg. Amb. Temp. | Std. Dev. Temp | Avg. Amb. Pressure | Std. Dev Amb. Pressure | Initial Pressure | Final Pressure | deltaP | Duration | Leak Rate | Initial Pressure | Final Pressure | Leak Rate | Duration | Leak Rate | Leak Rate | E-Log
Short Link URL
| S/N |  | (YYYY/MM/DD) | (Deg C) | (Deg C) | (mbar) | (mbar) | (mbar) | (mbar) | (mbar) | (hr) | (mbar/hr) | (mbar) | (mbar) | (mbar/hr) | (hr) | (mbar/hr) | (mbar/hr)
Short Detector | GE1/1-X-S-CERN-0001 | Tha chamber is leaking fast, the main problem was found in the corners of the small base of the trapezoid | 2017-05-03 | 24.02 | 0.08 | 972.7 | 0.24 | Not done |  |  |  |  | 23.03 | 0 | 23.03 | 0.75028 | 30.6953 | 30.6953 | http://cmsonline.cern.ch/cms-elog/979820
| GE1/1-X-S-CERN-0002 |  |  | 23.85 | 0.08 | 978.8 | 0.11 | [5 empty] | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[15 empty] | 0 |  | #DIV/0!
| GE1/1-X-S-CERN-0003 | With tape on leaking screw, the chamber was gas tight | 2017-06-30 | [9 empty] | 24.69 | 24.93 | -0.24 | 1 | -0.24
[15 empty] | 0 |  | #DIV/0!
| GE1/1-X-S-CERN-0004 | Chamber is leaking @ 1mbar/hr. | 2017-07-06
[15 empty] | 0 |  | #DIV/0!
[15 empty] | 0 |  | #DIV/0!
[15 empty] | 0 |  | #DIV/0!
Long Detector | GE1/1-X-L-CERN-0001 | After bending the RO and applying 1.3Nm torque on RO and DR screws, the chamber was leaking at 1.3mbar/hr | 2017-07-14 | [9 empty] | 26.01 | 24.71 | 1.3 | 1.0003 | 1.2996
| GE1/1-X-L-CERN-0002 | Chamber is gas tight | 2017-06-01 |  |  |  |  | 29 | 29.29 | -0.29 | 0.95028 |  | 20.38 | 19.17 | 1.21 | 1.0503 | 1.1521
| GE1/1-X-L-CERN-0003 | After applying 1.3Nm torque on the RO and DR screws, the chamber was leaking at 2.33 mbar/hr | 2017-07-12 | [9 empty] | 24.36 | 21.96 | 2.4 | 1.0172 | 2.3594
| GE1/1-X-L-CERN-0004 | Chamber is gas tight | 2017-07-27 | [9 empty] | 26.68 | 25.66 | 1.02 | 1.0003 | 1.0197
| GE1/1-X-L-CERN-0003 | After fixing the readout board in the clean room | 2017-08-01 | [9 empty] | 25.18 | 23.51 | 1.67 | 1.0169 | 1.6422

## QC3_Long_4
Test Date: | 2017-07-27
File: | QC3_GE11-X-L-CERN-0004_20170727.xls
Elog: | http://cmsonline.cern.ch/cms-elog/988203
Description: | Calibration (Stand2)
Max P (mbar) | 26.68
Min P (mbar) | 25.66
Delta P (mbar) | 1.02
Min Time | 10:42
Max Time | 11:42
Delta T | 1:00
Time (hr) | 1.0003
Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar)
10:42 | 1 | 26.68 | 22.3 | 965.77
10:43 | 62 | 26.62 | 22.28 | 966.16
10:44 | 123 | 26.57 | 22.29 | 966.14
10:45 | 184 | 26.54 | 22.28 | 966.07
10:46 | 245 | 26.47 | 22.28 | 966.09
10:47 | 306 | 26.46 | 22.28 | 966.11
10:48 | 367 | 26.35 | 22.28 | 966.07
10:49 | 428 | 26.41 | 22.28 | 966.05
10:50 | 489 | 26.34 | 22.28 | 966.04
10:51 | 550 | 26.33 | 22.3 | 966.01
10:52 | 611 | 26.3 | 22.3 | 965.97
10:53 | 672 | 26.25 | 22.28 | 965.98
10:54 | 733 | 26.28 | 22.3 | 965.95
10:55 | 794 | 26.28 | 22.3 | 965.91
10:56 | 855 | 26.22 | 22.3 | 965.94
10:57 | 916 | 26.14 | 22.29 | 965.97
10:58 | 977 | 26.17 | 22.28 | 966
10:59 | 1038 | 26.13 | 22.28 | 966
11:00 | 1099 | 26.12 | 22.29 | 966.02
11:01 | 1160 | 26.08 | 22.29 | 966.02
11:02 | 1221 | 26.01 | 22.29 | 966
11:03 | 1282 | 26.04 | 22.29 | 965.96
11:04 | 1343 | 26.03 | 22.29 | 965.98
11:05 | 1404 | 26.05 | 22.29 | 965.92
11:06 | 1465 | 26.07 | 22.28 | 965.94
11:07 | 1526 | 26.06 | 22.29 | 965.93
11:08 | 1587 | 26.02 | 22.29 | 965.9
11:09 | 1648 | 25.99 | 22.29 | 965.9
11:10 | 1709 | 26.04 | 22.31 | 965.89
11:11 | 1770 | 25.98 | 22.29 | 965.88
11:12 | 1831 | 25.99 | 22.29 | 965.89
11:13 | 1870 | 25.97 | 22.29 | 965.89
11:14 | 1953 | 25.95 | 22.3 | 965.87
11:15 | 2014 | 25.97 | 22.28 | 965.85
11:16 | 2075 | 25.94 | 22.29 | 965.83
11:17 | 2136 | 25.86 | 22.3 | 965.85
11:18 | 2197 | 25.83 | 22.29 | 965.85
11:19 | 2258 | 25.85 | 22.28 | 965.86
11:20 | 2319 | 25.81 | 22.29 | 965.88
11:21 | 2380 | 25.87 | 22.29 | 965.87
11:22 | 2441 | 25.8 | 22.29 | 965.88
11:23 | 2502 | 25.83 | 22.28 | 965.89
11:24 | 2563 | 25.84 | 22.28 | 965.84
11:25 | 2624 | 25.84 | 22.28 | 965.87
11:26 | 2685 | 25.82 | 22.28 | 965.88
11:27 | 2746 | 25.76 | 22.29 | 965.88
11:28 | 2807 | 25.78 | 22.28 | 965.85
11:29 | 2868 | 25.78 | 22.29 | 965.84
11:30 | 2929 | 25.77 | 22.29 | 965.86
11:31 | 2990 | 25.73 | 22.3 | 965.82
11:32 | 3051 | 25.72 | 22.29 | 965.86
11:33 | 3112 | 25.74 | 22.29 | 965.83
11:34 | 3173 | 25.81 | 22.28 | 965.81
11:35 | 3234 | 25.76 | 22.29 | 965.84
11:36 | 3295 | 25.72 | 22.3 | 965.81
11:37 | 3356 | 25.72 | 22.29 | 965.82
11:38 | 3417 | 25.69 | 22.29 | 965.81
11:39 | 3478 | 25.71 | 22.29 | 965.82
11:40 | 3539 | 25.7 | 22.3 | 965.82
11:41 | 3600 | 25.72 | 22.29 | 965.78
11:42 | 3661 | 25.66 | 22.29 | 965.8

## QC3_Summary_of_Setups
Most Recent Measurement is used

## QC3_Long_Detectors_Summary
Most Recent Measurement is used

## QC3_Short_Detectors_Summary
Most Recent Measurement is used

## QC3_Short_1
Test Date: | 2017-05-02 |  |  |  | Test Date: | 2017-05-02 |  |  |  | Test Date: | 2017-05-02 |  |  |  | Test Date: | 2017-03-05 |  |  |  | Test Date: | 2017-03-05 |  |  |  | Test Date: | 2017-03-05
File: | QC3_GE11-X-S-CERN-0001_20170502_test1.xlsm |  |  |  | File: | QC3_GE11-X-S-CERN-0001_20170502_test2.xlsm |  |  |  | File: | QC3_GE11-X-S-CERN-0001_20170502_test3.xlsm |  |  |  | File: | QC3_GE11-X-S-CERN-0001_20170503_test4.xlsm |  |  |  | File: | QC3_GE11-X-S-CERN-0001_20170503_test5.xlsm |  |  |  | File: | QC3_GE11-X-S-CERN-0001_20170503_test6.xlsm
Elog: | http://cmsonline.cern.ch/cms-elog/979820 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/979820 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/979820 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/979820 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/979820 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/979820
Description: | test 1: 1.3 Nm torque both on RO and Drift screws |  |  |  | Description: | test 2: green tape on one corner of the small base of the trapezoid |  |  |  | Description: | test 3: tape both on the readout and drift screws+ readout pins |  |  |  | Description: | test 4: tape on drift and RO screws + placing clamps in the four corners |  |  |  | Description: | test 5: tape on drift and RO screws + placing clamps every 10 cm all around the chamber |  |  |  | Description: | test 6: kapton tape between the external frame the readout and drift on the big and small base of the trapezoid
Max P (mbar) | 25.35 |  |  |  | Max P (mbar) | 20.38 |  |  |  | Max P (mbar) | 26.01 |  |  |  | Max P (mbar) | 25.68 |  |  |  | Max P (mbar) | 24.36 |  |  |  | Max P (mbar) | 23.03
Min P (mbar) | 0.3 |  |  |  | Min P (mbar) | 0 |  |  |  | Min P (mbar) | 0.15 |  |  |  | Min P (mbar) | 0.65 |  |  |  | Min P (mbar) | 1.27 |  |  |  | Min P (mbar) | 1.27
Delta P (mbar) | 25.05 |  |  |  | Delta P (mbar) | 20.38 |  |  |  | Delta P (mbar) | 25.86 |  |  |  | Delta P (mbar) | 25.03 |  |  |  | Delta P (mbar) | 23.09 |  |  |  | Delta P (mbar) | 21.76
Min Time | 11:15 |  |  |  | Min Time | 12:11 |  |  |  | Min Time | 15:20 |  |  |  | Min Time | 18:19 |  |  |  | Min Time | 10:27 |  |  |  | Min Time | 11:50
Max Time | 12:00 |  |  |  | Max Time | 12:56 |  |  |  | Max Time | 16:05 |  |  |  | Max Time | 19:04 |  |  |  | Max Time | 11:12 |  |  |  | Max Time | 12:35
Delta T | 0:45 |  |  |  | Delta T | 0:45 |  |  |  | Delta T | 0:45 |  |  |  | Delta T | 0:45 |  |  |  | Delta T | 0:45 |  |  |  | Delta T | 0:45
Time (hr) | 0.75028 |  |  |  | Time (hr) | 0.75028 |  |  |  | Time (hr) | 0.75028 |  |  |  | Time (hr) | 0.75 |  |  |  | Time (hr) | 0.75028 |  |  |  | Time (hr) | 0.75028
Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar)
11:15 | 1 | 25.35 | 22.3 | 967.94 | 12:11 | 1 | 20.38 | 22.3 | 966.86 | 15:20 | 1 | 26.01 | 23.27 | 967.94 | 18:19 | 1 | 25.68 | 21.81 | 969.03 | 10:27 | 1 | 24.36 | 22.3 | 970.12 | 11:50 | 1 | 23.03 | 22.3 | 970.12
11:16 | 62 | 23.77 | 22.24 | 967.42 | 12:12 | 62 | 18.83 | 22.25 | 967.43 | 15:21 | 62 | 24.68 | 22.28 | 967.11 | 18:20 | 62 | 23.98 | 22.24 | 968.06 | 10:28 | 62 | 23.08 | 22.24 | 969.9 | 11:51 | 62 | 21.79 | 22.25 | 969.78
11:17 | 123 | 22.3 | 22.22 | 967.42 | 12:13 | 123 | 17.69 | 22.25 | 967.46 | 15:22 | 123 | 23.18 | 22.28 | 967.11 | 18:21 | 123 | 22.61 | 22.24 | 968.1 | 10:29 | 123 | 21.84 | 22.25 | 969.92 | 11:52 | 123 | 20.73 | 22.25 | 969.77
11:18 | 184 | 20.96 | 22.22 | 967.46 | 12:14 | 184 | 16.51 | 22.25 | 967.39 | 15:23 | 184 | 21.83 | 22.28 | 967.11 | 18:22 | 184 | 21.41 | 22.25 | 968.09 | 10:30 | 184 | 20.71 | 22.25 | 969.91 | 11:53 | 184 | 19.73 | 22.24 | 969.72
11:19 | 245 | 19.72 | 22.2 | 967.49 | 12:15 | 245 | 15.4 | 22.25 | 967.38 | 15:24 | 245 | 20.52 | 22.29 | 967.11 | 18:23 | 245 | 20.26 | 22.24 | 968.03 | 10:31 | 245 | 19.72 | 22.25 | 969.93 | 11:54 | 245 | 18.69 | 22.23 | 969.74
11:20 | 306 | 18.37 | 22.2 | 967.46 | 12:16 | 306 | 14.45 | 22.25 | 967.39 | 15:25 | 306 | 19.22 | 22.28 | 967.11 | 18:24 | 306 | 19.06 | 22.24 | 968.07 | 10:32 | 306 | 18.65 | 22.25 | 969.94 | 11:55 | 306 | 17.79 | 22.23 | 969.72
11:21 | 367 | 17.21 | 22.21 | 967.47 | 12:17 | 367 | 13.39 | 22.26 | 967.37 | 15:26 | 367 | 18.04 | 22.29 | 967.15 | 18:25 | 367 | 17.96 | 22.24 | 968.08 | 10:33 | 367 | 17.69 | 22.25 | 969.94 | 11:56 | 367 | 16.82 | 22.23 | 969.74
11:22 | 428 | 16.11 | 22.21 | 967.5 | 12:18 | 428 | 12.43 | 22.25 | 967.44 | 15:27 | 428 | 16.89 | 22.28 | 967.14 | 18:26 | 428 | 17.01 | 22.25 | 968.08 | 10:34 | 428 | 16.71 | 22.25 | 969.93 | 11:57 | 428 | 16.04 | 22.22 | 969.73
11:23 | 489 | 15.1 | 22.22 | 967.55 | 12:19 | 489 | 11.6 | 22.26 | 967.41 | 15:28 | 489 | 15.87 | 22.29 | 967.12 | 18:27 | 489 | 15.99 | 22.27 | 968.08 | 10:35 | 489 | 15.81 | 22.25 | 969.93 | 11:58 | 489 | 15.28 | 22.22 | 969.69
11:24 | 550 | 14.06 | 22.23 | 967.53 | 12:20 | 550 | 10.77 | 22.24 | 967.44 | 15:29 | 550 | 14.76 | 22.28 | 967.13 | 18:28 | 550 | 15.06 | 22.25 | 968.04 | 10:36 | 550 | 15.01 | 22.25 | 969.9 | 11:59 | 550 | 14.43 | 22.24 | 969.7
11:25 | 611 | 13.07 | 22.23 | 967.5 | 12:21 | 611 | 9.98 | 22.24 | 967.49 | 15:30 | 611 | 13.78 | 22.28 | 967.1 | 18:29 | 611 | 14.1 | 22.26 | 968.06 | 10:37 | 611 | 14.17 | 22.26 | 969.95 | 12:00 | 611 | 13.67 | 22.23 | 969.66
11:26 | 672 | 12.2 | 22.23 | 967.47 | 12:22 | 672 | 9.28 | 22.24 | 967.42 | 15:31 | 672 | 12.79 | 22.27 | 967.15 | 18:30 | 672 | 13.28 | 22.25 | 968.09 | 10:38 | 672 | 13.33 | 22.26 | 969.93 | 12:01 | 672 | 12.95 | 22.23 | 969.71
11:27 | 733 | 11.39 | 22.22 | 967.42 | 12:23 | 733 | 8.52 | 22.25 | 967.45 | 15:32 | 733 | 11.92 | 22.25 | 967.14 | 18:31 | 733 | 12.42 | 22.25 | 968.07 | 10:39 | 733 | 12.61 | 22.26 | 969.95 | 12:02 | 733 | 12.16 | 22.23 | 969.65
11:28 | 794 | 10.55 | 22.23 | 967.4 | 12:24 | 794 | 7.92 | 22.25 | 967.39 | 15:33 | 794 | 11.09 | 22.27 | 967.11 | 18:32 | 794 | 11.67 | 22.25 | 968.07 | 10:40 | 794 | 11.91 | 22.25 | 969.95 | 12:03 | 794 | 11.54 | 22.23 | 969.6
11:29 | 855 | 9.8 | 22.24 | 967.38 | 12:25 | 855 | 7.25 | 22.24 | 967.39 | 15:34 | 855 | 10.25 | 22.27 | 967.1 | 18:33 | 855 | 10.86 | 22.26 | 968.09 | 10:41 | 855 | 11.2 | 22.25 | 969.97 | 12:04 | 855 | 10.91 | 22.23 | 969.57
11:30 | 916 | 9.06 | 22.24 | 967.42 | 12:26 | 916 | 6.66 | 22.25 | 967.38 | 15:35 | 916 | 9.44 | 22.26 | 967.1 | 18:34 | 916 | 10.18 | 22.24 | 968.06 | 10:42 | 916 | 10.57 | 22.24 | 969.96 | 12:05 | 916 | 10.19 | 22.22 | 969.63
11:31 | 977 | 8.39 | 22.24 | 967.46 | 12:27 | 977 | 6.17 | 22.26 | 967.32 | 15:36 | 977 | 8.75 | 22.27 | 967.08 | 18:35 | 977 | 9.53 | 22.25 | 968.04 | 10:43 | 977 | 10 | 22.25 | 969.96 | 12:06 | 977 | 9.7 | 22.23 | 969.58
11:32 | 1038 | 7.69 | 22.23 | 967.42 | 12:28 | 1038 | 5.66 | 22.26 | 967.3 | 15:37 | 1038 | 8.08 | 22.27 | 967.1 | 18:36 | 1038 | 8.88 | 22.24 | 968.06 | 10:44 | 1038 | 9.4 | 22.23 | 969.95 | 12:07 | 1038 | 9.14 | 22.21 | 969.61
11:33 | 1099 | 7.14 | 22.22 | 967.47 | 12:29 | 1099 | 5.15 | 22.25 | 967.33 | 15:38 | 1099 | 7.39 | 22.26 | 967.11 | 18:37 | 1099 | 8.24 | 22.25 | 968.05 | 10:45 | 1099 | 8.8 | 22.24 | 969.94 | 12:08 | 1099 | 8.54 | 22.22 | 969.63
11:34 | 1160 | 6.56 | 22.22 | 967.52 | 12:30 | 1160 | 4.81 | 22.23 | 967.29 | 15:39 | 1160 | 6.82 | 22.25 | 967.11 | 18:38 | 1160 | 7.69 | 22.26 | 968.06 | 10:46 | 1160 | 8.29 | 22.23 | 969.97 | 12:09 | 1160 | 8.1 | 22.21 | 969.62
11:35 | 1221 | 5.98 | 22.22 | 967.52 | 12:31 | 1221 | 4.27 | 22.24 | 967.3 | 15:40 | 1221 | 6.26 | 22.28 | 967.1 | 18:39 | 1221 | 7.14 | 22.26 | 968.04 | 10:47 | 1221 | 7.85 | 22.25 | 969.94 | 12:10 | 1221 | 7.58 | 22.22 | 969.61
11:36 | 1282 | 5.46 | 22.24 | 967.55 | 12:32 | 1282 | 4.03 | 22.24 | 967.35 | 15:41 | 1282 | 5.76 | 22.27 | 967.1 | 18:40 | 1282 | 6.64 | 22.25 | 968.05 | 10:48 | 1282 | 7.28 | 22.26 | 969.95 | 12:11 | 1282 | 7.17 | 22.2 | 969.65
11:37 | 1343 | 5.01 | 22.23 | 967.61 | 12:33 | 1343 | 3.65 | 22.25 | 967.36 | 15:42 | 1343 | 5.25 | 22.27 | 967.06 | 18:41 | 1343 | 6.09 | 22.24 | 968.04 | 10:49 | 1343 | 6.81 | 22.25 | 969.94 | 12:12 | 1343 | 6.71 | 22.19 | 969.64
11:38 | 1404 | 4.63 | 22.22 | 967.62 | 12:34 | 1404 | 3.34 | 22.25 | 967.36 | 15:43 | 1404 | 4.75 | 22.27 | 967.09 | 18:42 | 1404 | 5.75 | 22.26 | 968.03 | 10:50 | 1404 | 6.44 | 22.25 | 969.92 | 12:13 | 1404 | 6.33 | 22.23 | 969.62
11:39 | 1465 | 4.2 | 22.22 | 967.63 | 12:35 | 1465 | 3.06 | 22.23 | 967.34 | 15:44 | 1465 | 4.39 | 22.26 | 967.11 | 18:43 | 1465 | 5.31 | 22.26 | 968.04 | 10:51 | 1465 | 6.07 | 22.25 | 969.97 | 12:14 | 1465 | 5.87 | 22.2 | 969.67
11:40 | 1526 | 3.87 | 22.22 | 967.6 | 12:36 | 1526 | 2.79 | 22.23 | 967.31 | 15:45 | 1526 | 4.02 | 22.26 | 967.11 | 18:44 | 1526 | 4.91 | 22.25 | 968.06 | 10:52 | 1526 | 5.63 | 22.25 | 969.94 | 12:15 | 1526 | 5.5 | 22.19 | 969.63
11:41 | 1587 | 3.56 | 22.22 | 967.6 | 12:37 | 1587 | 2.51 | 22.21 | 967.32 | 15:46 | 1587 | 3.62 | 22.26 | 967.11 | 18:45 | 1587 | 4.56 | 22.23 | 968.06 | 10:53 | 1587 | 5.29 | 22.26 | 969.94 | 12:16 | 1587 | 5.19 | 22.19 | 969.6
11:42 | 1648 | 3.23 | 22.23 | 967.54 | 12:38 | 1648 | 2.23 | 22.21 | 967.32 | 15:47 | 1648 | 3.34 | 22.27 | 967.08 | 18:46 | 1648 | 4.2 | 22.24 | 968.05 | 10:54 | 1648 | 4.92 | 22.25 | 969.94 | 12:17 | 1648 | 4.87 | 22.18 | 969.61
11:43 | 1709 | 2.97 | 22.22 | 967.59 | 12:39 | 1709 | 2.05 | 22.2 | 967.37 | 15:48 | 1709 | 2.99 | 22.27 | 967.08 | 18:47 | 1709 | 3.93 | 22.23 | 968.04 | 10:55 | 1709 | 4.6 | 22.25 | 969.96 | 12:18 | 1709 | 4.52 | 22.18 | 969.58
11:44 | 1770 | 2.71 | 22.23 | 967.53 | 12:40 | 1770 | 1.84 | 22.21 | 967.41 | 15:49 | 1770 | 2.7 | 22.26 | 967.08 | 18:48 | 1770 | 3.57 | 22.24 | 968.04 | 10:56 | 1770 | 4.36 | 22.24 | 969.94 | 12:19 | 1770 | 4.27 | 22.17 | 969.58
11:45 | 1831 | 2.48 | 22.23 | 967.56 | 12:41 | 1831 | 1.61 | 22.21 | 967.44 | 15:50 | 1831 | 2.44 | 22.26 | 967.08 | 18:49 | 1831 | 3.27 | 22.25 | 968.03 | 10:57 | 1831 | 4.06 | 22.24 | 969.98 | 12:20 | 1831 | 4.02 | 22.17 | 969.54
11:46 | 1892 | 2.21 | 22.24 | 967.51 | 12:42 | 1892 | 1.44 | 22.2 | 967.37 | 15:51 | 1892 | 2.21 | 22.25 | 967.11 | 18:50 | 1892 | 3.1 | 22.25 | 968.02 | 10:58 | 1892 | 3.76 | 22.24 | 969.98 | 12:21 | 1892 | 3.72 | 22.17 | 969.52
11:47 | 1953 | 1.97 | 22.22 | 967.52 | 12:43 | 1953 | 1.35 | 22.17 | 967.35 | 15:52 | 1953 | 1.98 | 22.26 | 967.11 | 18:51 | 1953 | 2.76 | 22.25 | 968.01 | 10:59 | 1953 | 3.52 | 22.24 | 969.99 | 12:22 | 1953 | 3.47 | 22.17 | 969.5
11:48 | 2014 | 1.83 | 22.23 | 967.52 | 12:44 | 2014 | 1.19 | 22.16 | 967.32 | 15:53 | 2014 | 1.73 | 22.28 | 967.07 | 18:52 | 2014 | 2.57 | 22.24 | 968.03 | 11:00 | 2014 | 3.28 | 22.25 | 969.96 | 12:23 | 2014 | 3.28 | 22.16 | 969.46
11:49 | 2075 | 1.59 | 22.24 | 967.5 | 12:45 | 2075 | 1 | 22.13 | 967.34 | 15:54 | 2075 | 1.52 | 22.27 | 967.09 | 18:53 | 2075 | 2.3 | 22.24 | 968.02 | 11:01 | 2075 | 3.08 | 22.24 | 969.98 | 12:24 | 2075 | 3.06 | 22.15 | 969.44
11:50 | 2136 | 1.4 | 22.24 | 967.53 | 12:46 | 2136 | 0.85 | 22.1 | 967.34 | 15:55 | 2136 | 1.34 | 22.28 | 967.06 | 18:54 | 2136 | 2.08 | 22.23 | 968.03 | 11:02 | 2136 | 2.88 | 22.24 | 969.95 | 12:25 | 2136 | 2.8 | 22.18 | 969.4
11:51 | 2197 | 1.24 | 22.24 | 967.5 | 12:47 | 2197 | 0.74 | 22.08 | 967.37 | 15:56 | 2197 | 1.18 | 22.27 | 967.08 | 18:55 | 2197 | 1.88 | 22.21 | 968.05 | 11:03 | 2197 | 2.66 | 22.24 | 969.97 | 12:26 | 2197 | 2.66 | 22.19 | 969.4
11:52 | 2258 | 1.14 | 22.22 | 967.53 | 12:48 | 2258 | 0.58 | 22.06 | 967.32 | 15:57 | 2258 | 0.97 | 22.27 | 967.09 | 18:56 | 2258 | 1.67 | 22.22 | 968.02 | 11:04 | 2258 | 2.54 | 22.25 | 969.99 | 12:27 | 2258 | 2.46 | 22.19 | 969.39
11:53 | 2319 | 1.02 | 22.23 | 967.52 | 12:49 | 2319 | 0.57 | 22.06 | 967.3 | 15:58 | 2319 | 0.8 | 22.27 | 967.1 | 18:57 | 2319 | 1.58 | 22.23 | 968.03 | 11:05 | 2319 | 2.23 | 22.26 | 969.97 | 12:28 | 2319 | 2.29 | 22.16 | 969.38
11:54 | 2380 | 0.85 | 22.25 | 967.48 | 12:50 | 2380 | 0.43 | 22.06 | 967.29 | 15:59 | 2380 | 0.73 | 22.26 | 967.07 | 18:58 | 2380 | 1.41 | 22.23 | 968.06 | 11:06 | 2380 | 2.13 | 22.25 | 969.96 | 12:29 | 2380 | 2.12 | 22.13 | 969.37
11:55 | 2441 | 0.76 | 22.24 | 967.52 | 12:51 | 2441 | 0.36 | 22.08 | 967.23 | 16:00 | 2441 | 0.55 | 22.27 | 967.07 | 18:59 | 2441 | 1.24 | 22.23 | 968.06 | 11:07 | 2441 | 1.98 | 22.26 | 969.97 | 12:30 | 2441 | 1.98 | 22.11 | 969.34
11:56 | 2502 | 0.7 | 22.24 | 967.52 | 12:52 | 2502 | 0.31 | 22.06 | 967.24 | 16:01 | 2502 | 0.45 | 22.27 | 967.09 | 19:00 | 2502 | 1.12 | 22.23 | 968.07 | 11:08 | 2502 | 1.82 | 22.26 | 969.99 | 12:31 | 2502 | 1.81 | 22.11 | 969.34
11:57 | 2563 | 0.51 | 22.24 | 967.52 | 12:53 | 2563 | 0.22 | 22.07 | 967.21 | 16:02 | 2563 | 0.41 | 22.26 | 967.05 | 19:01 | 2563 | 0.95 | 22.21 | 968.08 | 11:09 | 2563 | 1.64 | 22.26 | 969.96 | 12:32 | 2563 | 1.7 | 22.12 | 969.32
11:58 | 2624 | 0.43 | 22.23 | 967.59 | 12:54 | 2624 | 0.11 | 22.05 | 967.23 | 16:03 | 2624 | 0.28 | 22.27 | 967.08 | 19:02 | 2624 | 0.8 | 22.21 | 968.1 | 11:10 | 2624 | 1.58 | 22.26 | 969.98 | 12:33 | 2624 | 1.52 | 22.12 | 969.3
11:59 | 2685 | 0.38 | 22.25 | 967.54 | 12:55 | 2685 | 0.13 | 22.06 | 967.19 | 16:04 | 2685 | 0.19 | 22.27 | 967.08 | 19:03 | 2685 | 0.77 | 22.19 | 968.12 | 11:11 | 2685 | 1.4 | 22.25 | 970 | 12:34 | 2685 | 1.41 | 22.12 | 969.32
12:00 | 2746 | 0.3 | 22.26 | 967.46 | 12:56 | 2746 | 0 | 22.06 | 967.2 | 16:05 | 2746 | 0.15 | 22.27 | 967.08 | 19:04 | 2746 | 0.65 | 22.19 | 968.1 | 11:12 | 2746 | 1.27 | 22.26 | 969.96 | 12:35 | 2746 | 1.27 | 22.1 | 969.32
[10 empty] | 16:06 | 2807 | 0.04 | 22.27 | 967.09 | 19:05 | 2807 | 0.51 | 22.22 | 968.1 | 11:13 | 2807 | 1.18 | 22.25 | 969.98 | 12:36 | 2807 | 1.19 | 22.13 | 969.3
[15 empty] | 19:06 | 2868 | 0.49 | 22.22 | 968.12 | 11:14 | 2868 | 1.08 | 22.25 | 969.97 | 12:37 | 2868 | 1.08 | 22.13 | 969.28
[15 empty] | 19:07 | 2929 | 0.39 | 22.23 | 968.08 | 11:15 | 2929 | 1.03 | 22.26 | 969.96 | 12:38 | 2929 | 0.96 | 22.1 | 969.28
[15 empty] | 19:08 | 2990 | 0.25 | 22.22 | 968.13 | 11:16 | 2990 | 0.86 | 22.27 | 969.98 | 12:39 | 2990 | 0.9 | 22.1 | 969.27
[15 empty] | 19:09 | 3051 | 0.21 | 22.25 | 968.11 | 11:17 | 3051 | 0.79 | 22.26 | 969.98 | 12:40 | 3051 | 0.78 | 22.1 | 969.27
[15 empty] | 19:10 | 3112 | 0.15 | 22.23 | 968.16 | 11:18 | 3112 | 0.72 | 22.26 | 969.96 | 12:41 | 3112 | 0.7 | 22.1 | 969.25
[15 empty] | 19:11 | 3173 | 0.13 | 22.25 | 968.14 | 11:19 | 3173 | 0.76 | 22.25 | 969.99 | 12:42 | 3173 | 0.58 | 22.11 | 969.25
[15 empty] | 19:12 | 3234 | 0.03 | 22.25 | 968.14 | 11:20 | 3234 | 0.6 | 22.27 | 969.91 | 12:43 | 3234 | 0.55 | 22.12 | 969.25
[20 empty] | 11:21 | 3295 | 0.57 | 22.26 | 969.92 | 12:44 | 3295 | 0.52 | 22.14 | 969.25
[20 empty] | 11:22 | 3356 | 0.49 | 22.26 | 969.93 | 12:45 | 3356 | 0.46 | 22.12 | 969.26
[20 empty] | 11:23 | 3417 | 0.4 | 22.26 | 969.95 | 12:46 | 3417 | 0.42 | 22.1 | 969.25
[20 empty] | 11:24 | 3478 | 0.4 | 22.25 | 969.94 | 12:47 | 3478 | 0.38 | 22.08 | 969.26
[20 empty] | 11:25 | 3539 | 0.32 | 22.24 | 969.96 | 12:48 | 3539 | 0.28 | 22.07 | 969.26
[25 empty] | 12:49 | 3600 | 0.23 | 22.09 | 969.25
[25 empty] | 12:50 | 3661 | 0.16 | 22.1 | 969.25
[25 empty] | 12:51 | 3722 | 0.17 | 22.1 | 969.27
[25 empty] | 12:52 | 3783 | 0.16 | 22.07 | 969.26
[25 empty] | 12:53 | 3844 | 0.07 | 22.07 | 969.24
[25 empty] | 12:54 | 3905 | 0.03 | 22.06 | 969.26
[25 empty] | 12:55 | 3966 | 0 | 22.09 | 969.25

## QC3_Short_2
Test Date: | 06/20/2017
File: | QC3_GE11-X-S-CERN-0002_2017_06_20.xlsm
Elog:
Description:
Max P (mbar) | 25.35
Min P (mbar) | -0.43
Delta P (mbar) | 25.78
Min Time | 9:43
Max Time | 10:28
Delta T | 0:45
Time (hr) | 0.75
Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar)
9:43 | 1 | 25.35 | 22.3 | 971.2
9:44 | 62 | 23.04 | 22.2 | 971.39
9:45 | 123 | 20.79 | 22.2 | 971.46
9:46 | 184 | 18.85 | 22.25 | 971.45
9:47 | 245 | 17.11 | 22.23 | 971.44
9:48 | 306 | 15.43 | 22.24 | 971.47
9:49 | 367 | 13.99 | 22.25 | 971.47
9:50 | 428 | 12.6 | 22.25 | 971.44
9:51 | 489 | 11.34 | 22.24 | 971.43
9:52 | 550 | 10.09 | 22.25 | 971.4
9:53 | 611 | 9.09 | 22.26 | 971.41
9:54 | 672 | 8.09 | 22.27 | 971.4
9:55 | 733 | 7.2 | 22.29 | 971.4
9:56 | 794 | 6.37 | 22.27 | 971.39
9:57 | 855 | 5.7 | 22.27 | 971.39
9:58 | 916 | 5 | 22.29 | 971.41
9:59 | 977 | 4.38 | 22.29 | 971.43
10:00 | 1038 | 3.91 | 22.29 | 971.43
10:01 | 1099 | 3.4 | 22.29 | 971.37
10:02 | 1160 | 2.96 | 22.29 | 971.4
10:03 | 1221 | 2.63 | 22.29 | 971.4
10:04 | 1282 | 2.22 | 22.3 | 971.4
10:05 | 1343 | 1.97 | 22.3 | 971.36
10:06 | 1404 | 1.7 | 22.31 | 971.36
10:07 | 1465 | 1.42 | 22.31 | 971.39
10:08 | 1526 | 1.27 | 22.31 | 971.39
10:09 | 1587 | 1.07 | 22.3 | 971.4
10:10 | 1648 | 0.78 | 22.31 | 971.4
10:11 | 1709 | 0.65 | 22.31 | 971.4
10:12 | 1770 | 0.56 | 22.31 | 971.36
10:13 | 1831 | 0.46 | 22.31 | 971.36
10:14 | 1892 | 0.35 | 22.31 | 971.37
10:15 | 1953 | 0.2 | 22.32 | 971.38
10:16 | 2014 | 0.12 | 22.31 | 971.35
10:17 | 2075 | 0.07 | 22.31 | 971.36
10:18 | 2136 | 0.01 | 22.32 | 971.36
10:19 | 2197 | -0.11 | 22.31 | 971.38
10:20 | 2258 | -0.16 | 22.31 | 971.36
10:21 | 2319 | -0.19 | 22.31 | 971.35
10:22 | 2380 | -0.26 | 22.31 | 971.37
10:23 | 2441 | -0.22 | 22.32 | 971.33
10:24 | 2502 | -0.18 | 22.32 | 971.35
10:25 | 2563 | -0.3 | 22.32 | 971.35
10:26 | 2624 | -0.33 | 22.34 | 971.38
10:27 | 2685 | -0.42 | 22.33 | 971.37
10:28 | 2746 | -0.43 | 22.33 | 971.42
10:29 | 2807 | -0.46 | 22.33 | 971.4
10:30 | 2868 | -0.45 | 22.33 | 971.37
10:31 | 2929 | -0.49 | 22.32 | 971.38
10:32 | 2990 | -0.47 | 22.32 | 971.4
10:33 | 3051 | -0.5 | 22.34 | 971.38
10:34 | 3112 | -0.51 | 22.35 | 971.38
10:35 | 3173 | -0.55 | 22.35 | 971.38
10:36 | 3234 | -0.48 | 22.33 | 971.36
10:37 | 3295 | -0.53 | 22.33 | 971.36
10:38 | 3356 | -0.48 | 22.35 | 971.34
10:39 | 3417 | -0.51 | 22.36 | 971.33
10:40 | 3478 | -0.53 | 22.35 | 971.34
10:41 | 3539 | -0.5 | 22.36 | 971.33
10:42 | 3600 | -0.5 | 22.36 | 971.35
10:43 | 3661 | -0.52 | 22.36 | 971.35
10:44 | 3722 | -0.55 | 22.39 | 971.33
10:45 | 3783 | -0.54 | 22.41 | 971.35
10:46 | 3844 | -0.47 | 22.39 | 971.34
10:47 | 3905 | -0.56 | 22.4 | 971.36
10:48 | 3966 | -0.51 | 22.4 | 971.36
10:49 | 4027 | -0.57 | 22.38 | 971.33
10:50 | 4088 | -0.56 | 22.4 | 971.31
10:51 | 4149 | -0.56 | 22.41 | 971.31
10:52 | 4210 | -0.52 | 22.4 | 971.31
10:53 | 4271 | -0.54 | 22.39 | 971.31
10:54 | 4332 | -0.52 | 22.42 | 971.31
10:55 | 4393 | -0.57 | 22.42 | 971.32
10:56 | 4454 | -0.59 | 22.41 | 971.32

## QC3_Short_3
Test Date: | 2017-06-07 |  |  |  | Test Date: | 2017-06-07 |  |  |  | Test Date: | 2017-06-07 |  |  |  | Test Date: | 2017-06-08 |  |  |  | Test Date: | 2017-06-29 |  |  |  | Test Date: | 2017-06-29 |  |  |  | Test Date: | 2017-06-30 |  |  |  | Test Date: | 2017-07-06 |  |  |  | Test Date: | 2017-07-06
File: | QC3_GE11-X-S-CERN-0003_20170607_test2.xls |  |  |  | File: | QC3_GE11-X-S-CERN-0003_20170607_test3.xls |  |  |  | File: | QC3_GE11-X-S-CERN-0003_20170607.xls |  |  |  | File: | QC3_GE11-X-S-CERN-0003_20170608_test4.xls |  |  |  | File: | QC3_GE11-X-S-CERN-0003_20170629.xlsx |  |  |  | File: | QC3_GE11-X-S-CERN-0003_20170629_WITH_TAPE_ON_RO.xlsx |  |  |  | File: | QC3_GE11-X-S-CERN-0003_20170630_TapeOnLeakingScrewRO.xlsx |  |  |  | File: | QC3_GE11-X-S-CERN-0003_Calibration_Stand2 |  |  |  | File: | QC3_GE11-X-S-CERN-0003_20170706
Elog: | http://cmsonline.cern.ch/cms-elog/989551 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/989551 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/989551 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/989551 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/995372 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/995372 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/995372 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/995373 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/995373
Description: | test2:After putting tape on the RO screws |  |  |  | Description: | test3:After putting tape on the RO screws+Drift screws |  |  |  | Description: | Initial |  |  |  | Description: | Test4: Removed tape, Apply 1.3 Nm both on RO and Dr screws |  |  |  | Description: | Initial2 |  |  |  | Description: | With tape on entire RO |  |  |  | Description: | With tape on the leaking screw |  |  |  | Description: | Calibration |  |  |  | Description: | tape on Leaking screw on RO
Max P (mbar) | 24.13 |  |  |  | Max P (mbar) | 22.47 |  |  |  | Max P (mbar) | 24.13 |  |  |  | Max P (mbar) | 21.81 |  |  |  | Max P (mbar) | 22.37 |  |  |  | Max P (mbar) | 22.23 |  |  |  | Max P (mbar) | 24.88 |  |  |  | Max P (mbar) | 25.52 |  |  |  | Max P (mbar) | 24.13
Min P (mbar) | 18.62 |  |  |  | Min P (mbar) | 20.86 |  |  |  | Min P (mbar) | 19.59 |  |  |  | Min P (mbar) | 18.09 |  |  |  | Min P (mbar) | 18.79 |  |  |  | Min P (mbar) | 21.74 |  |  |  | Min P (mbar) | 24.52 |  |  |  | Min P (mbar) | 25.1 |  |  |  | Min P (mbar) | 22.75
Delta P (mbar) | 5.51 |  |  |  | Delta P (mbar) | 1.61 |  |  |  | Delta P (mbar) | 4.54 |  |  |  | Delta P (mbar) | 3.72 |  |  |  | Delta P (mbar) | 3.58 |  |  |  | Delta P (mbar) | 0.49 |  |  |  | Delta P (mbar) | 0.36 |  |  |  | Delta P (mbar) | 0.42 |  |  |  | Delta P (mbar) | 1.38
Min Time | 15:47 |  |  |  | Min Time | 17:22 |  |  |  | Min Time | 15:04 |  |  |  | Min Time | 9:51 |  |  |  | Min Time | 11:54 |  |  |  | Min Time | 14:29 |  |  |  | Min Time | 11:16 |  |  |  | Min Time | 15:09 |  |  |  | Min Time | 16:18
Max Time | 16:46 |  |  |  | Max Time | 18:21 |  |  |  | Max Time | 0:00 |  |  |  | Max Time | 10:50 |  |  |  | Max Time | 12:53 |  |  |  | Max Time | 15:28 |  |  |  | Max Time | 12:15 |  |  |  | Max Time | 16:08 |  |  |  | Max Time | 17:17
Delta T | 0:59 |  |  |  | Delta T | 0:59 |  |  |  | Delta T | -0.62758 |  |  |  | Delta T | 0:59 |  |  |  | Delta T | 0:59 |  |  |  | Delta T | 0:59 |  |  |  | Delta T | 0:59 |  |  |  | Delta T | 0:59 |  |  |  | Delta T | 0:59
Time (hr) | 0.98222 |  |  |  | Time (hr) | 0.98194 |  |  |  | Time (hr) | #NUM! |  |  |  | Time (hr) | 0.98222 |  |  |  | Time (hr) | 0.98361 |  |  |  | Time (hr) | 0.98361 |  |  |  | Time (hr) | 0.98361 |  |  |  | Time (hr) | 0.9825 |  |  |  | Time (hr) | 0.98194
Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar)
15:47 | 1 | 24.13 | 21.81 | 972.29 | 17:22 | 1 | 22.47 | 21.32 | 973.37 | 15:04 | 1 | 24.13 | 21.32 | 972.29 | 9:51 | 1 | 21.81 | 21.81 | 972.29 | 0.49601 | 1 | 22.37 | 22.3 | 956 | 0.60331 | 1 | 22.04 | 22.3 | 958.17 | 0.46917 | 1 | 24.69 | 22.3 | 962.51 | 0.63138 | 1 | 25.13 | 22.3 | 970.12 | 16:18 | 1 | 24.13 | 22.3 | 970.12
15:48 | 62 | 23.84 | 21.49 | 972.45 | 17:23 | 62 | 22.31 | 21.62 | 972.56 | 15:05 | 62 | 24.02 | 21.53 | 972.37 | 9:52 | 62 | 21.55 | 21.65 | 971.94 | 0.4967 | 62 | 22.16 | 22.32 | 956.17 | 0.604 | 62 | 22.2 | 22.36 | 958.15 | 0.46986 | 62 | 24.52 | 22.51 | 962.13 | 0.63207 | 62 | 25.1 | 22.25 | 970.05 | 16:19 | 62 | 24.02 | 22.24 | 969.63
15:49 | 123 | 23.74 | 21.54 | 972.43 | 17:24 | 123 | 22.27 | 21.6 | 972.54 | 15:06 | 123 | 23.79 | 21.53 | 972.36 | 9:53 | 123 | 21.46 | 21.65 | 971.98 | 0.4974 | 123 | 22.1 | 22.33 | 956.22 | 0.6047 | 123 | 22.21 | 22.37 | 958.19 | 0.47056 | 123 | 24.56 | 22.52 | 962.08 | 0.63277 | 123 | 25.14 | 22.24 | 970.07 | 16:20 | 123 | 23.86 | 22.25 | 969.63
15:50 | 184 | 23.59 | 21.53 | 972.43 | 17:25 | 184 | 22.24 | 21.62 | 972.5 | 15:07 | 184 | 23.64 | 21.54 | 972.37 | 9:54 | 184 | 21.32 | 21.67 | 971.98 | 0.49809 | 184 | 21.97 | 22.33 | 956.23 | 0.60539 | 184 | 22.2 | 22.39 | 958.18 | 0.47125 | 184 | 24.52 | 22.52 | 962.07 | 0.63346 | 184 | 25.13 | 22.24 | 970.03 | 16:21 | 184 | 23.82 | 22.24 | 969.63
15:51 | 245 | 23.5 | 21.53 | 972.48 | 17:26 | 245 | 22.18 | 21.63 | 972.49 | 15:08 | 245 | 23.42 | 21.52 | 972.36 | 9:55 | 245 | 21.23 | 21.66 | 971.98 | 0.49878 | 245 | 21.88 | 22.32 | 956.24 | 0.60609 | 245 | 22.22 | 22.36 | 958.16 | 0.47194 | 245 | 24.55 | 22.5 | 962.02 | 0.63416 | 245 | 25.22 | 22.23 | 970 | 16:22 | 245 | 23.71 | 22.22 | 969.65
15:52 | 306 | 23.27 | 21.54 | 972.48 | 17:27 | 306 | 22.19 | 21.61 | 972.47 | 15:09 | 306 | 23.38 | 21.52 | 972.33 | 9:56 | 306 | 21.15 | 21.66 | 971.95 | 0.49948 | 306 | 21.78 | 22.32 | 956.24 | 0.60678 | 306 | 22.23 | 22.35 | 958.16 | 0.47264 | 306 | 24.54 | 22.49 | 961.99 | 0.63485 | 306 | 25.22 | 22.24 | 970 | 16:23 | 306 | 23.71 | 22.25 | 969.65
15:53 | 367 | 23.15 | 21.55 | 972.49 | 17:28 | 367 | 22.14 | 21.62 | 972.48 | 15:10 | 367 | 23.12 | 21.49 | 972.37 | 9:57 | 367 | 21.09 | 21.66 | 971.96 | 0.50017 | 367 | 21.68 | 22.33 | 956.27 | 0.60748 | 367 | 22.22 | 22.36 | 958.17 | 0.47333 | 367 | 24.57 | 22.5 | 962.02 | 0.63554 | 367 | 25.31 | 22.25 | 970 | 16:24 | 367 | 23.67 | 22.23 | 969.69
15:54 | 428 | 23.03 | 21.56 | 972.52 | 17:29 | 428 | 22.12 | 21.65 | 972.43 | 15:11 | 428 | 22.99 | 21.51 | 972.37 | 9:58 | 428 | 21 | 21.67 | 971.89 | 0.50087 | 428 | 21.64 | 22.33 | 956.35 | 0.60817 | 428 | 22.17 | 22.38 | 958.2 | 0.47403 | 428 | 24.57 | 22.51 | 962.01 | 0.63624 | 428 | 25.42 | 22.23 | 969.99 | 16:25 | 428 | 23.71 | 22.25 | 969.69
15:55 | 489 | 22.9 | 21.57 | 972.53 | 17:30 | 489 | 22.04 | 21.65 | 972.43 | 15:12 | 489 | 22.85 | 21.48 | 972.38 | 9:59 | 489 | 20.83 | 21.67 | 971.91 | 0.50156 | 489 | 21.52 | 22.33 | 956.41 | 0.60887 | 489 | 22.11 | 22.42 | 958.17 | 0.47472 | 489 | 24.52 | 22.52 | 962.02 | 0.63693 | 489 | 25.4 | 22.24 | 969.98 | 16:26 | 489 | 23.61 | 22.25 | 969.63
15:56 | 550 | 22.78 | 21.57 | 972.53 | 17:31 | 550 | 21.97 | 21.63 | 972.43 | 15:13 | 550 | 22.7 | 21.51 | 972.38 | 10:00 | 550 | 20.82 | 21.68 | 971.9 | 0.50226 | 550 | 21.43 | 22.32 | 956.45 | 0.60956 | 550 | 22.14 | 22.45 | 958.21 | 0.47542 | 550 | 24.57 | 22.53 | 962.06 | 0.63763 | 550 | 25.48 | 22.24 | 969.91 | 16:27 | 550 | 23.55 | 22.24 | 969.62
15:57 | 611 | 22.66 | 21.57 | 972.53 | 17:32 | 611 | 21.94 | 21.62 | 972.44 | 15:14 | 611 | 22.41 | 21.52 | 972.39 | 10:01 | 611 | 20.76 | 21.68 | 971.84 | 0.50295 | 611 | 21.37 | 22.33 | 956.54 | 0.61025 | 611 | 22.13 | 22.43 | 958.17 | 0.47611 | 611 | 24.57 | 22.56 | 962.03 | 0.63832 | 611 | 25.42 | 22.25 | 969.93 | 16:28 | 611 | 23.58 | 22.23 | 969.65
15:58 | 672 | 22.5 | 21.55 | 972.53 | 17:33 | 672 | 21.84 | 21.62 | 972.43 | 15:15 | 672 | 22.28 | 21.49 | 972.39 | 10:02 | 672 | 20.67 | 21.66 | 971.85 | 0.50365 | 672 | 21.26 | 22.32 | 956.63 | 0.61095 | 672 | 22.19 | 22.44 | 958.17 | 0.47681 | 672 | 24.61 | 22.56 | 962.04 | 0.63902 | 672 | 25.47 | 22.24 | 969.93 | 16:29 | 671 | 23.58 | 22.23 | 969.65
15:59 | 733 | 22.41 | 21.57 | 972.55 | 17:34 | 733 | 21.85 | 21.62 | 972.41 | 15:16 | 733 | 22.17 | 21.51 | 972.4 | 10:03 | 733 | 20.63 | 21.66 | 971.85 | 0.50434 | 733 | 21.15 | 22.32 | 956.73 | 0.61164 | 733 | 22.14 | 22.46 | 958.16 | 0.47751 | 733 | 24.64 | 22.58 | 962.04 | 0.63971 | 733 | 25.51 | 22.23 | 969.94 | 16:30 | 733 | 23.54 | 22.24 | 969.64
16:00 | 794 | 22.21 | 21.56 | 972.56 | 17:35 | 794 | 21.88 | 21.6 | 972.45 | 15:17 | 794 | 22.01 | 21.5 | 972.42 | 10:04 | 794 | 20.5 | 21.68 | 971.86 | 0.50503 | 794 | 21.13 | 22.32 | 956.69 | 0.61234 | 794 | 22.13 | 22.44 | 958.15 | 0.47821 | 794 | 24.59 | 22.6 | 961.97 | 0.64041 | 794 | 25.44 | 22.25 | 969.9 | 16:31 | 794 | 23.49 | 22.24 | 969.66
16:01 | 855 | 22.11 | 21.55 | 972.51 | 17:36 | 855 | 21.76 | 21.64 | 972.4 | 15:18 | 855 | 21.82 | 21.53 | 972.42 | 10:05 | 855 | 20.39 | 21.67 | 971.9 | 0.50573 | 855 | 21.03 | 22.32 | 956.72 | 0.61303 | 855 | 22.08 | 22.45 | 958.18 | 0.4789 | 855 | 24.65 | 22.59 | 962.06 | 0.6411 | 855 | 25.5 | 22.25 | 969.93 | 16:32 | 855 | 23.43 | 22.25 | 969.57
16:02 | 916 | 21.96 | 21.57 | 972.5 | 17:37 | 916 | 21.79 | 21.63 | 972.42 | 15:19 | 916 | 21.7 | 21.5 | 972.43 | 10:06 | 916 | 20.32 | 21.68 | 971.87 | 0.50642 | 916 | 20.96 | 22.31 | 956.79 | 0.61373 | 916 | 22.11 | 22.42 | 958.2 | 0.47959 | 916 | 24.61 | 22.6 | 962.1 | 0.64179 | 916 | 25.45 | 22.24 | 969.92 | 16:33 | 916 | 23.44 | 22.24 | 969.59
16:03 | 977 | 21.87 | 21.54 | 972.53 | 17:38 | 977 | 21.82 | 21.63 | 972.42 | 15:20 | 977 | 21.58 | 21.54 | 972.39 | 10:07 | 977 | 20.21 | 21.67 | 971.91 | 0.50712 | 977 | 20.9 | 22.32 | 956.8 | 0.61442 | 977 | 22.08 | 22.42 | 958.18 | 0.48029 | 977 | 24.66 | 22.64 | 962.02 | 0.64248 | 977 | 25.5 | 22.24 | 969.92 | 16:34 | 977 | 23.44 | 22.24 | 969.54
16:04 | 1038 | 21.76 | 21.54 | 972.52 | 17:39 | 1038 | 21.72 | 21.63 | 972.45 | 15:21 | 1038 | 21.4 | 21.51 | 972.42 | 10:08 | 1038 | 20.12 | 21.68 | 971.91 | 0.50781 | 1038 | 20.74 | 22.33 | 956.83 | 0.61512 | 1038 | 22.12 | 22.45 | 958.19 | 0.48098 | 1038 | 24.63 | 22.63 | 962 | 0.64317 | 1038 | 25.37 | 22.25 | 969.93 | 16:35 | 1038 | 23.37 | 22.24 | 969.58
16:05 | 1099 | 21.59 | 21.54 | 972.5 | 17:40 | 1099 | 21.69 | 21.63 | 972.42 | 15:22 | 1099 | 21.21 | 21.53 | 972.42 | 10:09 | 1099 | 20.06 | 21.68 | 971.92 | 0.50851 | 1099 | 20.69 | 22.31 | 956.81 | 0.61581 | 1099 | 22.08 | 22.43 | 958.17 | 0.48168 | 1099 | 24.66 | 22.6 | 962.08 | 0.64387 | 1099 | 25.41 | 22.23 | 969.92 | 16:36 | 1099 | 23.41 | 22.24 | 969.6
16:06 | 1160 | 21.53 | 21.55 | 972.53 | 17:41 | 1160 | 21.61 | 21.63 | 972.43 | 15:23 | 1160 | 21.08 | 21.55 | 972.44 | 10:10 | 1160 | 20 | 21.67 | 971.92 | 0.5092 | 1160 | 20.64 | 22.33 | 956.8 | 0.6165 | 1160 | 22.08 | 22.43 | 958.19 | 0.48237 | 1160 | 24.59 | 22.62 | 962.04 | 0.64456 | 1160 | 25.39 | 22.22 | 969.95 | 16:37 | 1160 | 23.37 | 22.24 | 969.56
16:07 | 1221 | 21.35 | 21.56 | 972.61 | 17:42 | 1221 | 21.6 | 21.63 | 972.44 | 15:24 | 1221 | 21 | 21.54 | 972.44 | 10:11 | 1221 | 19.87 | 21.69 | 971.93 | 0.5099 | 1221 | 20.54 | 22.31 | 956.73 | 0.6172 | 1221 | 22.06 | 22.45 | 958.19 | 0.48307 | 1221 | 24.67 | 22.59 | 961.99 | 0.64525 | 1221 | 25.4 | 22.22 | 969.95 | 16:38 | 1221 | 23.36 | 22.24 | 969.53
16:08 | 1282 | 21.29 | 21.57 | 972.58 | 17:43 | 1282 | 21.58 | 21.64 | 972.44 | 15:25 | 1282 | 20.88 | 21.55 | 972.49 | 10:12 | 1282 | 19.86 | 21.67 | 971.94 | 0.51059 | 1282 | 20.48 | 22.31 | 956.7 | 0.61789 | 1282 | 21.97 | 22.46 | 958.22 | 0.48376 | 1282 | 24.73 | 22.59 | 962.04 | 0.64595 | 1282 | 25.46 | 22.23 | 969.89 | 16:39 | 1282 | 23.34 | 22.25 | 969.52
16:09 | 1343 | 21.15 | 21.58 | 972.59 | 17:44 | 1343 | 21.56 | 21.63 | 972.44 | 15:26 | 1343 | 20.73 | 21.53 | 972.52 | 10:13 | 1343 | 19.77 | 21.68 | 971.95 | 0.51128 | 1343 | 20.46 | 22.32 | 956.69 | 0.61859 | 1343 | 21.98 | 22.48 | 958.23 | 0.48446 | 1343 | 24.76 | 22.62 | 961.94 | 0.64664 | 1343 | 25.37 | 22.24 | 969.89 | 16:40 | 1343 | 23.32 | 22.24 | 969.55
16:10 | 1404 | 21.03 | 21.59 | 972.57 | 17:45 | 1404 | 21.59 | 21.65 | 972.43 | 15:27 | 1404 | 20.57 | 21.56 | 972.47 | 10:14 | 1404 | 19.68 | 21.7 | 971.92 | 0.51198 | 1404 | 20.39 | 22.32 | 956.7 | 0.61928 | 1404 | 21.96 | 22.49 | 958.23 | 0.48515 | 1404 | 24.75 | 22.65 | 961.93 | 0.64734 | 1404 | 25.47 | 22.23 | 969.86 | 16:41 | 1404 | 23.23 | 22.24 | 969.46
16:11 | 1465 | 20.9 | 21.6 | 972.55 | 17:46 | 1465 | 21.55 | 21.64 | 972.41 | 15:28 | 1465 | 20.41 | 21.55 | 972.49 | 10:15 | 1465 | 19.62 | 21.68 | 971.92 | 0.51267 | 1465 | 20.3 | 22.33 | 956.64 | 0.61998 | 1465 | 21.95 | 22.46 | 958.25 | 0.48584 | 1465 | 24.7 | 22.67 | 962.04 | 0.64803 | 1465 | 25.51 | 22.24 | 969.82 | 16:42 | 1465 | 23.27 | 22.24 | 969.44
16:12 | 1526 | 20.81 | 21.58 | 972.55 | 17:47 | 1526 | 21.53 | 21.64 | 972.4 | 15:29 | 1526 | 20.32 | 21.55 | 972.49 | 10:16 | 1526 | 19.56 | 21.68 | 971.95 | 0.51337 | 1526 | 20.28 | 22.34 | 956.62 | 0.62067 | 1526 | 21.93 | 22.45 | 958.22 | 0.48654 | 1526 | 24.79 | 22.66 | 961.96 | 0.64873 | 1526 | 25.48 | 22.25 | 969.87 | 16:43 | 1526 | 23.26 | 22.24 | 969.48
16:13 | 1587 | 20.73 | 21.58 | 972.52 | 17:48 | 1587 | 21.52 | 21.64 | 972.42 | 15:30 | 1587 | 20.18 | 21.55 | 972.49 | 10:17 | 1587 | 19.56 | 21.67 | 971.98 | 0.51406 | 1587 | 20.12 | 22.32 | 956.62 | 0.62138 | 1587 | 21.9 | 22.48 | 958.25 | 0.48723 | 1587 | 24.75 | 22.69 | 961.93 | 0.64942 | 1587 | 25.45 | 22.24 | 969.78 | 16:44 | 1587 | 23.26 | 22.24 | 969.41
16:14 | 1648 | 20.57 | 21.57 | 972.56 | 17:49 | 1648 | 21.45 | 21.66 | 972.39 | 15:31 | 1648 | 20.09 | 21.54 | 972.44 | 10:18 | 1648 | 19.39 | 21.69 | 971.95 | 0.51477 | 1648 | 20.09 | 22.34 | 956.66 | 0.62207 | 1648 | 21.93 | 22.48 | 958.25 | 0.48793 | 1648 | 24.77 | 22.7 | 961.89 | 0.65012 | 1648 | 25.48 | 22.23 | 969.8 | 16:45 | 1648 | 23.16 | 22.24 | 969.42
16:15 | 1709 | 20.4 | 21.58 | 972.57 | 17:50 | 1709 | 21.36 | 21.64 | 972.4 | 15:32 | 1709 | 19.85 | 21.55 | 972.45 | 10:19 | 1709 | 19.3 | 21.69 | 971.94 | 0.51546 | 1709 | 20.03 | 22.32 | 956.67 | 0.62277 | 1709 | 21.85 | 22.45 | 958.29 | 0.48862 | 1709 | 24.73 | 22.7 | 961.89 | 0.65081 | 1709 | 25.52 | 22.24 | 969.81 | 16:46 | 1709 | 23.17 | 22.25 | 969.37
16:16 | 1770 | 20.38 | 21.57 | 972.6 | 17:51 | 1770 | 21.36 | 21.64 | 972.44 | 15:33 | 1770 | 19.78 | 21.54 | 972.45 | 10:20 | 1770 | 19.21 | 21.7 | 971.96 | 0.51616 | 1770 | 19.94 | 22.33 | 956.66 | 0.62346 | 1770 | 21.9 | 22.44 | 958.3 | 0.48932 | 1770 | 24.77 | 22.72 | 961.88 | 0.6515 | 1770 | 25.46 | 22.25 | 969.83 | 16:47 | 1770 | 23.13 | 22.24 | 969.37
16:17 | 1831 | 20.24 | 21.58 | 972.61 | 17:52 | 1831 | 21.31 | 21.65 | 972.37 | 15:34 | 1831 | 19.59 | 21.5 | 972.44 | 10:21 | 1831 | 19.13 | 21.7 | 971.92 | 0.51685 | 1831 | 19.85 | 22.33 | 956.64 | 0.62416 | 1831 | 21.84 | 22.49 | 958.34 | 0.49001 | 1831 | 24.79 | 22.71 | 961.86 | 0.6522 | 1831 | 25.44 | 22.23 | 969.84 | 16:48 | 1831 | 23.11 | 22.22 | 969.36
16:18 | 1892 | 20.1 | 21.58 | 972.58 | 17:53 | 1892 | 21.25 | 21.65 | 972.4 | [5 empty] | 10:22 | 1892 | 19.06 | 21.69 | 971.89 | 0.51755 | 1892 | 19.81 | 22.34 | 956.67 | 0.62485 | 1892 | 21.83 | 22.49 | 958.35 | 0.49071 | 1892 | 24.81 | 22.72 | 961.86 | 0.65289 | 1892 | 25.49 | 22.24 | 969.89 | 16:49 | 1892 | 23.08 | 22.24 | 969.31
16:19 | 1953 | 20.08 | 21.57 | 972.65 | 17:54 | 1953 | 21.2 | 21.66 | 972.36 | [5 empty] | 10:23 | 1953 | 18.97 | 21.68 | 971.89 | 0.51824 | 1953 | 19.76 | 22.33 | 956.68 | 0.62554 | 1953 | 21.79 | 22.51 | 958.35 | 0.4914 | 1953 | 24.73 | 22.75 | 961.9 | 0.65359 | 1953 | 25.46 | 22.25 | 969.86 | 16:50 | 1953 | 23.04 | 22.24 | 969.37
16:20 | 2014 | 19.83 | 21.57 | 972.64 | 17:55 | 2014 | 21.24 | 21.64 | 972.37 | [5 empty] | 10:24 | 2014 | 18.94 | 21.69 | 971.84 | 0.51894 | 2014 | 19.73 | 22.33 | 956.68 | 0.62624 | 2014 | 21.82 | 22.49 | 958.35 | 0.49209 | 2014 | 24.78 | 22.75 | 961.85 | 0.65427 | 2014 | 25.5 | 22.23 | 969.87 | 16:51 | 2014 | 23.05 | 22.23 | 969.32
16:21 | 2075 | 19.76 | 21.57 | 972.63 | 17:56 | 2075 | 21.15 | 21.66 | 972.38 | [5 empty] | 10:25 | 2075 | 18.89 | 21.68 | 971.83 | 0.51963 | 2075 | 19.65 | 22.33 | 956.73 | 0.62693 | 2075 | 21.78 | 22.5 | 958.35 | 0.49279 | 2075 | 24.79 | 22.74 | 961.93 | 0.65497 | 2075 | 25.51 | 22.23 | 969.85 | 16:52 | 2075 | 23.01 | 22.23 | 969.32
16:22 | 2136 | 19.61 | 21.56 | 972.6 | 17:57 | 2136 | 21.14 | 21.64 | 972.35 | [5 empty] | 10:26 | 2136 | 18.72 | 21.69 | 971.85 | 0.52032 | 2136 | 19.48 | 22.33 | 956.75 | 0.62763 | 2136 | 21.74 | 22.46 | 958.35 | 0.49348 | 2136 | 24.78 | 22.74 | 961.92 | 0.65566 | 2136 | 25.45 | 22.24 | 969.85 | 16:53 | 2136 | 23 | 22.24 | 969.3
16:23 | 2197 | 19.56 | 21.58 | 972.58 | 17:58 | 2197 | 21.11 | 21.64 | 972.4 | [5 empty] | 10:27 | 2197 | 18.68 | 21.67 | 971.81 | 0.52102 | 2197 | 19.48 | 22.32 | 956.76 | 0.62832 | 2197 | 21.79 | 22.5 | 958.37 | 0.49418 | 2197 | 24.77 | 22.75 | 961.85 | 0.65635 | 2197 | 25.5 | 22.23 | 969.84 | 16:54 | 2197 | 22.94 | 22.24 | 969.26
16:24 | 2258 | 19.48 | 21.57 | 972.59 | 17:59 | 2258 | 21.09 | 21.63 | 972.4 | [5 empty] | 10:28 | 2258 | 18.61 | 21.69 | 971.8 | 0.52171 | 2258 | 19.37 | 22.33 | 956.77 | 0.62902 | 2258 | 21.81 | 22.47 | 958.36 | 0.49487 | 2258 | 24.83 | 22.74 | 961.88 | 0.65705 | 2258 | 25.44 | 22.23 | 969.81 | 16:55 | 2258 | 22.88 | 22.24 | 969.27
16:25 | 2319 | 19.3 | 21.57 | 972.54 | 18:00 | 2319 | 21 | 21.66 | 972.36 | [5 empty] | 10:29 | 2319 | 18.49 | 21.68 | 971.82 | 0.52241 | 2319 | 19.39 | 22.33 | 956.79 | 0.62971 | 2319 | 21.82 | 22.49 | 958.34 | 0.49557 | 2319 | 24.82 | 22.74 | 961.92 | 0.65774 | 2319 | 25.41 | 22.22 | 969.82 | 16:56 | 2319 | 22.91 | 22.21 | 969.28
16:26 | 2380 | 19.24 | 21.56 | 972.56 | 18:01 | 2380 | 21 | 21.65 | 972.36 | [5 empty] | 10:30 | 2380 | 18.43 | 21.69 | 971.78 | 0.5231 | 2380 | 19.27 | 22.33 | 956.78 | 0.63041 | 2380 | 21.79 | 22.48 | 958.37 | 0.49626 | 2380 | 24.86 | 22.75 | 961.89 | 0.65844 | 2380 | 25.46 | 22.24 | 969.78 | 16:57 | 2380 | 22.82 | 22.21 | 969.24
16:27 | 2441 | 19.07 | 21.55 | 972.58 | 18:02 | 2441 | 20.97 | 21.64 | 972.37 | [5 empty] | 10:31 | 2441 | 18.35 | 21.68 | 971.82 | 0.5238 | 2441 | 19.18 | 22.33 | 956.82 | 0.6311 | 2441 | 21.81 | 22.47 | 958.3 | 0.49696 | 2441 | 24.87 | 22.76 | 961.93 | 0.65913 | 2441 | 25.45 | 22.23 | 969.74 | 16:58 | 2441 | 22.78 | 22.24 | 969.25
16:28 | 2502 | 18.99 | 21.56 | 972.66 | 18:03 | 2502 | 20.93 | 21.65 | 972.37 | [5 empty] | 10:32 | 2502 | 18.28 | 21.67 | 971.85 | 0.52449 | 2502 | 19.07 | 22.31 | 956.84 | 0.63179 | 2502 | 21.8 | 22.45 | 958.25 | 0.49765 | 2502 | 24.88 | 22.75 | 961.9 | 0.65983 | 2502 | 25.45 | 22.23 | 969.71 | 16:59 | 2502 | 22.81 | 22.22 | 969.24
16:29 | 2563 | 18.89 | 21.56 | 972.62 | 18:04 | 2563 | 20.92 | 21.66 | 972.37 | [5 empty] | 10:33 | 2563 | 18.22 | 21.66 | 971.8 | 0.52519 | 2563 | 18.91 | 22.31 | 956.93 | 0.63249 | 2563 | 21.79 | 22.47 | 958.22 | 0.49834 | 2563 | 24.82 | 22.74 | 961.89 | 0.66052 | 2563 | 25.42 | 22.22 | 969.69 | 17:00 | 2563 | 22.81 | 22.23 | 969.24
16:30 | 2624 | 18.84 | 21.54 | 972.62 | 18:05 | 2624 | 20.92 | 21.65 | 972.36 | [5 empty] | 10:34 | 2624 | 18.14 | 21.67 | 971.81 | 0.52588 | 2624 | 18.93 | 22.32 | 956.96 | 0.63318 | 2624 | 21.82 | 22.47 | 958.23 | 0.49904 | 2624 | 24.86 | 22.76 | 961.88 | 0.66122 | 2624 | 25.42 | 22.22 | 969.7 | 17:01 | 2624 | 22.81 | 22.23 | 969.22
16:31 | 2685 | 18.66 | 21.54 | 972.6 | 18:06 | 2685 | 20.86 | 21.65 | 972.35 | [5 empty] | 10:35 | 2685 | 18.12 | 21.68 | 971.8 | 0.52657 | 2685 | 18.87 | 22.33 | 956.97 | 0.63388 | 2685 | 21.8 | 22.49 | 958.22 | 0.49973 | 2685 | 24.86 | 22.74 | 961.9 | 0.66191 | 2685 | 25.47 | 22.23 | 969.68 | 17:02 | 2685 | 22.75 | 22.21 | 969.22
16:32 | 2746 | 18.62 | 21.57 | 972.61 | 18:07 | 2746 | 20.89 | 21.67 | 972.44 | [5 empty] | 10:36 | 2746 | 18.09 | 21.67 | 971.71 | 0.52727 | 2746 | 18.79 | 22.31 | 957.02 | 0.63457 | 2746 | 21.83 | 22.49 | 958.23 | 0.50043 | 2746 | 24.87 | 22.73 | 961.84 | 0.6626 | 2746 | 25.44 | 22.21 | 969.69 | 17:03 | 2746 | 22.81 | 22.21 | 969.24
16:33 | 2807 | 18.38 | 21.54 | 972.56 | 18:08 | 2807 | 20.85 | 21.65 | 972.37 | [5 empty] | 10:37 | 2807 | 18.01 | 21.68 | 971.68 | 0.52796 | 2807 | 18.67 | 22.32 | 957.01 | 0.63527 | 2807 | 21.79 | 22.49 | 958.25 | 0.50112 | 2807 | 24.93 | 22.72 | 961.84 | 0.6633 | 2807 | 25.42 | 22.23 | 969.66 | 17:04 | 2807 | 22.79 | 22.23 | 969.18
16:34 | 2868 | 18.32 | 21.56 | 972.58 | 18:09 | 2868 | 20.8 | 21.65 | 972.35 | [5 empty] | 10:38 | 2868 | 17.94 | 21.68 | 971.69 | 0.52866 | 2868 | 18.64 | 22.33 | 956.99 | 0.63596 | 2868 | 21.79 | 22.51 | 958.22 | 0.50182 | 2868 | 24.91 | 22.72 | 961.84 | 0.66399 | 2868 | 25.39 | 22.21 | 969.71 | 17:05 | 2868 | 22.72 | 22.21 | 969.2
16:35 | 2929 | 18.22 | 21.56 | 972.57 | 18:10 | 2929 | 20.73 | 21.68 | 972.35 | [5 empty] | 10:39 | 2929 | 17.86 | 21.68 | 971.67 | 0.52935 | 2929 | 18.56 | 22.32 | 956.99 | 0.63666 | 2929 | 21.78 | 22.53 | 958.22 | 0.50251 | 2929 | 24.9 | 22.71 | 961.86 | 0.66469 | 2929 | 25.34 | 22.22 | 969.69 | 17:06 | 2929 | 22.64 | 22.21 | 969.16
16:36 | 2990 | 18.11 | 21.55 | 972.65 | 18:11 | 2990 | 20.75 | 21.67 | 972.36 | [5 empty] | 10:40 | 2990 | 17.83 | 21.68 | 971.65 | 0.53005 | 2990 | 18.53 | 22.34 | 957 | 0.63735 | 2990 | 21.75 | 22.51 | 958.18 | 0.50321 | 2990 | 24.9 | 22.72 | 961.83 | 0.66537 | 2990 | 25.34 | 22.22 | 969.7 | 17:07 | 2990 | 22.68 | 22.22 | 969.17
16:37 | 3051 | 18 | 21.55 | 972.62 | 18:12 | 3051 | 20.63 | 21.68 | 972.35 | [5 empty] | 10:41 | 3051 | 17.73 | 21.68 | 971.61 | 0.53074 | 3051 | 18.5 | 22.31 | 957.05 | 0.63804 | 3051 | 21.82 | 22.46 | 958.15 | 0.5039 | 3051 | 24.97 | 22.71 | 961.79 | 0.66606 | 3051 | 25.27 | 22.22 | 969.71 | 17:08 | 3051 | 22.6 | 22.22 | 969.13
16:38 | 3112 | 17.94 | 21.55 | 972.58 | 18:13 | 3112 | 20.64 | 21.69 | 972.34 | [5 empty] | 10:42 | 3112 | 17.71 | 21.68 | 971.65 | 0.53144 | 3112 | 18.36 | 22.31 | 957.03 | 0.63874 | 3112 | 21.81 | 22.41 | 958.1 | 0.50459 | 3112 | 24.95 | 22.71 | 961.81 | 0.66676 | 3112 | 25.25 | 22.21 | 969.69 | 17:09 | 3112 | 22.64 | 22.23 | 969.15
16:39 | 3173 | 17.83 | 21.55 | 972.62 | 18:14 | 3173 | 20.56 | 21.67 | 972.35 | [5 empty] | 10:43 | 3173 | 17.67 | 21.69 | 971.63 | 0.53213 | 3173 | 18.35 | 22.3 | 957.06 | 0.63943 | 3173 | 21.81 | 22.42 | 958.03 | 0.50529 | 3173 | 24.89 | 22.71 | 961.76 | 0.66745 | 3173 | 25.28 | 22.22 | 969.7 | 17:10 | 3173 | 22.66 | 22.21 | 969.17
16:40 | 3234 | 17.73 | 21.56 | 972.6 | 18:15 | 3234 | 20.53 | 21.68 | 972.33 | [5 empty] | 10:44 | 3234 | 17.64 | 21.7 | 971.63 | 0.53282 | 3234 | 18.3 | 22.31 | 957.05 | 0.64013 | 3234 | 21.79 | 22.44 | 958.01 | 0.50598 | 3234 | 24.95 | 22.71 | 961.73 | 0.66815 | 3234 | 25.22 | 22.23 | 969.67 | 17:11 | 3234 | 22.54 | 22.23 | 969.14
16:41 | 3295 | 17.67 | 21.58 | 972.58 | 18:16 | 3295 | 20.5 | 21.66 | 972.32 | [5 empty] | 10:45 | 3295 | 17.55 | 21.69 | 971.62 | 0.53352 | 3295 | 18.22 | 22.32 | 957.05 | 0.64082 | 3295 | 21.78 | 22.49 | 958.06 | 0.50668 | 3295 | 24.92 | 22.68 | 961.79 | 0.66884 | 3295 | 25.23 | 22.23 | 969.68 | 17:12 | 3295 | 22.57 | 22.22 | 969.16
16:42 | 3356 | 17.53 | 21.58 | 972.63 | 18:17 | 3356 | 20.49 | 21.69 | 972.3 | [5 empty] | 10:46 | 3356 | 17.5 | 21.7 | 971.65 | 0.53421 | 3356 | 18.16 | 22.33 | 957.06 | 0.64152 | 3356 | 21.65 | 22.46 | 958.12 | 0.50737 | 3356 | 24.93 | 22.65 | 961.86 | 0.66954 | 3356 | 25.25 | 22.23 | 969.63 | 17:13 | 3356 | 22.47 | 22.22 | 969.15
16:43 | 3417 | 17.45 | 21.58 | 972.62 | 18:18 | 3417 | 20.44 | 21.66 | 972.3 | [5 empty] | 10:47 | 3417 | 17.48 | 21.69 | 971.68 | 0.53491 | 3417 | 18.06 | 22.32 | 957.08 | 0.64221 | 3417 | 21.74 | 22.43 | 958.13 | 0.50807 | 3417 | 24.91 | 22.63 | 961.91 | 0.67023 | 3417 | 25.27 | 22.22 | 969.65 | 17:14 | 3417 | 22.44 | 22.22 | 969.13
16:44 | 3478 | 17.42 | 21.58 | 972.61 | 18:19 | 3478 | 20.43 | 21.69 | 972.3 | [5 empty] | 10:48 | 3478 | 17.36 | 21.71 | 971.66 | 0.5356 | 3478 | 18.06 | 22.31 | 957.07 | 0.64291 | 3478 | 21.68 | 22.47 | 958.11 | 0.50876 | 3478 | 24.93 | 22.64 | 961.78 | 0.67093 | 3478 | 25.27 | 22.23 | 969.64 | 17:15 | 3478 | 22.5 | 22.22 | 969.09
16:45 | 3539 | 17.23 | 21.57 | 972.62 | 18:20 | 3539 | 20.4 | 21.68 | 972.31 | [5 empty] | 10:49 | 3539 | 17.25 | 21.71 | 971.62 | 0.5363 | 3539 | 18.02 | 22.31 | 957.07 | 0.6436 | 3539 | 21.68 | 22.44 | 958.13 | 0.50946 | 3539 | 24.93 | 22.62 | 961.75 | 0.67162 | 3539 | 25.28 | 22.22 | 969.63 | 17:16 | 3539 | 22.41 | 22.21 | 969.14
16:46 | 3600 | 17.15 | 21.6 | 972.62 | 18:21 | 3600 | 20.4 | 21.68 | 972.29 | [5 empty] | 10:50 | 3600 | 17.16 | 21.71 | 971.67 | 0.53699 | 3600 | 17.96 | 22.31 | 957.09 | 0.64429 | 3600 | 21.71 | 22.43 | 958.13 | 0.51015 | 3600 | 24.92 | 22.63 | 961.62 | 0.67231 | 3600 | 25.32 | 22.23 | 969.59 | 17:17 | 3600 | 22.48 | 22.2 | 969.16
[15 empty] | 10:51 | 3661 | 17.23 | 21.71 | 971.66 | 0.53769 | 3661 | 17.85 | 22.31 | 957.09 | 0.64499 | 3661 | 21.61 | 22.44 | 958.12 | 0.51084 | 3661 | 24.93 | 22.63 | 961.59 | [5 empty] | 17:18 | 3661 | 22.41 | 22.21 | 969.15
[15 empty] | 10:52 | 3722 | 17.13 | 21.72 | 971.64 | 0.53838 | 3722 | 17.8 | 22.32 | 957.07 | 0.64568 | 3722 | 21.68 | 22.48 | 958.13 | [10 empty] | 17:19 | 3722 | 22.47 | 22.19 | 969.13
[20 empty] | 0.53907 | 3783 | 17.76 | 22.31 | 957.07 | [15 empty] | 17:20 | 3783 | 22.36 | 22.2 | 969.09
[20 empty] | 0.53977 | 3844 | 17.73 | 22.31 | 957.07

## QC3_Short_4
Test Date: | 2017-07-05 |  |  |  | Test Date: | 2017-07-06
File: | QC3_GE11-X-S-CERN-0004_20170705.xlsx |  |  |  | File: | QC3_GE11-X-S-CERN-0004_20170706
Elog: | http://cmsonline.cern.ch/cms-elog/995243 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/995287
Description: | Initial |  |  |  | Description: | After Applying 1.3 Nm on RO and DR screws
Max P (mbar) | 23.36 |  |  |  | Max P (mbar) | 22.7
Min P (mbar) | 19.51 |  |  |  | Min P (mbar) | 21.67
Delta P (mbar) | 3.85 |  |  |  | Delta P (mbar) | 1.03
Min Time | 18:30 |  |  |  | Min Time | 14:20
Max Time | 19:29 |  |  |  | Max Time | 15:19
Delta T | 0:59 |  |  |  | Delta T | 0:59
Time (hr) | 0.98361 |  |  |  | Time (hr) | 0.98333
Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar)
0.77086 | 1 | 23.36 | 22.3 | 971.2 | 14:20 | 1 | 23.03 | 22.3 | 970.12
0.77155 | 62 | 23.14 | 22.3 | 970.74 | 14:21 | 62 | 22.7 | 22.61 | 970.97
0.77225 | 123 | 23.02 | 22.31 | 970.76 | 14:22 | 123 | 22.65 | 22.62 | 970.94
0.77294 | 184 | 22.88 | 22.29 | 970.75 | 14:23 | 184 | 22.55 | 22.61 | 970.94
0.77365 | 245 | 22.8 | 22.3 | 970.74 | 14:24 | 245 | 22.58 | 22.6 | 970.91
0.77433 | 306 | 22.72 | 22.3 | 970.74 | 14:25 | 306 | 22.48 | 22.58 | 970.9
0.77503 | 367 | 22.54 | 22.3 | 970.72 | 14:26 | 367 | 22.44 | 22.58 | 970.97
0.77573 | 428 | 22.46 | 22.3 | 970.75 | 14:27 | 428 | 22.47 | 22.58 | 970.93
0.77642 | 489 | 22.42 | 22.3 | 970.74 | 14:28 | 489 | 22.5 | 22.56 | 970.93
0.77712 | 550 | 22.36 | 22.29 | 970.72 | 14:29 | 550 | 22.43 | 22.58 | 970.94
0.77781 | 611 | 22.25 | 22.31 | 970.68 | 14:30 | 611 | 22.37 | 22.57 | 970.9
0.77851 | 672 | 22.16 | 22.31 | 970.69 | 14:31 | 672 | 22.4 | 22.58 | 970.93
0.7792 | 733 | 22.1 | 22.29 | 970.68 | 14:32 | 733 | 22.37 | 22.57 | 970.88
0.7799 | 794 | 22.04 | 22.31 | 970.66 | 14:33 | 794 | 22.36 | 22.59 | 970.85
0.78059 | 855 | 21.98 | 22.32 | 970.63 | 14:34 | 855 | 22.32 | 22.56 | 970.89
0.78128 | 916 | 21.9 | 22.32 | 970.6 | 14:35 | 916 | 22.31 | 22.58 | 970.89
0.78198 | 977 | 21.83 | 22.33 | 970.56 | 14:36 | 977 | 22.2 | 22.57 | 970.89
0.78267 | 1038 | 21.77 | 22.36 | 970.56 | 14:37 | 1038 | 22.26 | 22.58 | 970.85
0.78337 | 1099 | 21.65 | 22.38 | 970.53 | 14:38 | 1099 | 22.18 | 22.57 | 970.88
0.78406 | 1160 | 21.58 | 22.41 | 970.56 | 14:39 | 1160 | 22.15 | 22.56 | 970.83
0.78476 | 1221 | 21.55 | 22.41 | 970.55 | 14:40 | 1221 | 22.21 | 22.56 | 970.82
0.78545 | 1282 | 21.46 | 22.41 | 970.56 | 14:41 | 1282 | 22.12 | 22.56 | 970.86
0.78615 | 1343 | 21.45 | 22.43 | 970.53 | 14:42 | 1343 | 22.11 | 22.57 | 970.87
0.78684 | 1404 | 21.37 | 22.47 | 970.51 | 14:43 | 1404 | 22.08 | 22.57 | 970.92
0.78753 | 1465 | 21.29 | 22.46 | 970.51 | 14:44 | 1465 | 22.08 | 22.59 | 970.89
0.78823 | 1526 | 21.3 | 22.49 | 970.52 | 14:45 | 1526 | 22.1 | 22.59 | 970.84
0.78892 | 1587 | 21.17 | 22.51 | 970.46 | 14:46 | 1587 | 22.07 | 22.58 | 970.87
0.78962 | 1648 | 21.16 | 22.5 | 970.47 | 14:47 | 1648 | 22.05 | 22.59 | 970.86
0.79031 | 1709 | 21.11 | 22.5 | 970.46 | 14:48 | 1709 | 22.03 | 22.58 | 970.86
0.79101 | 1770 | 21.03 | 22.51 | 970.47 | 14:49 | 1770 | 22 | 22.59 | 970.81
0.7917 | 1831 | 21.02 | 22.51 | 970.45 | 14:50 | 1831 | 21.96 | 22.59 | 970.81
0.7924 | 1892 | 20.97 | 22.47 | 970.44 | 14:51 | 1892 | 21.99 | 22.58 | 970.78
0.79309 | 1953 | 20.93 | 22.47 | 970.41 | 14:52 | 1953 | 21.95 | 22.6 | 970.79
0.79382 | 2014 | 20.9 | 22.48 | 970.43 | 14:53 | 2014 | 21.95 | 22.6 | 970.79
0.79448 | 2075 | 20.73 | 22.49 | 970.44 | 14:54 | 2075 | 21.98 | 22.59 | 970.81
0.79517 | 2136 | 20.7 | 22.5 | 970.44 | 14:55 | 2136 | 21.97 | 22.59 | 970.75
0.79587 | 2197 | 20.67 | 22.51 | 970.4 | 14:56 | 2197 | 21.88 | 22.61 | 970.79
0.79656 | 2258 | 20.62 | 22.48 | 970.47 | 14:57 | 2258 | 21.92 | 22.58 | 970.76
0.79726 | 2319 | 20.53 | 22.5 | 970.46 | 14:58 | 2319 | 21.86 | 22.6 | 970.81
0.79795 | 2380 | 20.55 | 22.47 | 970.46 | 14:59 | 2380 | 21.9 | 22.59 | 970.85
0.79865 | 2441 | 20.49 | 22.48 | 970.49 | 15:00 | 2441 | 21.81 | 22.6 | 970.84
0.79934 | 2502 | 20.46 | 22.49 | 970.51 | 15:01 | 2502 | 21.76 | 22.61 | 970.84
0.80003 | 2563 | 20.41 | 22.52 | 970.48 | 15:02 | 2563 | 21.79 | 22.62 | 970.81
0.80073 | 2624 | 20.36 | 22.51 | 970.51 | 15:03 | 2624 | 21.79 | 22.61 | 970.82
0.80142 | 2685 | 20.34 | 22.52 | 970.54 | 15:04 | 2685 | 21.79 | 22.62 | 970.75
0.80212 | 2746 | 20.32 | 22.53 | 970.5 | 15:05 | 2746 | 21.79 | 22.6 | 970.73
0.80281 | 2807 | 20.29 | 22.53 | 970.51 | 15:06 | 2807 | 21.76 | 22.59 | 970.73
0.80351 | 2868 | 20.13 | 22.55 | 970.51 | 15:07 | 2868 | 21.73 | 22.57 | 970.76
0.8042 | 2929 | 20.14 | 22.55 | 970.52 | 15:08 | 2929 | 21.79 | 22.56 | 970.72
0.8049 | 2990 | 20.1 | 22.55 | 970.53 | 15:09 | 2990 | 21.74 | 22.54 | 970.71
0.80559 | 3051 | 20.07 | 22.54 | 970.56 | 15:10 | 3051 | 21.71 | 22.58 | 970.71
0.80628 | 3112 | 20 | 22.54 | 970.55 | 15:11 | 3112 | 21.72 | 22.59 | 970.64
0.80698 | 3173 | 19.97 | 22.55 | 970.58 | 15:12 | 3173 | 21.7 | 22.6 | 970.59
0.80767 | 3234 | 19.92 | 22.56 | 970.52 | 15:13 | 3234 | 21.71 | 22.54 | 970.54
0.80837 | 3295 | 19.87 | 22.58 | 970.56 | 15:14 | 3295 | 21.72 | 22.53 | 970.5
0.80906 | 3356 | 19.77 | 22.55 | 970.56 | 15:15 | 3356 | 21.68 | 22.55 | 970.46
0.80976 | 3417 | 19.69 | 22.53 | 970.6 | 15:16 | 3417 | 21.72 | 22.55 | 970.48
0.81045 | 3478 | 19.68 | 22.53 | 970.59 | 15:17 | 3478 | 21.71 | 22.56 | 970.39
0.81115 | 3539 | 19.62 | 22.53 | 970.61 | 15:18 | 3539 | 21.71 | 22.54 | 970.37
0.81184 | 3600 | 19.51 | 22.52 | 970.59 | 15:19 | 3600 | 21.67 | 22.57 | 970.37
[5 empty] | 15:20 | 3661 | 21.67 | 22.56 | 970.35
[5 empty] | 15:21 | 3722 | 21.65 | 22.56 | 970.35
[5 empty] | 15:22 | 3783 | 21.64 | 22.57 | 970.35
[5 empty] | 15:23 | 3844 | 21.65 | 22.57 | 970.35
[5 empty] | 15:24 | 3905 | 21.63 | 22.56 | 970.36

## QC3_Long_1
Test Date: | 2017-05-17 |  |  |  | Test Date: | 2017-05-17 |  |  |  | Test Date: | 2017-05-22 |  |  |  | Test Date: | 2017-05-22 |  |  |  | Test Date: | 2017-05-30 |  |  |  | Test Date: | 2017-06-01 |  |  |  | Test Date: | 2017-07-13 |  |  |  | Test Date: | 2017-07-14
File: | QC3_GE11-X-L-CERN-0001_Calibration_2017_05_17.xls |  |  |  | File: | QC3_GE11-X-L-CERN-0001_2017_05_17.xls |  |  |  | File: | QC3_GE11-X-L-CERN-0001_2017_05_22_test1.xls |  |  |  | File: | QC3_GE11-X-L-CERN-0001_2017_05_22_test2.xls |  |  |  | File: | QC3_GE11-X-L-CERN-0001_2017_05_30.xls |  |  |  | File: | QC3_GE11-X-L-CERN-0001_20170601.xls |  |  |  | File: | QC3_GE11-X-L-CERN-0001_20170714_Test1.xlsx |  |  |  | File: | QC3_GE11-X-L-CERN-0001_20170714_Test2.xlsx
Elog: | http://cmsonline.cern.ch/cms-elog/983820 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/983820 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/984997 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/985202 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/987828 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/988020 |  |  |  | Elog: |  |  |  |  | Elog:
Description: | Calibration run |  |  |  | Description: | Initial test |  |  |  | Description: | After Replacing the RO and EF |  |  |  | Description: | Found leaking screw replaced |  |  |  | Description: | After fixing the short found in GEM1 |  |  |  | Description: | After fixing the bending found on the RO |  |  |  | Description: | After bending the RO |  |  |  | Description: | After applying 1.3Nm torque on RO and DR screws
Max P (mbar) | 28.45 |  |  |  | Max P (mbar) | 23.69 |  |  |  | Max P (mbar) | 22.04 |  |  |  | Max P (mbar) | 22.45 |  |  |  | Max P (mbar) | 25.16 |  |  |  | Max P (mbar) | 22.22 |  |  |  | Max P (mbar) | 27.34 |  |  |  | Max P (mbar) | 26.07
Min P (mbar) | 27.5 |  |  |  | Min P (mbar) | 8.69 |  |  |  | Min P (mbar) | 1.36 |  |  |  | Min P (mbar) | 21.6 |  |  |  | Min P (mbar) | 24.49 |  |  |  | Min P (mbar) | 21.87 |  |  |  | Min P (mbar) | 23.62 |  |  |  | Min P (mbar) | 24.7
Delta P (mbar) | 0.95 |  |  |  | Delta P (mbar) | 15 |  |  |  | Delta P (mbar) | 20.68 |  |  |  | Delta P (mbar) | 0.85 |  |  |  | Delta P (mbar) | 0.67 |  |  |  | Delta P (mbar) | 0.35 |  |  |  | Delta P (mbar) | 3.72 |  |  |  | Delta P (mbar) | 1.37
Min Time | 14:44 |  |  |  | Min Time | 15:50 |  |  |  | Min Time | 18:06 |  |  |  | Min Time | 15:49 |  |  |  | Min Time | 16:30 |  |  |  | Min Time | 14:21 |  |  |  | Min Time | 17:56 |  |  |  | Min Time | 9:07
Max Time | 15:29 |  |  |  | Max Time | 16:35 |  |  |  | Max Time | 18:51 |  |  |  | Max Time | 16:34 |  |  |  | Max Time | 17:15 |  |  |  | Max Time | 15:24 |  |  |  | Max Time | 18:56 |  |  |  | Max Time | 10:07
Delta T | 0:45 |  |  |  | Delta T | 0:45 |  |  |  | Delta T | 0:45 |  |  |  | Delta T | 0:45 |  |  |  | Delta T | 0:45 |  |  |  | Delta T | 1:03 |  |  |  | Delta T | 1:00 |  |  |  | Delta T | 1:00
Time (hr) | 0.75028 |  |  |  | Time (hr) | 0.75028 |  |  |  | Time (hr) | 0.75028 |  |  |  | Time (hr) | 0.75 |  |  |  | Time (hr) | 0.75 |  |  |  | Time (hr) | 1.0503 |  |  |  | Time (hr) | 1.0006 |  |  |  | Time (hr) | 1.0003
Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar)
14:44 | 1 | 27.67 | 21.81 | 971.2 | 15:50 | 1 | 23.69 | 21.81 | 970.12 | 18:06 | 1 | 22.04 | 21.81 | 970.12 | 15:49 | 1 | 22.37 | 21.81 | 970.12 | 16:30 | 1 | 25.02 | 21.81 | 967.94 | 14:21 | 1 | 22.04 | 22.78 | 367.31 | 17:56 | 1 | 27.34 | 22.3 | 970.12 | 9:07 | 1 | 26.01 | 22.3 | 971.2
14:45 | 62 | 27.5 | 21.9 | 971.12 | 15:51 | 62 | 22.99 | 21.97 | 970.04 | 18:07 | 62 | 20.96 | 21.93 | 970.07 | 15:50 | 62 | 22.45 | 21.92 | 969.22 | 16:31 | 62 | 25.16 | 22.08 | 968.12 | 14:22 | 62 | 22.13 | 22.31 | 276.32 | 17:57 | 62 | 26.93 | 22.37 | 969.2 | 9:08 | 62 | 26.07 | 22.31 | 971.1
14:46 | 123 | 27.61 | 21.9 | 971.11 | 15:52 | 123 | 22.52 | 21.97 | 970.06 | 18:08 | 123 | 19.99 | 21.93 | 970.06 | 15:51 | 123 | 22.45 | 21.91 | 969.25 | 16:32 | 123 | 25.14 | 22.13 | 968.08 | 14:23 | 123 | 22.03 | 22.3 | 275.93 | 17:58 | 123 | 26.69 | 22.39 | 969.2 | 9:09 | 123 | 26.06 | 22.3 | 971.09
14:47 | 184 | 27.61 | 21.9 | 971.08 | 15:53 | 184 | 21.97 | 21.98 | 970.03 | 18:09 | 184 | 19.08 | 21.91 | 970.04 | 15:52 | 184 | 22.38 | 21.9 | 969.26 | 16:33 | 184 | 25.12 | 22.13 | 968.09 | 14:24 | 184 | 21.99 | 22.31 | 275.8 | 17:59 | 184 | 26.6 | 22.38 | 969.2 | 9:10 | 184 | 26.01 | 22.29 | 971.12
14:48 | 245 | 27.63 | 21.9 | 971.04 | 15:54 | 245 | 21.51 | 21.98 | 970 | 18:10 | 245 | 18.13 | 21.91 | 970.06 | 15:53 | 245 | 22.4 | 21.91 | 969.25 | 16:34 | 245 | 25.06 | 22.14 | 968.08 | 14:25 | 245 | 22.05 | 22.31 | 275.73 | 18:00 | 245 | 26.39 | 22.36 | 969.19 | 9:11 | 245 | 26.02 | 22.28 | 971.13
14:49 | 306 | 27.61 | 21.88 | 971.05 | 15:55 | 306 | 21.01 | 21.97 | 970 | 18:11 | 306 | 17.3 | 21.91 | 970.07 | 15:54 | 306 | 22.33 | 21.91 | 969.25 | 16:35 | 306 | 25.11 | 22.1 | 968.08 | 14:26 | 306 | 21.96 | 22.31 | 275.65 | 18:01 | 306 | 26.35 | 22.36 | 969.17 | 9:12 | 306 | 26.05 | 22.29 | 971.13
14:50 | 367 | 27.65 | 21.89 | 971.07 | 15:56 | 367 | 20.56 | 21.97 | 970.01 | 18:12 | 367 | 16.46 | 21.91 | 970.05 | 15:55 | 367 | 22.29 | 21.91 | 969.23 | 16:36 | 367 | 25.05 | 22.1 | 968.1 | 14:27 | 367 | 22 | 22.3 | 275.6 | 18:02 | 367 | 26.2 | 22.35 | 969.19 | 9:13 | 367 | 26.01 | 22.31 | 971.1
14:51 | 428 | 27.64 | 21.9 | 970.99 | 15:57 | 428 | 20.2 | 21.97 | 970.03 | 18:13 | 428 | 15.62 | 21.91 | 970.07 | 15:56 | 428 | 22.27 | 21.91 | 969.22 | 16:37 | 428 | 25.05 | 22.12 | 968.11 | 14:28 | 428 | 21.98 | 22.3 | 275.52 | 18:03 | 428 | 26.07 | 22.33 | 969.15 | 9:14 | 428 | 25.93 | 22.31 | 971.1
14:52 | 489 | 27.59 | 21.89 | 971.02 | 15:58 | 489 | 19.74 | 21.98 | 970.02 | 18:14 | 489 | 14.86 | 21.92 | 970.08 | 15:57 | 489 | 22.19 | 21.91 | 969.24 | 16:38 | 489 | 25.03 | 22.1 | 968.11 | 14:29 | 489 | 22.01 | 22.3 | 275.47 | 18:04 | 489 | 25.95 | 22.34 | 969.18 | 9:15 | 489 | 25.8 | 22.3 | 971.11
14:53 | 550 | 27.57 | 21.9 | 971 | 15:59 | 550 | 19.35 | 21.99 | 969.98 | 18:15 | 550 | 14.09 | 21.93 | 970.06 | 15:58 | 550 | 22.18 | 21.9 | 969.22 | 16:39 | 550 | 25.06 | 22.13 | 968.13 | 14:30 | 550 | 21.96 | 22.3 | 275.45 | 18:05 | 550 | 25.82 | 22.35 | 969.18 | 9:16 | 550 | 25.89 | 22.3 | 971.12
14:54 | 611 | 27.56 | 21.89 | 971.01 | 16:00 | 611 | 18.93 | 21.99 | 970.01 | 18:16 | 611 | 13.4 | 21.93 | 970.04 | 15:59 | 611 | 22.2 | 21.91 | 969.22 | 16:40 | 611 | 25.05 | 22.1 | 968.12 | 14:31 | 611 | 21.95 | 22.3 | 275.36 | 18:06 | 611 | 25.8 | 22.36 | 969.17 | 9:17 | 611 | 25.8 | 22.31 | 971.13
14:55 | 672 | 27.56 | 21.88 | 970.98 | 16:01 | 672 | 18.49 | 21.98 | 970 | 18:17 | 672 | 12.69 | 21.92 | 970.07 | 16:00 | 672 | 22.14 | 21.88 | 969.23 | 16:41 | 672 | 25.02 | 22.1 | 968.08 | 14:32 | 672 | 21.9 | 22.29 | 275.36 | 18:07 | 672 | 25.66 | 22.34 | 969.15 | 9:18 | 672 | 25.74 | 22.29 | 971.13
14:56 | 733 | 27.54 | 21.88 | 970.98 | 16:02 | 733 | 18.11 | 21.96 | 969.99 | 18:18 | 733 | 12.1 | 21.95 | 970.07 | 16:01 | 733 | 22.16 | 21.9 | 969.2 | 16:42 | 733 | 25.02 | 22.09 | 968.15 | 14:33 | 733 | 21.9 | 22.3 | 275.39 | 18:08 | 733 | 25.53 | 22.33 | 969.19 | 9:19 | 733 | 25.72 | 22.3 | 971.11
14:57 | 794 | 27.6 | 21.91 | 970.96 | 16:03 | 794 | 17.72 | 21.96 | 970 | 18:19 | 794 | 11.4 | 21.93 | 970.13 | 16:02 | 794 | 22.09 | 21.89 | 969.24 | 16:43 | 794 | 24.96 | 22.1 | 968.11 | 14:34 | 794 | 21.97 | 22.3 | 275.36 | 18:09 | 794 | 25.44 | 22.32 | 969.2 | 9:20 | 794 | 25.71 | 22.3 | 971.15
14:58 | 855 | 27.61 | 21.9 | 970.96 | 16:04 | 855 | 17.41 | 21.97 | 969.93 | 18:20 | 855 | 10.76 | 21.94 | 970.1 | 16:03 | 855 | 22.07 | 21.88 | 969.22 | 16:44 | 855 | 24.96 | 22.09 | 968.1 | 14:35 | 855 | 21.9 | 22.31 | 275.37 | 18:10 | 855 | 25.38 | 22.33 | 969.17 | 9:21 | 855 | 25.71 | 22.31 | 971.1
14:59 | 916 | 27.59 | 21.89 | 970.93 | 16:05 | 916 | 17.02 | 21.99 | 969.91 | 18:21 | 916 | 10.19 | 21.94 | 970.1 | 16:04 | 916 | 22.09 | 21.89 | 969.24 | 16:45 | 916 | 24.96 | 22.09 | 968.13 | 14:36 | 916 | 21.94 | 22.3 | 275.36 | 18:11 | 916 | 25.32 | 22.34 | 969.15 | 9:22 | 916 | 25.7 | 22.29 | 971.08
15:00 | 977 | 27.65 | 21.9 | 970.88 | 16:06 | 977 | 16.68 | 21.98 | 969.95 | 18:22 | 977 | 9.59 | 21.94 | 970.1 | 16:05 | 977 | 22.03 | 21.89 | 969.22 | 16:46 | 977 | 24.94 | 22.11 | 968.15 | 14:37 | 977 | 21.92 | 22.3 | 275.35 | 18:12 | 977 | 25.2 | 22.33 | 969.17 | 9:23 | 977 | 25.66 | 22.29 | 971.1
15:01 | 1038 | 27.64 | 21.9 | 970.81 | 16:07 | 1038 | 16.32 | 21.98 | 969.91 | 18:23 | 1038 | 9.1 | 21.95 | 970.11 | 16:06 | 1038 | 22 | 21.89 | 969.21 | 16:47 | 1038 | 24.96 | 22.1 | 968.09 | 14:38 | 1038 | 21.93 | 22.3 | 275.42 | 18:13 | 1038 | 25.16 | 22.35 | 969.16 | 9:24 | 1038 | 25.68 | 22.32 | 971.06
15:02 | 1099 | 27.64 | 21.9 | 970.77 | 16:08 | 1099 | 15.97 | 21.98 | 969.92 | 18:24 | 1099 | 8.6 | 21.97 | 970.1 | 16:07 | 1099 | 21.98 | 21.88 | 969.22 | 16:48 | 1099 | 24.92 | 22.1 | 968.11 | 14:39 | 1099 | 22.03 | 22.31 | 275.39 | 18:14 | 1099 | 25.07 | 22.36 | 969.13 | 9:25 | 1099 | 25.59 | 22.31 | 971.08
15:03 | 1160 | 27.71 | 21.9 | 970.7 | 16:09 | 1160 | 15.6 | 21.97 | 969.86 | 18:25 | 1160 | 8.1 | 21.97 | 970.12 | 16:08 | 1160 | 21.95 | 21.89 | 969.2 | 16:49 | 1160 | 24.89 | 22.1 | 968.1 | 14:40 | 1160 | 21.94 | 22.31 | 275.4 | 18:15 | 1160 | 25.05 | 22.36 | 969.16 | 9:26 | 1160 | 25.61 | 22.31 | 971.07
15:04 | 1221 | 27.67 | 21.88 | 970.65 | 16:10 | 1221 | 15.28 | 21.98 | 969.88 | 18:26 | 1221 | 7.62 | 21.97 | 970.13 | 16:09 | 1221 | 21.9 | 21.89 | 969.2 | 16:50 | 1221 | 24.88 | 22.08 | 968.12 | 14:41 | 1221 | 21.93 | 22.3 | 275.4 | 18:16 | 1221 | 24.91 | 22.36 | 969.15 | 9:27 | 1221 | 25.48 | 22.3 | 971.06
15:05 | 1282 | 27.7 | 21.89 | 970.62 | 16:11 | 1282 | 15.01 | 21.98 | 969.85 | 18:27 | 1282 | 7.15 | 21.95 | 970.12 | 16:10 | 1282 | 21.96 | 21.9 | 969.22 | 16:51 | 1282 | 24.88 | 22.1 | 968.09 | 14:42 | 1282 | 21.96 | 22.29 | 275.42 | 18:17 | 1282 | 24.88 | 22.35 | 969.14 | 9:28 | 1282 | 25.5 | 22.3 | 971.07
15:06 | 1343 | 27.74 | 21.89 | 970.53 | 16:12 | 1343 | 14.59 | 21.99 | 969.81 | 18:28 | 1343 | 6.79 | 21.94 | 970.12 | 16:11 | 1343 | 21.93 | 21.9 | 969.22 | 16:52 | 1343 | 24.82 | 22.1 | 968.09 | 14:43 | 1343 | 21.87 | 22.31 | 275.46 | 18:18 | 1343 | 24.78 | 22.34 | 969.16 | 9:29 | 1343 | 25.47 | 22.3 | 971.07
15:07 | 1404 | 27.73 | 21.89 | 970.5 | 16:13 | 1404 | 14.39 | 21.99 | 969.79 | 18:29 | 1404 | 6.31 | 22.02 | 970.12 | 16:12 | 1404 | 21.9 | 21.91 | 969.19 | 16:53 | 1404 | 24.81 | 22.1 | 968.07 | 14:44 | 1404 | 22 | 22.3 | 275.49 | 18:19 | 1404 | 24.76 | 22.34 | 969.18 | 9:30 | 1404 | 25.46 | 22.3 | 971.07
15:08 | 1465 | 27.7 | 21.89 | 970.48 | 16:14 | 1465 | 13.99 | 22 | 969.79 | 18:30 | 1465 | 5.97 | 22.12 | 970.07 | 16:13 | 1465 | 21.97 | 21.9 | 969.2 | 16:54 | 1465 | 24.82 | 22.11 | 968.06 | 14:45 | 1465 | 22.03 | 22.32 | 275.54 | 18:20 | 1465 | 24.66 | 22.33 | 969.16 | 9:31 | 1465 | 25.42 | 22.3 | 971.07
15:09 | 1526 | 27.81 | 21.9 | 970.5 | 16:15 | 1526 | 13.69 | 21.99 | 969.76 | 18:31 | 1526 | 5.55 | 22.07 | 970.12 | 16:14 | 1526 | 21.91 | 21.9 | 969.21 | 16:55 | 1526 | 24.83 | 22.06 | 968.07 | 14:46 | 1526 | 22.04 | 22.31 | 275.52 | 18:21 | 1526 | 24.62 | 22.33 | 969.15 | 9:32 | 1526 | 25.44 | 22.31 | 971.05
15:10 | 1587 | 27.91 | 21.91 | 970.5 | 16:16 | 1587 | 13.4 | 21.99 | 969.71 | 18:32 | 1587 | 5.21 | 21.99 | 970.12 | 16:15 | 1587 | 21.86 | 21.91 | 969.2 | 16:56 | 1587 | 24.81 | 22.07 | 968.09 | 14:47 | 1587 | 22.03 | 22.3 | 275.56 | 18:22 | 1587 | 24.51 | 22.32 | 969.17 | 9:33 | 1587 | 25.38 | 22.31 | 971.04
15:11 | 1648 | 27.98 | 21.92 | 970.48 | 16:17 | 1648 | 13.1 | 22.01 | 969.69 | 18:33 | 1648 | 4.88 | 21.95 | 970.11 | 16:16 | 1648 | 21.84 | 21.9 | 969.19 | 16:57 | 1648 | 24.78 | 22.08 | 968.08 | 14:48 | 1648 | 22.05 | 22.31 | 275.61 | 18:23 | 1648 | 24.48 | 22.34 | 969.18 | 9:34 | 1648 | 25.4 | 22.31 | 971.02
15:12 | 1709 | 28.02 | 21.91 | 970.44 | 16:18 | 1709 | 12.83 | 21.99 | 969.67 | 18:34 | 1709 | 4.58 | 21.93 | 970.14 | 16:17 | 1709 | 21.84 | 21.9 | 969.15 | 16:58 | 1709 | 24.72 | 22.05 | 968.09 | 14:49 | 1709 | 22.03 | 22.31 | 275.62 | 18:24 | 1709 | 24.42 | 22.32 | 969.21 | 9:35 | 1709 | 25.36 | 22.3 | 971.02
15:13 | 1770 | 28.08 | 21.9 | 970.45 | 16:19 | 1770 | 12.51 | 22 | 969.64 | 18:35 | 1770 | 4.24 | 21.9 | 970.13 | 16:18 | 1770 | 21.81 | 21.89 | 969.2 | 16:59 | 1770 | 24.72 | 22.04 | 968.08 | 14:50 | 1770 | 22.04 | 22.3 | 275.66 | 18:25 | 1770 | 24.4 | 22.33 | 969.2 | 9:36 | 1770 | 25.36 | 22.32 | 971.01
15:14 | 1831 | 28.13 | 21.93 | 970.37 | 16:20 | 1831 | 12.24 | 21.99 | 969.62 | 18:36 | 1831 | 4.05 | 21.9 | 970.13 | 16:19 | 1831 | 21.82 | 21.88 | 969.17 | 17:00 | 1831 | 24.82 | 22.01 | 968.08 | 14:51 | 1831 | 22.09 | 22.29 | 275.68 | 18:26 | 1831 | 24.29 | 22.33 | 969.15 | 9:37 | 1831 | 25.27 | 22.31 | 970.99
15:15 | 1892 | 28.24 | 21.93 | 970.34 | 16:21 | 1892 | 11.99 | 21.99 | 969.64 | 18:37 | 1892 | 3.81 | 21.89 | 970.12 | 16:20 | 1892 | 21.8 | 21.9 | 969.17 | 17:01 | 1892 | 24.76 | 21.99 | 968.09 | 14:52 | 1892 | 22.06 | 22.29 | 275.7 | 18:27 | 1892 | 24.2 | 22.32 | 969.15 | 9:38 | 1892 | 25.29 | 22.3 | 971.01
15:16 | 1953 | 28.38 | 21.92 | 970.31 | 16:22 | 1953 | 11.75 | 21.99 | 969.6 | 18:38 | 1953 | 3.51 | 21.9 | 970.13 | 16:21 | 1953 | 21.81 | 21.9 | 969.17 | 17:02 | 1953 | 24.71 | 21.99 | 968.12 | 14:53 | 1953 | 22.07 | 22.3 | 275.69 | 18:28 | 1953 | 24.22 | 22.32 | 969.18 | 9:39 | 1953 | 25.18 | 22.31 | 971.03
15:17 | 2014 | 28.39 | 21.95 | 970.33 | 16:23 | 2014 | 11.4 | 22 | 969.59 | 18:39 | 2014 | 3.34 | 21.89 | 970.12 | 16:22 | 2014 | 21.79 | 21.89 | 969.18 | 17:03 | 2014 | 24.64 | 21.99 | 968.1 | 14:54 | 2014 | 22.07 | 22.29 | 275.7 | 18:29 | 2014 | 24.11 | 22.32 | 969.15 | 9:40 | 2014 | 25.19 | 22.31 | 971
15:18 | 2075 | 28.34 | 21.95 | 970.29 | 16:24 | 2075 | 11.14 | 21.99 | 969.55 | 18:40 | 2075 | 3.04 | 21.88 | 970.15 | 16:23 | 2075 | 21.81 | 21.89 | 969.16 | 17:04 | 2075 | 24.7 | 21.99 | 968.09 | 14:55 | 2075 | 22.07 | 22.29 | 275.72 | 18:30 | 2075 | 24.06 | 22.33 | 969.14 | 9:41 | 2075 | 25.18 | 22.3 | 971
15:19 | 2136 | 28.31 | 21.94 | 970.26 | 16:25 | 2136 | 10.93 | 22 | 969.51 | 18:41 | 2136 | 2.84 | 21.92 | 970.13 | 16:24 | 2136 | 21.77 | 21.88 | 969.16 | 17:05 | 2136 | 24.66 | 21.99 | 968.07 | 14:56 | 2136 | 22.07 | 22.3 | 275.77 | 18:31 | 2136 | 24 | 22.33 | 969.17 | 9:42 | 2136 | 25.13 | 22.31 | 971.01
15:20 | 2197 | 28.33 | 21.94 | 970.3 | 16:26 | 2197 | 10.69 | 22.01 | 969.5 | 18:42 | 2197 | 2.72 | 21.91 | 970.11 | 16:25 | 2197 | 21.73 | 21.9 | 969.16 | 17:06 | 2197 | 24.69 | 21.99 | 968.08 | 14:57 | 2197 | 22.1 | 22.3 | 275.84 | 18:32 | 2197 | 24 | 22.34 | 969.18 | 9:43 | 2197 | 25.16 | 22.31 | 970.99
15:21 | 2258 | 28.35 | 21.94 | 970.24 | 16:27 | 2258 | 10.5 | 22 | 969.56 | 18:43 | 2258 | 2.5 | 21.9 | 970.12 | 16:26 | 2258 | 21.71 | 21.89 | 969.15 | 17:07 | 2258 | 24.65 | 21.97 | 968.11 | 14:58 | 2258 | 22.18 | 22.3 | 275.99 | 18:33 | 2258 | 23.96 | 22.36 | 969.19 | 9:44 | 2258 | 25.19 | 22.31 | 970.99
15:22 | 2319 | 28.38 | 21.94 | 970.24 | 16:28 | 2319 | 10.23 | 22 | 969.53 | 18:44 | 2319 | 2.38 | 21.9 | 970.11 | 16:27 | 2319 | 21.73 | 21.89 | 969.19 | 17:08 | 2319 | 24.63 | 21.97 | 968.09 | 14:59 | 2319 | 22.07 | 22.3 | 276.02 | 18:34 | 2319 | 23.84 | 22.33 | 969.17 | 9:45 | 2319 | 25.06 | 22.3 | 970.99
15:23 | 2380 | 28.34 | 21.95 | 970.24 | 16:29 | 2380 | 10 | 22.01 | 969.46 | 18:45 | 2380 | 2.12 | 21.89 | 970.11 | 16:28 | 2380 | 21.72 | 21.9 | 969.16 | 17:09 | 2380 | 24.59 | 22 | 968.04 | 15:00 | 2380 | 22.16 | 22.29 | 276.05 | 18:35 | 2380 | 23.82 | 22.34 | 969.2 | 9:46 | 2380 | 25.12 | 22.31 | 971.01
15:24 | 2441 | 28.4 | 21.95 | 970.22 | 16:30 | 2441 | 9.77 | 21.99 | 969.48 | 18:46 | 2441 | 2.03 | 21.9 | 970.13 | 16:29 | 2441 | 21.65 | 21.9 | 969.16 | 17:10 | 2441 | 24.56 | 22 | 968.05 | 15:01 | 2441 | 22.16 | 22.32 | 276.13 | 18:36 | 2441 | 23.85 | 22.32 | 969.17 | 9:47 | 2441 | 25.09 | 22.31 | 970.98
15:25 | 2502 | 28.43 | 21.95 | 970.22 | 16:31 | 2502 | 9.49 | 22 | 969.48 | 18:47 | 2502 | 1.84 | 21.9 | 970.11 | 16:30 | 2502 | 21.69 | 21.9 | 969.16 | 17:11 | 2502 | 24.6 | 21.99 | 968.04 | 15:02 | 2502 | 22.22 | 22.32 | 276.15 | 18:37 | 2502 | 23.77 | 22.32 | 969.21 | 9:48 | 2502 | 25.05 | 22.32 | 970.96
15:26 | 2563 | 28.44 | 21.97 | 970.22 | 16:32 | 2563 | 9.33 | 22.01 | 969.47 | 18:48 | 2563 | 1.76 | 21.91 | 970.1 | 16:31 | 2563 | 21.66 | 21.88 | 969.18 | 17:12 | 2563 | 24.56 | 22 | 968.03 | 15:03 | 2563 | 22.21 | 22.31 | 276.21 | 18:38 | 2563 | 23.74 | 22.33 | 969.21 | 9:49 | 2563 | 25.07 | 22.31 | 970.95
15:27 | 2624 | 28.45 | 21.95 | 970.21 | 16:33 | 2624 | 9.07 | 22.02 | 969.45 | 18:49 | 2624 | 1.67 | 21.96 | 970.12 | 16:32 | 2624 | 21.7 | 21.9 | 969.18 | 17:13 | 2624 | 24.58 | 22.01 | 968.05 | 15:04 | 2624 | 22.13 | 22.3 | 276.13 | 18:39 | 2624 | 23.69 | 22.33 | 969.21 | 9:50 | 2624 | 25.06 | 22.3 | 970.93
15:28 | 2685 | 28.4 | 21.96 | 970.22 | 16:34 | 2685 | 8.81 | 22.01 | 969.45 | 18:50 | 2685 | 1.51 | 21.99 | 970.13 | 16:33 | 2685 | 21.65 | 21.88 | 969.16 | 17:14 | 2685 | 24.49 | 21.99 | 968.04 | 15:05 | 2685 | 22.14 | 22.31 | 276.22 | 18:40 | 2685 | 23.67 | 22.32 | 969.2 | 9:51 | 2685 | 25.01 | 22.31 | 970.96
15:29 | 2746 | 28.43 | 21.96 | 970.21 | 16:35 | 2746 | 8.69 | 22.02 | 969.42 | 18:51 | 2746 | 1.36 | 22.07 | 970.1 | 16:34 | 2746 | 21.6 | 21.89 | 969.17 | 17:15 | 2746 | 24.52 | 22.01 | 968.04 | 15:06 | 2746 | 22.21 | 22.31 | 276.26 | 18:41 | 2746 | 23.62 | 22.31 | 969.21 | 9:52 | 2746 | 24.98 | 22.31 | 970.91
15:30 | 2807 | 28.41 | 21.96 | 970.22 | 16:36 | 2807 | 8.42 | 22 | 969.4 | 18:52 | 2807 | 1.28 | 22.14 | 970.08 | 16:35 | 2807 | 21.62 | 21.9 | 969.14 | 17:16 | 2807 | 24.47 | 22.03 | 968.05 | 15:07 | 2807 | 22.21 | 22.29 | 276.28 | 18:42 | 2807 | 23.61 | 22.32 | 969.2 | 9:53 | 2807 | 24.95 | 22.31 | 970.95
15:31 | 2868 | 28.48 | 21.96 | 970.2 | 16:37 | 2868 | 8.26 | 22 | 969.4 | 18:53 | 2868 | 1.25 | 22.17 | 970.11 | 16:36 | 2868 | 21.6 | 21.9 | 969.19 | 17:17 | 2868 | 24.44 | 22 | 968.05 | 15:08 | 2868 | 22.21 | 22.29 | 276.27 | 18:43 | 2868 | 23.5 | 22.32 | 969.18 | 9:54 | 2868 | 24.95 | 22.31 | 970.94
15:32 | 2929 | 28.49 | 21.95 | 970.17 | 16:38 | 2929 | 8.13 | 22.02 | 969.38 | 18:54 | 2929 | 1.04 | 22.22 | 970.1 | 16:37 | 2929 | 21.61 | 21.89 | 969.16 | 17:18 | 2929 | 24.47 | 21.97 | 968.07 | 15:09 | 2929 | 22.23 | 22.31 | 276.33 | 18:44 | 2929 | 23.52 | 22.33 | 969.2 | 9:55 | 2929 | 24.96 | 22.32 | 970.94
15:33 | 2990 | 28.56 | 21.95 | 970.17 | 16:39 | 2990 | 7.87 | 22.03 | 969.34 | 18:55 | 2990 | 1.02 | 22.23 | 970.11 | 16:38 | 2990 | 21.59 | 21.9 | 969.18 | 17:19 | 2990 | 24.5 | 21.99 | 968.07 | 15:10 | 2990 | 22.31 | 22.3 | 276.37 | 18:45 | 2990 | 23.38 | 22.32 | 969.2 | 9:56 | 2990 | 24.88 | 22.32 | 970.93
15:34 | 3051 | 28.49 | 21.96 | 970.16 | 16:40 | 3051 | 7.67 | 22.03 | 969.36 | 18:56 | 3051 | 0.96 | 22.23 | 970.11 | 16:39 | 3051 | 21.54 | 21.89 | 969.15 | 17:20 | 3051 | 24.39 | 21.98 | 968.07 | 15:11 | 3051 | 22.28 | 22.29 | 276.37 | 18:46 | 3051 | 23.36 | 22.32 | 969.21 | 9:57 | 3051 | 24.9 | 22.32 | 970.95
15:35 | 3112 | 28.57 | 21.96 | 970.15 | 16:41 | 3112 | 7.47 | 22.01 | 969.33 | 18:57 | 3112 | 0.88 | 22.24 | 970.09 | 16:40 | 3112 | 21.55 | 21.89 | 969.15 | 17:21 | 3112 | 24.41 | 21.99 | 968.07 | 15:12 | 3112 | 22.19 | 22.31 | 276.43 | 18:47 | 3112 | 23.33 | 22.31 | 969.19 | 9:58 | 3112 | 24.85 | 22.31 | 970.94
15:36 | 3173 | 28.58 | 21.98 | 970.14 | 16:42 | 3173 | 7.31 | 22.03 | 969.3 | 18:58 | 3173 | 0.72 | 22.24 | 970.1 | 16:41 | 3173 | 21.54 | 21.87 | 969.16 | 17:22 | 3173 | 24.41 | 22.01 | 968.09 | 15:13 | 3173 | 22.4 | 22.31 | 276.44 | 18:48 | 3173 | 23.33 | 22.32 | 969.2 | 9:59 | 3173 | 24.83 | 22.33 | 970.97
15:37 | 3234 | 28.58 | 21.96 | 970.14 | 16:43 | 3234 | 7.12 | 22.03 | 969.3 | 18:59 | 3234 | 0.75 | 22.23 | 970.12 | 16:42 | 3234 | 21.48 | 21.88 | 969.13 | 17:23 | 3234 | 24.42 | 21.99 | 968.12 | 15:14 | 3234 | 22.32 | 22.31 | 276.47 | 18:49 | 3234 | 23.27 | 22.34 | 969.2 | 10:00 | 3234 | 24.84 | 22.35 | 970.98
15:38 | 3295 | 28.63 | 21.97 | 970.14 | 16:44 | 3295 | 6.96 | 22.02 | 969.26 | 19:00 | 3295 | 0.67 | 22.23 | 970.09 | 16:43 | 3295 | 21.49 | 21.89 | 969.15 | 17:24 | 3295 | 24.39 | 22 | 968.16 | 15:15 | 3295 | 22.33 | 22.32 | 276.49 | 18:50 | 3295 | 23.25 | 22.36 | 969.21 | 10:01 | 3295 | 24.79 | 22.34 | 970.98
15:39 | 3356 | 28.67 | 21.96 | 970.11 | 16:45 | 3356 | 6.81 | 22.02 | 969.27 | 19:01 | 3356 | 0.54 | 22.24 | 970.11 | 16:44 | 3356 | 21.48 | 21.89 | 969.13 | 17:25 | 3356 | 24.42 | 22.02 | 968.14 | 15:16 | 3356 | 22.32 | 22.31 | 276.53 | 18:51 | 3356 | 23.15 | 22.4 | 969.17 | 10:02 | 3356 | 24.76 | 22.34 | 970.97
15:40 | 3417 | 28.66 | 21.96 | 970.15 | 16:46 | 3417 | 6.62 | 22.04 | 969.28 | 19:02 | 3417 | 0.52 | 22.25 | 970.11 | 16:45 | 3417 | 21.48 | 21.89 | 969.13 | 17:26 | 3417 | 24.44 | 22.05 | 968.15 | 15:17 | 3417 | 22.26 | 22.32 | 276.56 | 18:52 | 3417 | 23.1 | 22.47 | 969.19 | 10:03 | 3417 | 24.84 | 22.34 | 970.97
15:41 | 3478 | 28.68 | 21.98 | 970.15 | 16:47 | 3478 | 6.49 | 22.04 | 969.27 | 19:03 | 3478 | 0.45 | 22.25 | 970.12 | 16:46 | 3478 | 21.49 | 21.89 | 969.15 | 17:27 | 3478 | 24.42 | 22.07 | 968.15 | 15:18 | 3478 | 22.43 | 22.31 | 276.7 | 18:53 | 3478 | 23.12 | 22.5 | 969.18 | 10:04 | 3478 | 24.74 | 22.34 | 970.89
15:42 | 3539 | 28.71 | 21.99 | 970.12 | 16:48 | 3539 | 6.34 | 22.04 | 969.24 | 19:04 | 3539 | 0.42 | 22.25 | 970.11 | 16:47 | 3539 | 21.47 | 21.88 | 969.1 | 17:28 | 3539 | 24.43 | 22.07 | 968.16 | 15:19 | 3539 | 22.38 | 22.33 | 276.76 | 18:54 | 3539 | 23.09 | 22.54 | 969.15 | 10:05 | 3539 | 24.76 | 22.34 | 970.91
15:43 | 3600 | 28.76 | 21.98 | 970.1 | 16:49 | 3600 | 6.13 | 22.05 | 969.24 | 19:05 | 3600 | 0.41 | 22.25 | 970.14 | 16:48 | 3600 | 21.42 | 21.88 | 969.15 | 17:29 | 3600 | 24.38 | 22.05 | 968.17 | 15:20 | 3600 | 22.36 | 22.34 | 276.8 | 18:55 | 3600 | 23.07 | 22.55 | 969.16 | 10:06 | 3600 | 24.7 | 22.32 | 970.89
[5 empty] | 16:50 | 3661 | 6.02 | 22.05 | 969.23 | [5 empty] | 16:49 | 3661 | 21.45 | 21.88 | 969.14 | 17:30 | 3661 | 24.36 | 22.05 | 968.21 | 15:21 | 3661 | 22.4 | 22.31 | 276.93 | 18:56 | 3661 | 23.04 | 22.55 | 969.18 | 10:07 | 3661 | 24.71 | 22.32 | 970.93
[5 empty] | 16:51 | 3722 | 5.84 | 22.06 | 969.21 | [10 empty] | 17:31 | 3722 | 24.37 | 22.07 | 968.18 | 15:22 | 3722 | 22.62 | 22.31 | 277
[5 empty] | 16:52 | 3783 | 5.72 | 22.08 | 969.21 | [10 empty] | 17:32 | 3783 | 24.42 | 22.07 | 968.19 | 15:23 | 3783 | 22.42 | 22.32 | 277.04
[5 empty] | 16:53 | 3844 | 5.57 | 22.05 | 969.2 | [10 empty] | 17:33 | 3844 | 24.3 | 22.06 | 968.18 | 15:24 | 3844 | 22.36 | 22.32 | 276.99
[5 empty] | 16:54 | 3905 | 5.42 | 22.06 | 969.2 | [10 empty] | 17:34 | 3905 | 24.3 | 22.07 | 968.18
[5 empty] | 16:55 | 3966 | 5.24 | 22.06 | 969.22 | [10 empty] | 17:35 | 3966 | 24.33 | 22.05 | 968.18
[5 empty] | 16:56 | 4027 | 5.15 | 22.06 | 969.15 | [10 empty] | 17:36 | 4027 | 24.33 | 22.07 | 968.16
[5 empty] | 16:57 | 4088 | 5.02 | 22.07 | 969.18
[5 empty] | 16:58 | 4149 | 4.81 | 22.08 | 969.18
[5 empty] | 16:59 | 4210 | 4.77 | 22.06 | 969.16
[5 empty] | 17:00 | 4271 | 4.58 | 22.07 | 969.18
[5 empty] | 17:01 | 4332 | 4.45 | 22.08 | 969.15
[5 empty] | 17:02 | 4393 | 4.36 | 22.09 | 969.09
[5 empty] | 17:03 | 4454 | 4.24 | 22.08 | 969.14
[5 empty] | 17:04 | 4515 | 4.14 | 22.07 | 969.12
[5 empty] | 17:05 | 4576 | 4.02 | 22.05 | 969.1
[5 empty] | 17:06 | 4637 | 3.92 | 22.05 | 969.11
[5 empty] | 17:07 | 4698 | 3.8 | 22.06 | 969.1

## QC3_Long_2
Test Date: | 2017-06-01 |  |  |  | Test Date: | 2017-06-01
File: | QC3_GE11-X-L-CERN-0002_Calibration_2017_06_01.xls |  |  |  | File: | QC3_GE11-X-L-CERN-0002_2017_06_01.xls
Elog: | http://cmsonline.cern.ch/cms-elog/988203 |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/988203
Description: | Calibration (Stand2) |  |  |  | Description: | After fixing the bending on the RO
Max P (mbar) | 29.3 |  |  |  | Max P (mbar) | 20.38
Min P (mbar) | 28.72 |  |  |  | Min P (mbar) | 19.39
Delta P (mbar) | 0.58 |  |  |  | Delta P (mbar) | 0.99
Min Time | 10:04 |  |  |  | Min Time | 15:18
Max Time | 11:01 |  |  |  | Max Time | 16:21
Delta T | 0:57 |  |  |  | Delta T | 1:03
Time (hr) | 0.95028 |  |  |  | Time (hr) | 1.0503
Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar)
10:04 | 1 | 29 | 21.81 | 972.29 | 15:18 | 1 | 20.38 | 21.81 | 972.29
10:05 | 62 | 28.72 | 21.75 | 971.91 | 15:19 | 62 | 20.26 | 21.67 | 972.27
10:06 | 123 | 28.75 | 21.73 | 971.89 | 15:20 | 123 | 20.13 | 21.68 | 972.26
10:07 | 184 | 28.79 | 21.75 | 971.91 | 15:21 | 184 | 20.14 | 21.66 | 972.26
10:08 | 245 | 28.81 | 21.74 | 971.87 | 15:22 | 245 | 20.11 | 21.67 | 972.24
10:09 | 306 | 28.79 | 21.75 | 971.88 | 15:23 | 306 | 20.11 | 21.65 | 972.23
10:10 | 367 | 28.85 | 21.73 | 971.92 | 15:24 | 367 | 20.04 | 21.64 | 972.22
10:11 | 428 | 28.82 | 21.75 | 971.84 | 15:25 | 428 | 20.05 | 21.63 | 972.22
10:12 | 489 | 28.8 | 21.74 | 971.84 | 15:26 | 489 | 20.03 | 21.64 | 972.21
10:13 | 550 | 28.85 | 21.74 | 971.83 | 15:27 | 550 | 20.07 | 21.64 | 972.23
10:14 | 611 | 28.83 | 21.74 | 971.88 | 15:28 | 611 | 20.05 | 21.62 | 972.21
10:15 | 672 | 28.93 | 21.75 | 971.88 | 15:29 | 672 | 20.07 | 21.63 | 972.2
10:16 | 733 | 28.87 | 21.75 | 971.87 | 15:30 | 733 | 20.01 | 21.62 | 972.21
10:17 | 794 | 28.84 | 21.76 | 971.92 | 15:31 | 794 | 20.07 | 21.61 | 972.22
10:18 | 855 | 28.84 | 21.76 | 971.92 | 15:32 | 855 | 19.98 | 21.62 | 972.19
10:19 | 916 | 28.89 | 21.76 | 971.92 | 15:33 | 916 | 19.98 | 21.59 | 972.22
10:20 | 977 | 28.92 | 21.75 | 971.92 | 15:34 | 977 | 19.93 | 21.6 | 972.17
10:21 | 1038 | 28.89 | 21.75 | 971.95 | 15:35 | 1038 | 19.96 | 21.61 | 972.18
10:22 | 1099 | 28.89 | 21.74 | 971.95 | 15:36 | 1099 | 19.84 | 21.6 | 972.17
10:23 | 1160 | 28.9 | 21.74 | 971.91 | 15:37 | 1160 | 19.79 | 21.61 | 972.13
10:24 | 1221 | 28.89 | 21.74 | 971.91 | 15:38 | 1221 | 19.81 | 21.59 | 972.16
10:25 | 1282 | 28.89 | 21.74 | 971.92 | 15:39 | 1282 | 19.81 | 21.6 | 972.12
10:26 | 1343 | 28.87 | 21.75 | 971.91 | 15:40 | 1343 | 19.83 | 21.59 | 972.12
10:27 | 1404 | 28.92 | 21.74 | 971.91 | 15:41 | 1404 | 19.76 | 21.59 | 972.09
10:28 | 1465 | 29 | 21.74 | 971.89 | 15:42 | 1465 | 19.84 | 21.61 | 972.06
10:29 | 1526 | 28.9 | 21.75 | 971.88 | 15:43 | 1526 | 19.76 | 21.6 | 972.06
10:30 | 1587 | 28.9 | 21.74 | 971.9 | 15:44 | 1587 | 19.74 | 21.58 | 972.06
10:31 | 1648 | 28.88 | 21.74 | 971.89 | 15:45 | 1648 | 19.74 | 21.6 | 972.04
10:32 | 1709 | 28.92 | 21.73 | 971.91 | 15:46 | 1709 | 19.73 | 21.59 | 971.97
10:33 | 1770 | 28.94 | 21.74 | 971.94 | 15:47 | 1770 | 19.68 | 21.59 | 971.97
10:34 | 1831 | 28.96 | 21.75 | 971.89 | 15:48 | 1831 | 19.66 | 21.59 | 971.95
10:35 | 1892 | 28.94 | 21.75 | 971.91 | 15:49 | 1892 | 19.7 | 21.58 | 971.94
10:36 | 1953 | 28.95 | 21.75 | 971.86 | 15:50 | 1953 | 19.64 | 21.58 | 971.89
10:37 | 2014 | 28.9 | 21.74 | 971.85 | 15:51 | 2014 | 19.62 | 21.59 | 971.91
10:38 | 2075 | 28.92 | 21.74 | 971.83 | 15:52 | 2075 | 19.6 | 21.58 | 971.85
10:39 | 2136 | 28.94 | 21.74 | 971.84 | 15:53 | 2136 | 19.56 | 21.58 | 971.79
10:40 | 2197 | 28.88 | 21.74 | 971.82 | 15:54 | 2197 | 19.56 | 21.6 | 971.81
10:41 | 2258 | 28.96 | 21.74 | 971.81 | 15:55 | 2258 | 19.53 | 21.59 | 971.77
10:42 | 2319 | 29.12 | 21.73 | 971.81 | 15:56 | 2319 | 19.49 | 21.61 | 971.74
10:43 | 2380 | 29.23 | 21.74 | 971.74 | 15:57 | 2380 | 19.53 | 21.6 | 971.7
10:44 | 2441 | 29.3 | 21.73 | 971.69 | 15:58 | 2441 | 19.45 | 21.56 | 971.68
10:45 | 2502 | 29.23 | 21.74 | 971.7 | 15:59 | 2502 | 19.48 | 21.58 | 971.63
10:46 | 2563 | 29.28 | 21.75 | 971.67 | 16:00 | 2563 | 19.45 | 21.55 | 971.63
10:47 | 2624 | 29.3 | 21.75 | 971.61 | 16:01 | 2624 | 19.43 | 21.57 | 971.63
10:48 | 2685 | 29.28 | 21.74 | 971.64 | 16:02 | 2685 | 19.39 | 21.57 | 971.61
10:49 | 2746 | 29.3 | 21.75 | 971.62 | 16:03 | 2746 | 19.39 | 21.57 | 971.58
10:50 | 2807 | 29.23 | 21.75 | 971.64 | 16:04 | 2807 | 19.42 | 21.57 | 971.57
10:51 | 2868 | 29.22 | 21.75 | 971.62 | 16:05 | 2868 | 19.39 | 21.56 | 971.59
10:52 | 2929 | 29.29 | 21.74 | 971.64 | 16:06 | 2929 | 19.37 | 21.57 | 971.55
10:53 | 2990 | 29.3 | 21.74 | 971.61 | 16:07 | 2990 | 19.39 | 21.56 | 971.55
10:54 | 3051 | 29.27 | 21.74 | 971.61 | 16:08 | 3051 | 19.33 | 21.58 | 971.52
10:55 | 3112 | 29.28 | 21.74 | 971.57 | 16:09 | 3112 | 19.39 | 21.58 | 971.51
10:56 | 3173 | 29.28 | 21.74 | 971.56 | 16:10 | 3173 | 19.32 | 21.57 | 971.45
10:57 | 3234 | 29.31 | 21.74 | 971.6 | 16:11 | 3234 | 19.32 | 21.57 | 971.44
10:58 | 3295 | 29.3 | 21.74 | 971.55 | 16:12 | 3295 | 19.35 | 21.57 | 971.42
10:59 | 3356 | 29.31 | 21.74 | 971.53 | 16:13 | 3356 | 19.23 | 21.57 | 971.4
11:00 | 3417 | 29.33 | 21.73 | 971.55 | 16:14 | 3417 | 19.25 | 21.55 | 971.4
11:01 | 3478 | 29.29 | 21.72 | 971.54 | 16:15 | 3478 | 19.25 | 21.58 | 971.4
[5 empty] | 16:16 | 3539 | 19.29 | 21.55 | 971.41
[5 empty] | 16:17 | 3600 | 19.2 | 21.56 | 971.39
[5 empty] | 16:18 | 3661 | 19.15 | 21.56 | 971.41
[5 empty] | 16:19 | 3722 | 19.18 | 21.57 | 971.36
[5 empty] | 16:20 | 3783 | 19.1 | 21.56 | 971.36
[5 empty] | 16:21 | 3844 | 19.17 | 21.55 | 971.36

## QC3_Long_3
Test Date: | 2017-07-11 |  |  |  | Test Date: | 2017-07-12 |  |  |  | Test Date: | 2017-08-01
File: | QC3_GE11-X-L-CERN-0003_20170711.xlsx |  |  |  | File: | QC3_GE11-X-L-CERN-0003_20170712.xlsx |  |  |  | File: | QC3_GE11-X-L-CERN-0003_20170801.xlsx
Elog: |  |  |  |  | Elog: |  |  |  |  | Elog:
Description: | After applying 1.3Nm torque on the RO and DR screws |  |  |  | Description: | After applying Kapton tape on faulty screws on the DR |  |  |  | Description: | After fixing the readout board in the clean room
Max P (mbar) | 26.34 |  |  |  | Max P (mbar) | 24.48 |  |  |  | Max P (mbar) | 25.35
Min P (mbar) | 20.21 |  |  |  | Min P (mbar) | 22.03 |  |  |  | Min P (mbar) | 23.51
Delta P (mbar) | 6.13 |  |  |  | Delta P (mbar) | 2.45 |  |  |  | Delta P (mbar) | 1.84
Min Time | 19:37 |  |  |  | Min Time | 11:49 |  |  |  | Min Time | 9:03
Max Time | 20:36 |  |  |  | Max Time | 12:50 |  |  |  | Max Time | 10:04
Delta T | 0:59 |  |  |  | Delta T | 1:01 |  |  |  | Delta T | 1:01
Time (hr) | 0.98361 |  |  |  | Time (hr) | 1.0172 |  |  |  | Time (hr) | 1.0169
Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar) | Time | Seconds | Pressure (mBar) | Temperature (C) | Atm Pressure (mBar)
19:37 | 1 | 26.34 | 22.3 | 966.86 | 11:49 | 1 | 24.36 | 22.3 | 969.03 | 9:03 | 1 | 25.35 | 22.3 | 967.94
19:38 | 62 | 26.09 | 22.19 | 966.99 | 11:50 | 62 | 24.46 | 22.29 | 968.2 | 9:04 | 62 | 25.18 | 22.34 | 967.94
19:39 | 123 | 26 | 22.19 | 966.95 | 11:51 | 123 | 24.48 | 22.31 | 968.19 | 9:05 | 123 | 25.27 | 22.34 | 967.96
19:40 | 184 | 25.81 | 22.19 | 966.98 | 11:52 | 184 | 24.43 | 22.3 | 968.18 | 9:06 | 184 | 25.17 | 22.33 | 967.95
19:41 | 245 | 25.77 | 22.2 | 966.99 | 11:53 | 245 | 24.36 | 22.3 | 968.13 | 9:07 | 245 | 25.12 | 22.33 | 967.94
19:42 | 306 | 25.56 | 22.2 | 966.99 | 11:54 | 306 | 24.33 | 22.29 | 968.13 | 9:08 | 306 | 25.05 | 22.34 | 967.94
19:43 | 367 | 25.46 | 22.22 | 967 | 11:55 | 367 | 24.34 | 22.29 | 968.13 | 9:09 | 367 | 25.01 | 22.35 | 967.94
19:44 | 428 | 25.33 | 22.21 | 966.98 | 11:56 | 428 | 24.2 | 22.29 | 968.12 | 9:10 | 428 | 24.97 | 22.33 | 967.91
19:45 | 489 | 25.22 | 22.23 | 966.98 | 11:57 | 489 | 24.19 | 22.3 | 968.1 | 9:11 | 489 | 24.91 | 22.35 | 967.92
19:46 | 550 | 25.16 | 22.26 | 966.99 | 11:58 | 550 | 24.06 | 22.29 | 968.14 | 9:12 | 550 | 24.81 | 22.35 | 967.91
19:47 | 611 | 24.99 | 22.25 | 966.94 | 11:59 | 611 | 24.05 | 22.29 | 968.12 | 9:13 | 611 | 24.82 | 22.36 | 967.94
19:48 | 672 | 24.85 | 22.24 | 966.96 | 12:00 | 672 | 24.04 | 22.29 | 968.13 | 9:14 | 672 | 24.76 | 22.36 | 967.94
19:49 | 733 | 24.76 | 22.23 | 966.98 | 12:01 | 733 | 24.03 | 22.3 | 968.11 | 9:15 | 733 | 24.75 | 22.35 | 967.95
19:50 | 794 | 24.65 | 22.24 | 966.96 | 12:02 | 794 | 23.95 | 22.3 | 968.06 | 9:16 | 794 | 24.76 | 22.35 | 967.95
19:51 | 855 | 24.55 | 22.24 | 966.95 | 12:03 | 855 | 23.91 | 22.29 | 968.09 | 9:17 | 855 | 24.65 | 22.37 | 967.92
19:52 | 916 | 24.47 | 22.24 | 967.01 | 12:04 | 916 | 23.84 | 22.29 | 968.08 | 9:18 | 916 | 24.64 | 22.35 | 967.94
19:53 | 977 | 24.33 | 22.23 | 967.03 | 12:05 | 977 | 23.83 | 22.29 | 968.12 | 9:19 | 977 | 24.6 | 22.35 | 967.95
19:54 | 1038 | 24.23 | 22.25 | 967.04 | 12:06 | 1038 | 23.8 | 22.3 | 968.13 | 9:20 | 1038 | 24.53 | 22.35 | 967.96
19:55 | 1099 | 24.1 | 22.28 | 967.08 | 12:07 | 1099 | 23.76 | 22.3 | 968.14 | 9:21 | 1099 | 24.58 | 22.34 | 967.99
19:56 | 1160 | 23.9 | 22.28 | 967.08 | 12:08 | 1160 | 23.68 | 22.3 | 968.13 | 9:22 | 1160 | 24.51 | 22.33 | 967.95
19:57 | 1221 | 23.86 | 22.25 | 967.11 | 12:09 | 1221 | 23.68 | 22.3 | 968.11 | 9:23 | 1221 | 24.46 | 22.34 | 967.96
19:58 | 1282 | 23.7 | 22.22 | 967.08 | 12:10 | 1282 | 23.7 | 22.3 | 968.08 | 9:24 | 1282 | 24.43 | 22.35 | 967.96
19:59 | 1343 | 23.64 | 22.23 | 967.08 | 12:11 | 1343 | 23.52 | 22.29 | 968.13 | 9:25 | 1343 | 24.42 | 22.37 | 967.95
20:00 | 1404 | 23.54 | 22.23 | 967.1 | 12:12 | 1404 | 23.54 | 22.29 | 968.15 | 9:26 | 1404 | 24.36 | 22.34 | 967.94
20:01 | 1465 | 23.41 | 22.24 | 967.07 | 12:13 | 1465 | 23.52 | 22.3 | 968.18 | 9:27 | 1465 | 24.37 | 22.36 | 967.95
20:02 | 1526 | 23.32 | 22.26 | 967.07 | 12:14 | 1526 | 23.4 | 22.3 | 968.2 | 9:28 | 1526 | 24.27 | 22.36 | 967.91
20:03 | 1587 | 23.23 | 22.26 | 967.08 | 12:15 | 1587 | 23.4 | 22.3 | 968.2 | 9:29 | 1587 | 24.31 | 22.36 | 967.91
20:04 | 1648 | 23.07 | 22.23 | 967.05 | 12:16 | 1648 | 23.32 | 22.29 | 968.2 | 9:30 | 1648 | 24.26 | 22.36 | 967.92
20:05 | 1709 | 23.02 | 22.23 | 967.03 | 12:17 | 1709 | 23.3 | 22.3 | 968.27 | 9:31 | 1709 | 24.28 | 22.36 | 967.9
20:06 | 1770 | 22.91 | 22.24 | 967.04 | 12:18 | 1770 | 23.27 | 22.3 | 968.23 | 9:32 | 1770 | 24.2 | 22.35 | 967.92
20:07 | 1831 | 22.81 | 22.24 | 967.05 | 12:19 | 1831 | 23.21 | 22.29 | 968.23 | 9:33 | 1831 | 24.19 | 22.35 | 967.9
20:08 | 1892 | 22.74 | 22.2 | 967.07 | 12:20 | 1892 | 23.14 | 22.3 | 968.24 | 9:34 | 1892 | 24.22 | 22.35 | 967.91
20:09 | 1953 | 22.65 | 22.19 | 967.03 | 12:21 | 1953 | 23.07 | 22.29 | 968.23 | 9:35 | 1953 | 24.07 | 22.36 | 967.94
20:10 | 2014 | 22.52 | 22.11 | 967.09 | 12:22 | 2014 | 23.11 | 22.29 | 968.19 | 9:36 | 2014 | 24.11 | 22.35 | 967.97
20:11 | 2075 | 22.46 | 22.07 | 967.07 | 12:23 | 2075 | 23.06 | 22.3 | 968.2 | 9:37 | 2075 | 24.08 | 22.33 | 967.97
20:12 | 2136 | 22.39 | 22.03 | 967.09 | 12:24 | 2136 | 23.08 | 22.29 | 968.18 | 9:38 | 2136 | 24 | 22.35 | 967.97
20:13 | 2197 | 22.3 | 22.05 | 967.1 | 12:25 | 2197 | 22.99 | 22.29 | 968.24 | 9:39 | 2197 | 23.99 | 22.34 | 967.99
20:14 | 2258 | 22.13 | 22 | 967.09 | 12:26 | 2258 | 22.93 | 22.29 | 968.29 | 9:40 | 2258 | 23.97 | 22.34 | 967.97
20:15 | 2319 | 22.07 | 22 | 967.09 | 12:27 | 2319 | 22.85 | 22.29 | 968.38 | 9:41 | 2319 | 23.98 | 22.33 | 968.01
20:16 | 2380 | 21.95 | 21.99 | 967.09 | 12:28 | 2380 | 22.81 | 22.29 | 968.37 | 9:42 | 2380 | 23.94 | 22.35 | 968.03
20:17 | 2441 | 21.84 | 22 | 967.06 | 12:29 | 2441 | 22.75 | 22.3 | 968.34 | 9:43 | 2441 | 23.94 | 22.37 | 967.98
20:18 | 2502 | 21.76 | 21.99 | 967.08 | 12:30 | 2502 | 22.8 | 22.3 | 968.29 | 9:44 | 2502 | 23.99 | 22.35 | 967.97
20:19 | 2563 | 21.65 | 22 | 967.1 | 12:31 | 2563 | 22.74 | 22.29 | 968.26 | 9:45 | 2563 | 23.91 | 22.34 | 967.96
20:20 | 2624 | 21.56 | 21.99 | 967.1 | 12:32 | 2624 | 22.67 | 22.3 | 968.22 | 9:46 | 2624 | 23.86 | 22.34 | 967.97
20:21 | 2685 | 21.46 | 22.02 | 967.1 | 12:33 | 2685 | 22.7 | 22.29 | 968.26 | 9:47 | 2685 | 23.86 | 22.35 | 967.93
20:22 | 2746 | 21.38 | 22 | 967.08 | 12:34 | 2746 | 22.59 | 22.3 | 968.26 | 9:48 | 2746 | 23.86 | 22.35 | 967.95
20:23 | 2807 | 21.27 | 22.02 | 967.11 | 12:35 | 2807 | 22.56 | 22.29 | 968.3 | 9:49 | 2807 | 23.8 | 22.35 | 967.95
20:24 | 2868 | 21.27 | 22.02 | 967.09 | 12:36 | 2868 | 22.46 | 22.31 | 968.31 | 9:50 | 2868 | 23.8 | 22.35 | 967.95
20:25 | 2929 | 21.09 | 22.02 | 967.06 | 12:37 | 2929 | 22.5 | 22.29 | 968.27 | 9:51 | 2929 | 23.74 | 22.36 | 967.95
20:26 | 2990 | 21.03 | 22.05 | 967.05 | 12:38 | 2990 | 22.43 | 22.3 | 968.29 | 9:52 | 2990 | 23.73 | 22.35 | 967.95
20:27 | 3051 | 20.89 | 22.06 | 967.05 | 12:39 | 3051 | 22.46 | 22.31 | 968.29 | 9:53 | 3051 | 23.63 | 22.35 | 967.95
20:28 | 3112 | 20.9 | 22.04 | 967.04 | 12:40 | 3112 | 22.43 | 22.3 | 968.32 | 9:54 | 3112 | 23.66 | 22.36 | 967.92
20:29 | 3173 | 20.75 | 22.02 | 967.03 | 12:41 | 3173 | 22.37 | 22.3 | 968.35 | 9:55 | 3173 | 23.63 | 22.35 | 967.91
20:30 | 3234 | 20.7 | 21.98 | 967.06 | 12:42 | 3234 | 22.19 | 22.29 | 968.35 | 9:56 | 3234 | 23.62 | 22.35 | 967.91
20:31 | 3295 | 20.58 | 22 | 967.03 | 12:43 | 3295 | 22.23 | 22.3 | 968.36 | 9:57 | 3295 | 23.56 | 22.35 | 967.89
20:32 | 3356 | 20.59 | 22.03 | 967.04 | 12:44 | 3356 | 22.21 | 22.3 | 968.25 | 9:58 | 3356 | 23.59 | 22.34 | 967.91
20:33 | 3417 | 20.46 | 22.01 | 967.04 | 12:45 | 3417 | 22.13 | 22.3 | 968.35 | 9:59 | 3417 | 23.58 | 22.35 | 967.93
20:34 | 3478 | 20.39 | 21.98 | 967.05 | 12:46 | 3478 | 22.09 | 22.29 | 968.4 | 10:00 | 3478 | 23.57 | 22.36 | 967.93
20:35 | 3539 | 20.26 | 21.99 | 967.06 | 12:47 | 3539 | 22.11 | 22.29 | 968.38 | 10:01 | 3539 | 23.54 | 22.35 | 967.91
20:36 | 3600 | 20.21 | 22.01 | 967.07 | 12:48 | 3600 | 22.03 | 22.28 | 968.44 | 10:02 | 3600 | 23.52 | 22.36 | 967.92
[5 empty] | 12:49 | 3661 | 21.98 | 22.29 | 968.47 | 10:03 | 3661 | 23.51 | 22.37 | 967.9
[5 empty] | 12:50 | 3722 | 21.96 | 22.28 | 968.54 | 10:04