## QC4_Summary
Instructions
1. Fill out the table below (columns D through V) based on the most up-to-date test for each detector
2. If a test would supercede the below information, update the table to reflect only the most up-to-date information
3. Follow the instructions on "QC4_Curves" tab that correspond with the test you are entering in this tab
4. Save the excel spreadsheet
| GE1/1 | Test Date | Note | Summary | Voltage Reached V_max | Current @ V_max | V_Drift | R_Equiv (Measured) | R_Equiv (Slope) | %Diff in R_Equiv | Preamplifier | Amplifier |  |  |  |  | Discriminator |  | Spurious Signal @ V_max | Filename | E-Log
[11 empty] | Make & Model | Make & Model | Coarse Gain | Fine Gain | Int. Time | Diff. Time | Make & Model | Threshold |  |  | Short Link URL
| S/N | (YYYY/MM/DD) |  |  | (V) | (uA) | (V) | (MOhm) | (MOhm) | [5 empty] | (ns) | (ns) |  | (mV) | (Hz)
Short Detector: | GE1/1-X-S-CERN-0001 | 2017-05-24 |  | Bipolar pulses are still present after 2.6kV | 4900 | 984.3 | 4626.21 | 5 | 4.9771 | 0.4601% | [7 empty] | 140 | 19.0333 | GE11-X-S-CERN-0001_QC4_20170524.xlsx
| GE1/1-X-S-CERN-0002 | 2017-05-18 |  | QC4 completed w/out problems, SS rate decreased after powering on the det for 3 days | 4900.2 | 984.7 | 4628.09 | 5 | 4.9757 | 0.4884% | CAEN A422A | ORTEC 474 | 4 | 4.5 | 500 | 500 |  | 140 | 22.4 | GE11-X-S-CERN-0002_QC4_20170518.xlsx
| GE1/1-X-S-CERN-0003 | 2017-06-13 |  | QC4 completed w/out problems, bipolar pulses present after 4KV | 4900.2 | 984.45 | 4626.915 | 5 | 4.9766 | 0.4702% | [7 empty] | 140 | 10.3667 | GE11-X-S-CERN-0003_QC4_20170613.xlsx
| GE1/1-X-S-CERN-0004 | 2017-07-06 | repeat qc4 after powering on over night at 3kV after noticing very high rate of counts the previous day | QC4 completed w/out problems, Max SS rate = 22 Hz | [5 empty] | #DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
Long Detector: | GE1/1-X-L-CERN-0001 | [8 empty] | #DIV/0!
| GE1/1-X-L-CERN-0002 | 2017-06-07 | Copper/Kapton sheet to shield the HV circuit on the R/O board | Bipolar pulses are still present after 2.6kV | 4900.6 | 979.45 | 4603.415 | 5 | 5.0017 | -0.03399% | [7 empty] | 140 | 9.7667 | GE11-X-L-CERN-0002-20170607-Test2.xlsx
|  |  |  | Chamber tripped: 1 time @ 4.2kV and 1 time @ 4.7kV | [5 empty] | #DIV/0!
|  |  |  | QC4 completed, SS Rate = 11Hz @ 4.9kV | [5 empty] | #DIV/0!
| GE1/1-X-L-CERN-0003 | 07/13/2017 | Powered on overnight at 4kV | QC4 completed, SS Rate = 19.85Hz @ 4.9kV | 4900.2 | 985.3 | 4630.91 | 5 | 4.9707 | 0.5895% | [7 empty] | 140 | 19.85 | GE11-X-S-CERN-0004_QC4_20170713.xlsx
|  | 2017-08-01 | After fixing RO in the clean room | QC4 stopped at 4.8kV because of high sp rate(85Hz) | 4800.2 | 965.55 | 4538.085 | 5 | 4.9697 | 0.6087% | [7 empty] | 140 | 84.5167
|  | 2017-08-01 | After powering on at 4kV for one hour | QC4 completed, SS Rate =14.45 Hz @ 4.9kV | 4900.2 | 992.4 | 4664.28 | 5 | 4.9627 | 0.7525% | [7 empty] | 140 | 14.45
| GE1/1-X-L-CERN-0004 | 2017-07-27 | Done after fixing faulty screw and doing QC3 | QC4 completed, SS Rate = 21Hz @ 4.9kV. There was trip at 4400V | 4900 | 990.05 | 4653.235 | 5 | 4.9648 | 0.0070899 | [7 empty] | 140 | 21 | GE11-X-L-CERN-0004-20172707-Test1.xlsx
|  | 2017-07-28 | Done after fixing faulty screw and doing QC3; done again because there was a short in GEM foil 3 | QC4 completed, SS Rate = 21.6Hz @ 4.9kV. There was trip at 4600V and multiple trips at 4900V | 4900 | 988.3 | 4645.01 | 5 | 4.9574 | 0.0085932 | [7 empty] | 140 | 21.567 | GE11-X-L-CERN-0004-20172707-Test2.xlsx

## QC4_Long_Detectors
Summary of long detectors Most recent measurement is used

## QC4_Short_Detectors
Summary of short detectors Most recent measurement is used

## QC4_all_detectors
Summary of short +long detectors Most recent measurement is used

## QC4_Short_1
Test Date: | 2017-04-05 |  |  | Test Date: | 2017-05-05 |  |  |  |  | Test Date: | 2017-08-05 |  |  |  |  | Test Date: | 2017-09-05 |  |  |  |  | Test Date: | 2017-09-05
File: | GE11-X-S-CERN-0001_QC4_20170504.xlsx |  |  | File: | GE11-X-S-CERN-0001_QC4_20170505.xlsx |  |  |  |  | File: | GE11-X-S-CERN-0001_QC4_20170508.xlsx |  |  |  |  | File: | GE11-X-S-CERN-0001_QC4_20170509.xlsx |  |  |  |  | File: | GE11-X-S-CERN-0001_QC4_20170524.xlsx
Elog: | http://cmsonline.cern.ch/cms-elog/979847 |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/980159 |  |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/981077 |  |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/981404 |  |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/985715
Description: | Initial |  |  | Description: | after adding HV filter |  |  |  |  | Description: | after powering on for 48h (2kV) |  |  |  |  | Description: | after powering on for 12h (4kV) |  |  |  |  | Description: | After Noticing the bipolar pulses in QC5
Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) | Author: | 140 |  |  |  | Disc Thresh (mV) |  | 140 |  |  |  | Disc Thresh (mV) |  | 140 |  |  |  | Disc Thresh (mV)
Measured R_Equiv (Mohm) | 4.711 |  |  | Measured R_Equiv (Mohm) | [5 empty] | Measured R_Equiv (Mohm) | 5 |  |  |  |  | Measured R_Equiv (Mohm) | 5 |  |  |  |  | Measured R_Equiv (Mohm) | 5
Status: | too high SS rate (expected value is zero) |  |  | Status: | too high SS rate (expected value is zero) |  |  |  |  | Status: | too high SS rate (expected value is zero) |  |  |  |  | Status: | too high SS rate (expected value is zero) |  |  |  |  | Status: | the bipolar pulses were present after 2.6 kV
Vmon | Imon | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Rate | Rate Err
(V) | (uA) | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz)
199 | 42.1 | 0 | 0 | 198.8 | 39.55 | 0 | 0 | 0 | 0 | 199.2 | 39.7 | 0 | 0 | 0 | 0 | 199 | 39.6 | 0 | 0 | 0 | 0 | 199.4 | 39.75 | 0.016667 | 0.016667
399 | 84.4 | 0 | 0 | 399 | 79.4 | 0 | 0 | 0 | 0 | 399 | 79.45 | 0 | 0 | 0 | 0 | 399 | 79.45 | 0 | 0 | 0 | 0 | 399.4 | 79.75 | 0 | 0
599.2 | 126.65 | 0 | 0 | 599 | 119.25 | 0 | 0 | 0 | 0 | 599.2 | 119.3 | 0 | 0 | 0 | 0 | 599.2 | 119.25 | 0 | 0 | 0 | 0 | 599.4 | 119.75 | 0 | 0
799.4 | 169.15 | 0.016667 | 0.016667 | 799.2 | 159.1 | 0 | 0 | 0 | 0 | 799.4 | 159.15 | 0 | 0 | 0 | 0 | 799.4 | 159.1 | 0 | 0 | 0 | 0 | 799.4 | 159.8 | 0 | 0
999.4 | 211.5 | 0.016667 | 0.016667 | 999.4 | 199.05 | 1 | 1 | 0.016667 | 0.016667 | 999.4 | 199.05 | 0 | 0 | 0 | 0 | 999.4 | 199.05 | 0 | 0 | 0 | 0 | 999.6 | 199.9 | 0 | 0
1199.6 | 253.9 | 0.81667 | 0.11667 | 1199.4 | 238.95 | 6 | 2.4495 | 0.1 | 0.040825 | 1199.6 | 239 | 0 | 0 | 0 | 0 | 1199.6 | 238.95 | 0 | 0 | 0 | 0 | 1199.8 | 240 | 0 | 0
1399.6 | 296.3 | 2.2167 | 0.19221 | 1400 | 278.9 | 14 | 3.7417 | 0.23333 | 0.062361 | 1399.6 | 278.9 | 0 | 0 | 0 | 0 | 1399.6 | 278.85 | 0 | 0 | 0 | 0 | 1399.6 | 280.1 | 0 | 0
1599.6 | 338.76 | 5.1833 | 0.29392 | 1599.6 | 318.85 | 74 | 8.6023 | 1.2333 | 0.14337 | 1599.6 | 318.85 | 3 | 1.7321 | 0.05 | 0.028868 | 1599.6 | 318.8 | 1 | 1 | 0.016667 | 0.016667 | 1599.8 | 320.3 | 0.05 | 0.028868
1799.6 | 381.2 | 8.7167 | 0.38115 | 1799.6 | 358.8 | 128 | 11.3137 | 2.1333 | 0.18856 | 1799.6 | 358.8 | 13 | 3.6056 | 0.21667 | 0.060093 | 1799.6 | 358.75 | 1 | 1 | 0.016667 | 0.016667 | 1799.8 | 360.45 | 0.31667 | 0.072648
1999.6 | 423.7 | 11.8833 | 0.44503 | 1999.6 | 398.75 | 193 | 13.8924 | 3.2167 | 0.23154 | 1999.6 | 398.75 | 42 | 6.4807 | 0.7 | 0.10801 | 1999.6 | 398.7 | 2 | 1.4142 | 0.033333 | 0.02357 | 2000 | 400.6 | 0.53333 | 0.094281
2199.6 | 466.2 | 15.2 | 0.50332 | 2199.8 | 438.75 | 269 | 16.4012 | 4.4833 | 0.27335 | 2199.8 | 438.75 | 89 | 9.434 | 1.4833 | 0.15723 | 2199.6 | 438.7 | 7 | 2.6458 | 0.11667 | 0.044096 | 2199.8 | 440.75 | 1.0667 | 0.13333
2399.6 | 508.6 | 17.85 | 0.54544 | 2399.6 | 478.65 | 311 | 17.6352 | 5.1833 | 0.29392 | 2399.6 | 478.7 | 146 | 12.083 | 2.4333 | 0.20138 | 2399.6 | 478.65 | 9 | 3 | 0.15 | 0.05 | 2400 | 480.95 | 2.0167 | 0.18333
2599.8 | 551.1 | 20.5667 | 0.58547 | 2599.6 | 518.7 | 372 | 19.2873 | 6.2 | 0.32146 | 2599.6 | 518.75 | 200 | 14.1421 | 3.3333 | 0.2357 | 2599.6 | 518.7 | 16 | 4 | 0.26667 | 0.066667 | 2599.8 | 521.1 | 1.85 | 0.17559
2799.6 | 593.65 | 21.4667 | 0.59815 | 2799.6 | 558.7 | 486 | 22.0454 | 8.1 | 0.36742 | 2799.6 | 558.75 | 249 | 15.7797 | 4.15 | 0.263 | 2799.6 | 558.65 | 40 | 6.3246 | 0.66667 | 0.10541 | 2800 | 561.35 | 1.8667 | 0.17638
2999.8 | 636.15 | 22.8167 | 0.61667 | 2999.8 | 598.7 | 556 | 23.5797 | 9.2667 | 0.39299 | 2999.8 | 598.75 | 277 | 16.6433 | 4.6167 | 0.27739 | 2999.8 | 598.75 | 51 | 7.1414 | 0.85 | 0.11902 | 3000 | 601.55 | 2.7167 | 0.21279
3099.8 | 657.4 | 23.8833 | 0.63092 | 3099.8 | 618.7 | 542 | 23.2809 | 9.0333 | 0.38801 | 3099.8 | 618.75 | 287 | 16.9411 | 4.7833 | 0.28235 | 3099.8 | 618.65 | 73 | 8.544 | 1.2167 | 0.1424 | 3100 | 621.6 | 2.7667 | 0.21473
3199.8 | 678.7 | 23 | 0.61914 | 3199.8 | 638.65 | 512 | 22.6274 | 8.5333 | 0.37712 | 3199.8 | 638.75 | 314 | 17.72 | 5.2333 | 0.29533 | 3199.8 | 638.6 | 87 | 9.3274 | 1.45 | 0.15546 | 3200 | 641.7 | 3.05 | 0.22546
3299.6 | 699.95 | 24.2 | 0.63509 | 3299.6 | 658.7 | 525 | 22.9129 | 8.75 | 0.38188 | 3299.6 | 658.75 | 325 | 18.0278 | 5.4167 | 0.30046 | 3299.6 | 658.6 | 91 | 9.5394 | 1.5167 | 0.15899 | 3300 | 661.8 | 3.4333 | 0.23921
3399.8 | 721.2 | 22.9333 | 0.61824 | 3399.8 | 678.75 | 574 | 23.9583 | 9.5667 | 0.3993 | 3399.8 | 678.75 | 357 | 18.8944 | 5.95 | 0.31491 | 3399.8 | 678.6 | 120 | 10.9545 | 2 | 0.18257 | 3400 | 681.9 | 6.9833 | 0.34116
3499.6 | 742.5 | 25.0667 | 0.64636 | 3499.6 | 698.7 | 618 | 24.8596 | 10.3 | 0.41433 | 3499.6 | 698.75 | 297 | 17.2337 | 4.95 | 0.28723 | 3499.6 | 698.6 | 118 | 10.8628 | 1.9667 | 0.18105 | 3500 | 702.05 | 4.25 | 0.26615
3599.8 | 763.8 | 25.15 | 0.64743 | 3599.8 | 718.75 | 654 | 25.5734 | 10.9 | 0.42622 | 3600 | 718.8 | 377 | 19.4165 | 6.2833 | 0.32361 | 3599.8 | 718.6 | 133 | 11.5326 | 2.2167 | 0.19221 | 3600 | 722.1 | 4.1667 | 0.26352
3699.8 | 785.1 | 25.1833 | 0.64786 | 3699.8 | 738.75 | 661 | 25.7099 | 11.0167 | 0.4285 | 3699.8 | 738.8 | 388 | 19.6977 | 6.4667 | 0.3283 | 3699.8 | 738.65 | 166 | 12.8841 | 2.7667 | 0.21473 | 3700 | 742.25 | 4.2833 | 0.26719
3799.8 | 806.45 | 27.5 | 0.677 | 3799.8 | 758.8 | 703 | 26.5141 | 11.7167 | 0.4419 | 3799.8 | 758.85 | 423 | 20.567 | 7.05 | 0.34278 | 3799.8 | 758.65 | 171 | 13.0767 | 2.85 | 0.21794 | 3800 | 762.35 | 5.3167 | 0.29768
3899.8 | 827.8 | 28.5167 | 0.6894 | 3899.8 | 778.85 | 708 | 26.6083 | 11.8 | 0.44347 | 3899.8 | 778.9 | 462 | 21.4942 | 7.7 | 0.35824 | 3900 | 778.7 | 200 | 14.1421 | 3.3333 | 0.2357 | 3900 | 782.45 | 5.9167 | 0.31402
4000 | 849.15 | 28.65 | 0.69101 | 4000 | 798.85 | 727 | 26.9629 | 12.1167 | 0.44938 | 4000 | 798.9 | 452 | 21.2603 | 7.5333 | 0.35434 | 4000 | 798.7 | 236 | 15.3623 | 3.9333 | 0.25604 | 4000 | 802.55 | 6.45 | 0.32787
4100 | 870.55 | 30.1333 | 0.70868 | 4100 | 818.95 | 945 | 30.7409 | 15.75 | 0.51235 | 4100 | 819 | 509 | 22.561 | 8.4833 | 0.37602 | 4100 | 818.8 | 254 | 15.9374 | 4.2333 | 0.26562 | 4099.8 | 822.65 | 6.7167 | 0.33458
4200 | 891.95 | 29.9 | 0.70593 | 4200 | 839 | 738 | 27.1662 | 12.3 | 0.45277 | 4200.2 | 839 | 490 | 22.1359 | 8.1667 | 0.36893 | 4200 | 838.85 | 281 | 16.7631 | 4.6833 | 0.27938 | 4200 | 842.85 | 6.8667 | 0.3383
4300 | 913.4 | 31.4833 | 0.72438 | 4300 | 859.1 | 874 | 29.5635 | 14.5667 | 0.49272 | 4300 | 859.1 | 546 | 23.3666 | 9.1 | 0.38944 | 4300.2 | 858.95 | 307 | 17.5214 | 5.1167 | 0.29202 | 4300 | 863 | 7.25 | 0.34761
4400.2 | 934.9 | 34.9667 | 0.7634 | 4400 | 879.2 | 886 | 29.7658 | 14.7667 | 0.4961 | 4400.2 | 879.2 | 607 | 24.6374 | 10.1167 | 0.41062 | 4400.2 | 879.05 | 373 | 19.3132 | 6.2167 | 0.32189 | 4400 | 883.15 | 7.55 | 0.35473
4500.2 | 956.4 | 38.95 | 0.80571 | 4500.2 | 899.35 | 988 | 31.4325 | 16.4667 | 0.52387 | 4500.4 | 899.35 | 661 | 25.7099 | 11.0167 | 0.4285 | 4500.2 | 899.2 | 450 | 21.2132 | 7.5 | 0.35355 | 4500 | 903.35 | 8.3167 | 0.37231
4600.2 | 978 | 45.0333 | 0.86635 | 4600 | 919.45 | 1108 | 33.2866 | 18.4667 | 0.55478 | 4600.2 | 919.45 | 744 | 27.2764 | 12.4 | 0.45461 | 4600.2 | 919.35 | 499 | 22.3383 | 8.3167 | 0.37231 | 4600 | 923.5 | 9.4667 | 0.39721
4700.2 | 999.8 | 49.1167 | 0.90477 | 4700.4 | 939.6 | 1342 | 36.6333 | 22.3667 | 0.61056 | 4700.2 | 939.6 | 930 | 30.4959 | 15.5 | 0.50827 | 4700.4 | 939.5 | 695 | 26.3629 | 11.5833 | 0.43938 | 4700 | 943.7 | 11.2667 | 0.43333
|  |  |  | 4800.4 | 959.8 | 1604 | 40.05 | 26.7333 | 0.6675 | 4800.4 | 959.8 | 1122 | 33.4963 | 18.7 | 0.55827 | 4800.4 | 959.7 | 974 | 31.209 | 16.2333 | 0.52015 | 4800 | 964.05 | 14.6667 | 0.49441
|  |  |  | 4900 | 979.85 | 1884 | 43.4051 | 31.4 | 0.72342 | 4900.2 | 979.9 | 1516 | 38.9358 | 25.2667 | 0.64893 | 4900.4 | 979.9 | 1223 | 34.9714 | 20.3833 | 0.58286 | 4900 | 984.3 | 19.0333 | 0.56322

## QC4_Short_2
Test Date: | 2017-08-05 |  |  |  |  | Test Date: | 2017-09-05 |  |  |  |  | Test Date: | 05/18/2017 |  |  |  |  | Test Date: | 2017-09-06
File: | GE11-X-S-CERN-0002_QC4_20170508.xlsx |  |  |  |  | File: | GE11-X-S-CERN-0002_QC4_20170509.xlsx |  |  |  |  | File: | GE11-X-S-CERN-0002_QC4_20170518.xlsx |  |  |  |  | File: | GE11-X-S-CERN-0002_QC4_20170609.xlsx
Elog: | http://cmsonline.cern.ch/cms-elog/980952 |  |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/981470 |  |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/981470 |  |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/989726
Description: | Initial Test |  |  |  |  | Description: | power on for 12h (4kV) |  |  |  |  | Description: | power on for 3days (4kV) |  |  |  |  | Description: | After QC2 fast test
Measured R_Equiv (Mohm) | 5 |  |  |  |  | Measured R_Equiv (Mohm) | 5 |  |  |  |  | Measured R_Equiv (Mohm) | 5 |  |  |  |  | Measured R_Equiv (Mohm) | 5
Disc Thresh (mV) |  | 140 |  |  |  | Disc Thresh (mV) |  | 140 |  |  |  | Disc Thresh (mV) |  | 140 |  |  |  | Disc Thresh (mV) |  | 140
Status: | Passed |  |  |  |  | Status: | Passed |  |  |  |  | Status: | Passed |  |  |  |  | Status: | Passed
Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err
(V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz)
199.6 | 39.85 | 0 | 0 | 0 | 0 | 199.8 | 39.9 | 0 | 0 | 0 | 0 | 199.6 | 39.85 |  |  | 0 | 0 | 199.2 | 39.65 | 0 | 0 | 0 | 0
399.6 | 79.8 | 0 | 0 | 0 | 0 | 399.6 | 79.85 | 0 | 0 | 0 | 0 | 399.6 | 79.8 |  |  | 0 | 0 | 399.2 | 79.55 | 0 | 0 | 0 | 0
599.4 | 119.75 | 0 | 0 | 0 | 0 | 599.4 | 119.85 | 0 | 0 | 0 | 0 | 599.4 | 119.8 |  |  | 0 | 0 | 599.6 | 119.6 | 0 | 0 | 0 | 0
799.6 | 159.8 | 0 | 0 | 0 | 0 | 799.6 | 159.9 | 0 | 0 | 0 | 0 | 799.4 | 159.85 |  |  | 0 | 0 | 799.8 | 159.6 | 0 | 0 | 0 | 0
999.6 | 199.85 | 0 | 0 | 0 | 0 | 999.6 | 200 | 0 | 0 | 0 | 0 | 999.6 | 199.95 |  |  | 0 | 0 | 999.8 | 199.6 | 0 | 0 | 0 | 0
1199.8 | 239.95 | 0 | 0 | 0 | 0 | 1199.8 | 240.05 | 0 | 0 | 0 | 0 | 1199.8 | 240 |  |  | 0 | 0 | 1199.8 | 239.6 | 0 | 0 | 0 | 0
1399.8 | 279.95 | 0 | 0 | 0 | 0 | 1399.6 | 280.15 | 0 | 0 | 0 | 0 | 1399.6 | 280.1 |  |  | 0 | 0 | 1399.8 | 279.6 | 0 | 0 | 0 | 0
1600 | 320.05 | 1 | 1 | 0.016667 | 0.016667 | 1599.8 | 320.35 | 1 | 1 | 0.016667 | 0.016667 | 1599.8 | 320.25 |  |  | 0 | 0 | 1599.8 | 319.7 | 0 | 0 | 0 | 0
1799.8 | 360.2 | 12 | 3.4641 | 0.2 | 0.057735 | 1800 | 360.5 | 1 | 1 | 0.016667 | 0.016667 | 1800 | 360.45 |  |  | 0 | 0 | 1800.2 | 359.85 | 0 | 0 | 0 | 0
1999.8 | 400.35 | 36 | 6 | 0.6 | 0.1 | 2000 | 400.7 | 1 | 1 | 0.016667 | 0.016667 | 2000 | 400.6 |  |  | 0 | 0 | 1999.8 | 399.95 | 0 | 0 | 0 | 0
2200 | 440.5 | 67 | 8.1854 | 1.1167 | 0.13642 | 2199.8 | 440.8 | 1 | 1 | 0.016667 | 0.016667 | 2200 | 440.75 |  |  | 0 | 0 | 2200.2 | 439.95 | 0 | 0 | 0 | 0
2400 | 480.6 | 116 | 10.7703 | 1.9333 | 0.17951 | 2399.8 | 481 | 1 | 1 | 0.016667 | 0.016667 | 2400 | 480.95 |  |  | 0.066667 | 0.033333 | 2400.2 | 480.15 | 1 | 1 | 0.016667 | 0.016667
2599.8 | 520.75 | 157 | 12.53 | 2.6167 | 0.20883 | 2599.6 | 521.2 | 3 | 1.7321 | 0.05 | 0.028868 | 2600 | 521.15 |  |  | 0.05 | 0.028868 | 2600.2 | 520.25 | 1 | 1 | 0.016667 | 0.016667
2800 | 560.95 | 219 | 14.7986 | 3.65 | 0.24664 | 2799.4 | 561.45 | 6 | 2.4495 | 0.1 | 0.040825 | 2800 | 561.4 |  |  | 0.033333 | 0.02357 | 2800.2 | 560.35 | 5 | 2.2361 | 0.083333 | 0.037268
2999.8 | 601.15 | 282 | 16.7929 | 4.7 | 0.27988 | 2999.8 | 601.65 | 39 | 6.245 | 0.65 | 0.10408 | 3000 | 601.6 |  |  | 0.15 | 0.05 | 3000.2 | 600.55 | 12 | 3.4641 | 0.2 | 0.057735
3100 | 621.25 | 288 | 16.9706 | 4.8 | 0.28284 | 3099 | 621.7 | 35 | 5.9161 | 0.58333 | 0.098601 | 3100 | 621.65 |  |  | 0.16667 | 0.052705 | 3100.2 | 620.6 | 24 | 4.899 | 0.4 | 0.08165
3200 | 641.35 | 313 | 17.6918 | 5.2167 | 0.29486 | 3200 | 641.8 | 58 | 7.6158 | 0.96667 | 0.12693 | 3200.2 | 641.75 |  |  | 0.21667 | 0.060093 | 3200.2 | 640.7 | 44 | 6.6332 | 0.73333 | 0.11055
3299.8 | 661.4 | 309 | 17.5784 | 5.15 | 0.29297 | 3300 | 661.9 | 83 | 9.1104 | 1.3833 | 0.15184 | 3300 | 661.85 |  |  | 0.35 | 0.076376 | 3300.4 | 660.75 | 56 | 7.4833 | 0.93333 | 0.12472
3400 | 681.5 | 312 | 17.6635 | 5.2 | 0.29439 | 3400 | 682 | 102 | 10.0995 | 1.7 | 0.16833 | 3400 | 681.95 |  |  | 0.61667 | 0.10138 | 3400.6 | 680.9 | 94 | 9.6954 | 1.5667 | 0.16159
3500 | 701.65 | 324 | 18 | 5.4 | 0.3 | 3500 | 702.15 | 110 | 10.4881 | 1.8333 | 0.1748 | 3500 | 702.1 |  |  | 0.68333 | 0.10672 | 3500.4 | 701 | 131 | 11.4455 | 2.1833 | 0.19076
3600 | 721.75 | 435 | 20.8567 | 7.25 | 0.34761 | 3600 | 722.25 | 143 | 11.9583 | 2.3833 | 0.1993 | 3600 | 722.2 |  |  | 0.86667 | 0.12019 | 3600.4 | 721.05 | 130 | 11.4018 | 2.1667 | 0.19003
3699.8 | 741.85 | 386 | 19.6469 | 6.4333 | 0.32745 | 3699.8 | 742.35 | 209 | 14.4568 | 3.4833 | 0.24095 | 3700 | 742.3 |  |  | 0.95 | 0.12583 | 3700.4 | 741.2 | 172 | 13.1149 | 2.8667 | 0.21858
3799.8 | 762 | 500 | 22.3607 | 8.3333 | 0.37268 | 3799.8 | 762.45 | 214 | 14.6287 | 3.5667 | 0.24381 | 3800 | 762.4 |  |  | 1.3167 | 0.14814 | 3800.4 | 761.3 | 174 | 13.1909 | 2.9 | 0.21985
3899.8 | 782.15 | 476 | 21.8174 | 7.9333 | 0.36362 | 3899.8 | 782.6 | 227 | 15.0665 | 3.7833 | 0.25111 | 3900 | 782.6 |  |  | 1.5 | 0.15811 | 3900.4 | 781.4 | 191 | 13.8203 | 3.1833 | 0.23034
3999.8 | 802.3 | 548 | 23.4094 | 9.1333 | 0.39016 | 3999.8 | 802.65 | 255 | 15.9687 | 4.25 | 0.26615 | 4000 | 802.65 |  |  | 1.6 | 0.1633 | 4000.4 | 801.45 | 237 | 15.3948 | 3.95 | 0.25658
4099.8 | 822.4 | 613 | 24.7588 | 10.2167 | 0.41265 | 4099.8 | 822.8 | 314 | 17.72 | 5.2333 | 0.29533 | 4099.8 | 822.75 |  |  | 1.6167 | 0.16415 | 4100.4 | 821.6 | 259 | 16.0935 | 4.3167 | 0.26822
4200 | 842.6 | 601 | 24.5153 | 10.0167 | 0.40859 | 4199.8 | 843 | 325 | 18.0278 | 5.4167 | 0.30046 | 4200 | 843 |  |  | 2.0167 | 0.18333 | 4200.4 | 841.75 | 301 | 17.3494 | 5.0167 | 0.28916
4300 | 862.85 | 678 | 26.0384 | 11.3 | 0.43397 | 4300 | 863.15 | 336 | 18.3303 | 5.6 | 0.30551 | 4300 | 863.2 |  |  | 2.4833 | 0.20344 | 4300.4 | 861.9 | 338 | 18.3848 | 5.6333 | 0.30641
4399.8 | 883 | 725 | 26.9258 | 12.0833 | 0.44876 | 4399.8 | 883.35 | 368 | 19.1833 | 6.1333 | 0.31972 | 4400 | 883.35 |  |  | 2.2833 | 0.19508 | 4400.4 | 882.1 | 367 | 19.1572 | 6.1167 | 0.31929
4500 | 903.25 | 768 | 27.7128 | 12.8 | 0.46188 | 4499.8 | 903.55 | 367 | 19.1572 | 6.1167 | 0.31929 | 4500 | 903.55 |  |  | 2.5667 | 0.20683 | 4500.4 | 902.35 | 391 | 19.7737 | 6.5167 | 0.32956
4599.8 | 923.5 | 675 | 25.9808 | 11.25 | 0.43301 | 4599.6 | 923.75 | 408 | 20.199 | 6.8 | 0.33665 | 4599.8 | 923.8 |  |  | 2.85 | 0.21794 | 4600.4 | 922.6 | 401 | 20.025 | 6.6833 | 0.33375
4699.8 | 943.75 | 858 | 29.2916 | 14.3 | 0.48819 | 4699.8 | 944 | 580 | 24.0832 | 9.6667 | 0.40139 | 4700 | 944.05 |  |  | 3.2333 | 0.23214 | 4700.4 | 942.9 | 514 | 22.6716 | 8.5667 | 0.37786
4800 | 964.1 | 1351 | 36.756 | 22.5167 | 0.6126 | 4800 | 964.35 | 909 | 30.1496 | 15.15 | 0.50249 | 4800.2 | 964.35 |  |  | 3.7333 | 0.24944 | 4800.4 | 963.15 | 716 | 26.7582 | 11.9333 | 0.44597
4899.8 | 984.5 | 2175 | 46.6369 | 36.25 | 0.77728 | 4899.8 | 984.6 | 1344 | 36.6606 | 22.4 | 0.61101 | 4900.2 | 984.7 |  |  | 5.8 | 0.31091 | 4900.6 | 983.5 | 965 | 31.0644 | 16.0833 | 0.51774

## QC4_Short_3
Test Date: | 2017-06-09 |  |  |  |  | Test Date: | 2017-06-12 |  |  |  |  | Test Date: | 2017-06-13 |  |  |  |  | Test Date: | 2017-06-13
File: | GE11-X-S-CERN-0003_QC4_20170609 |  |  |  |  | File: | GE11-X-S-CERN-0003_QC4_20170612.xlsx |  |  |  |  | File: | GE11-X-S-CERN-0003_QC4_20170613_Test_1.xlsx |  |  |  |  | File: | GE11-X-S-CERN-0003_QC4_20170613_Test_2.xlsx
Elog: | http://cmsonline.cern.ch/cms-elog/989693 |  |  |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/990188 |  |  |  |  | Elog: | [5 empty] | Elog:
Description: | Initial Test |  |  |  |  | Description: | power on during the weekend (at 3 kV) |  |  |  |  | Description: | power on overnight (at 4 kV) |  |  |  |  | Description:
Measured R_Equiv (Mohm) | 5 |  |  |  |  | Measured R_Equiv (Mohm) | 5 |  |  |  |  | Measured R_Equiv (Mohm) | 5 |  |  |  |  | Measured R_Equiv (Mohm) | 5
Disc Thresh (mV) |  | 140 |  |  |  | Disc Thresh (mV) |  | 140 |  |  |  | Disc Thresh (mV) |  | 140 |  |  |  | Disc Thresh (mV) |  | 140
Status: | Passed |  |  |  |  | Status: | Passed |  |  |  |  | Status: | Passed |  |  |  |  | Status: | Passed
Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err | Vmon | Imon | Counts | Count_Err | Rate | Rate Err
(V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz) | (V) | (uA) |  |  | (Hz) | (Hz)
199.2 | 39.7 | 0 | 0 | 0 | 0 | 199.2 | 39.75 | 0 | 0 | 0 | 0 | 199.8 | 39.9 | 0 | 0 | 0 | 0 | 199.4 | 39.75 | 0 | 0 | 0 | 0
399.4 | 79.75 | 0 | 0 | 0 | 0 | 399.2 | 79.75 | 0 | 0 | 0 | 0 | 399.6 | 79.85 | 0 | 0 | 0 | 0 | 399.4 | 79.8 | 0 | 0 | 0 | 0
599.4 | 119.7 | 2 | 1.4142 | 0.033333 | 0.02357 | 599.2 | 119.75 | 0 | 0 | 0 | 0 | 599.4 | 119.85 | 0 | 0 | 0 | 0 | 599.4 | 119.8 | 0 | 0 | 0 | 0
799.4 | 159.75 | 0 | 0 | 0 | 0 | 799.2 | 159.8 | 0 | 0 | 0 | 0 | 799.6 | 159.9 | 0 | 0 | 0 | 0 | 799.4 | 159.85 | 0 | 0 | 0 | 0
999.6 | 199.8 | 1 | 1 | 0.016667 | 0.016667 | 999.4 | 199.9 | 1 | 1 | 0.016667 | 0.016667 | 999.6 | 199.95 | 2 | 1.4142 | 0.033333 | 0.02357 | 999.6 | 199.95 | 0 | 0 | 0 | 0
1199.8 | 239.9 | 3 | 1.7321 | 0.05 | 0.028868 | 1199.6 | 240 | 0 | 0 | 0 | 0 | 1199.6 | 240.05 | 0 | 0 | 0 | 0 | 1199.8 | 240.05 | 0 | 0 | 0 | 0
1399.6 | 279.9 | 15 | 3.873 | 0.25 | 0.06455 | 1399.6 | 280.1 | 0 | 0 | 0 | 0 | 1399.6 | 280.1 | 1 | 1 | 0.016667 | 0.016667 | 1399.6 | 280.15 | 0 | 0 | 0 | 0
1599.8 | 320.05 | 30 | 5.4772 | 0.5 | 0.091287 | 1599.8 | 320.25 | 0 | 0 | 0 | 0 | 1599.8 | 320.3 | 0 | 0 | 0 | 0 | 1599.8 | 320.3 | 0 | 0 | 0 | 0
1799.8 | 360.2 | 62 | 7.874 | 1.0333 | 0.13123 | 1799.8 | 360.4 | 5 | 2.2361 | 0.083333 | 0.037268 | 1799.8 | 360.45 | 0 | 0 | 0 | 0 | 1799.8 | 360.5 | 2 | 1.4142 | 0.033333 | 0.02357
2000 | 400.35 | 99 | 9.9499 | 1.65 | 0.16583 | 1999.8 | 400.6 | 9 | 3 | 0.15 | 0.05 | 1999.8 | 400.6 | 1 | 1 | 0.016667 | 0.016667 | 1999.8 | 400.65 | 11 | 3.3166 | 0.18333 | 0.055277
2200 | 440.45 | 178 | 13.3417 | 2.9667 | 0.22236 | 2199.8 | 440.7 | 17 | 4.1231 | 0.28333 | 0.068718 | 2199.8 | 440.7 | 0 | 0 | 0 | 0 | 2199.8 | 440.8 | 26 | 5.099 | 0.43333 | 0.084984
2400 | 480.6 | 297 | 17.2337 | 4.95 | 0.28723 | 2399.8 | 480.9 | 28 | 5.2915 | 0.46667 | 0.088192 | 2399.8 | 480.9 | 1 | 1 | 0.016667 | 0.016667 | 2399.8 | 480.95 | 26 | 5.099 | 0.43333 | 0.084984
2600 | 520.7 | 438 | 20.9284 | 7.3 | 0.34881 | 2599.8 | 521.05 | 34 | 5.831 | 0.56667 | 0.097183 | 2599.8 | 521.05 | 2 | 1.4142 | 0.033333 | 0.02357 | 2599.8 | 521.15 | 45 | 6.7082 | 0.75 | 0.1118
2800 | 560.9 | 817 | 28.5832 | 13.6167 | 0.47639 | 2800 | 561.3 | 69 | 8.3066 | 1.15 | 0.13844 | 2800 | 561.25 | 11 | 3.3166 | 0.18333 | 0.055277 | 2800 | 561.35 | 56 | 7.4833 | 0.93333 | 0.12472
3000 | 601.1 | 936 | 30.5941 | 15.6 | 0.5099 | 3000 | 601.45 | 110 | 10.4881 | 1.8333 | 0.1748 | 3000 | 601.5 | 12 | 3.4641 | 0.2 | 0.057735 | 3000 | 601.55 | 89 | 9.434 | 1.4833 | 0.15723
3100 | 621.2 | 909 | 30.1496 | 15.15 | 0.50249 | 3100 | 621.55 | 88 | 9.3808 | 1.4667 | 0.15635 | 3100 | 621.55 | 17 | 4.1231 | 0.28333 | 0.068718 | 3100 | 621.6 | 94 | 9.6954 | 1.5667 | 0.16159
3200.2 | 641.3 | 1001 | 31.6386 | 16.6833 | 0.52731 | 3200 | 641.6 | 110 | 10.4881 | 1.8333 | 0.1748 | 3200 | 641.65 | 23 | 4.7958 | 0.38333 | 0.079931 | 3200 | 641.7 | 88 | 9.3808 | 1.4667 | 0.15635
3300 | 661.35 | 1044 | 32.311 | 17.4 | 0.53852 | 3300 | 661.7 | 142 | 11.9164 | 2.3667 | 0.19861 | 3300 | 661.7 | 30 | 5.4772 | 0.5 | 0.091287 | 3300 | 661.8 | 126 | 11.225 | 2.1 | 0.18708
3400.2 | 681.45 | 1076 | 32.8024 | 17.9333 | 0.54671 | 3400 | 681.8 | 135 | 11.619 | 2.25 | 0.19365 | 3400 | 681.8 | 44 | 6.6332 | 0.73333 | 0.11055 | 3400 | 681.9 | 124 | 11.1355 | 2.0667 | 0.18559
3500.2 | 701.6 | 1178 | 34.322 | 19.6333 | 0.57203 | 3500 | 701.95 | 172 | 13.1149 | 2.8667 | 0.21858 | 3500 | 701.95 | 65 | 8.0623 | 1.0833 | 0.13437 | 3500 | 702 | 143 | 11.9583 | 2.3833 | 0.1993
3600 | 721.7 | 1113 | 33.3617 | 18.55 | 0.55603 | 3600 | 722.05 | 198 | 14.0712 | 3.3 | 0.23452 | 3600 | 722 | 83 | 9.1104 | 1.3833 | 0.15184 | 3600 | 722.1 | 149 | 12.2066 | 2.4833 | 0.20344
3700.2 | 741.85 | 1142 | 33.7935 | 19.0333 | 0.56322 | 3700 | 742.15 | 169 | 13 | 2.8167 | 0.21667 | 3700 | 742.1 | 114 | 10.6771 | 1.9 | 0.17795 | 3700 | 742.2 | 158 | 12.5698 | 2.6333 | 0.2095
3800 | 761.95 | 1133 | 33.6601 | 18.8833 | 0.561 | 3800 | 762.25 | 200 | 14.1421 | 3.3333 | 0.2357 | 3800 | 762.2 | 128 | 11.3137 | 2.1333 | 0.18856 | 3800 | 762.35 | 189 | 13.7477 | 3.15 | 0.22913
3900 | 782.1 | 1238 | 35.1852 | 20.6333 | 0.58642 | 3900 | 782.4 | 204 | 14.2829 | 3.4 | 0.23805 | 3899.8 | 782.35 | 127 | 11.2694 | 2.1167 | 0.18782 | 3900 | 782.45 | 204 | 14.2829 | 3.4 | 0.23805
4000 | 802.25 | 1295 | 35.9861 | 21.5833 | 0.59977 | 4000 | 802.45 | 233 | 15.2643 | 3.8833 | 0.25441 | 4000 | 802.45 | 166 | 12.8841 | 2.7667 | 0.21473 | 4000 | 802.55 | 187 | 13.6748 | 3.1167 | 0.22791
4100 | 822.35 | 1265 | 35.5668 | 21.0833 | 0.59278 | 4099.8 | 822.6 | 220 | 14.8324 | 3.6667 | 0.24721 | 4099.8 | 822.55 | 207 | 14.3875 | 3.45 | 0.23979 | 4099.8 | 822.75 | 216 | 14.6969 | 3.6 | 0.24495
4200 | 842.6 | 1332 | 36.4966 | 22.2 | 0.60828 | 4200 | 842.8 | 232 | 15.2315 | 3.8667 | 0.25386 | 4200 | 842.7 | 183 | 13.5277 | 3.05 | 0.22546 | 4200 | 842.9 | 195 | 13.9642 | 3.25 | 0.23274
4300.2 | 862.6 | 1863 | 43.1625 | 31.05 | 0.71937 | 4300 | 863 | 258 | 16.0624 | 4.3 | 0.26771 | 4300 | 862.9 | 228 | 15.0997 | 3.8 | 0.25166 | 4300 | 863.1 | 204 | 14.2829 | 3.4 | 0.23805
4400 | 883.05 | 1610 | 40.1248 | 26.8333 | 0.66875 | 4400 | 883.2 | 260 | 16.1245 | 4.3333 | 0.26874 | 4400 | 883.1 | 229 | 15.1327 | 3.8167 | 0.25221 | 4400 | 883.3 | 287 | 16.9411 | 4.7833 | 0.28235
4500.2 | 903.3 | 1786 | 42.2611 | 29.7667 | 0.70435 | 4500 | 903.4 | 285 | 16.8819 | 4.75 | 0.28137 | 4500 | 903.3 | 212 | 14.5602 | 3.5333 | 0.24267 | 4500 | 903.55 | 542 | 23.2809 | 9.0333 | 0.38801
4600 | 923.55 | 1728 | 41.5692 | 28.8 | 0.69282 | 4600 | 923.65 | 301 | 17.3494 | 5.0167 | 0.28916 | 4599.8 | 923.55 | 255 | 15.9687 | 4.25 | 0.26615 | 4599.8 | 923.8 | 896 | 29.9333 | 14.9333 | 0.49889
4700 | 943.85 | 1933 | 43.9659 | 32.2167 | 0.73276 | 4700 | 943.85 | 298 | 17.2627 | 4.9667 | 0.28771 | 4699.8 | 943.8 | 297 | 17.2337 | 4.95 | 0.28723 | 4700 | 944.05 | 975 | 31.225 | 16.25 | 0.52042
4800.2 | 964.2 | 1958 | 44.2493 | 32.6333 | 0.73749 | 4800.2 | 964.2 | 397 | 19.9249 | 6.6167 | 0.33208 | 4800 | 964.15 | 435 | 20.8567 | 7.25 | 0.34761 | 4800.2 | 964.45 | 742 | 27.2397 | 12.3667 | 0.45399
4900.2 | 984.6 | 2729 | 52.2398 | 45.4833 | 0.87066 | 4900.2 | 984.6 | 588 | 24.2487 | 9.8 | 0.40415 | 4900.2 | 984.45 | 622 | 24.9399 | 10.3667 | 0.41567 | 4900.2 | 984.8 | 852 | 29.189 | 14.2 | 0.48648

## QC4_Short_4
Test Date: | 2017-07-06
File: | GE11-X-S-CERN-0004_QC4_20170706
Elog: | http://cmsonline.cern.ch/cms-elog/995257
Description: | After powering on over night at 3kV
Measured R_Equiv (Mohm) | 5
Disc Thresh (mV) |  | 140
Status: | Passed
Vmon | Imon | Counts | Count_Err | Rate | Rate Err
(V) | (uA) |  |  | (Hz) | (Hz)
199.4 | 39.8 | 0 | 0 | 0 | 0
399.4 | 79.75 | 0 | 0 | 0 | 0
599.6 | 119.75 | 1 | 1 | 0.016667 | 0.016667
799.6 | 159.75 | 0 | 0 | 0 | 0
999.8 | 199.85 | 0 | 0 | 0 | 0
1199.8 | 239.9 | 0 | 0 | 0 | 0
1399.8 | 279.9 | 0 | 0 | 0 | 0
1599.8 | 319.95 | 2 | 1.4142 | 0.033333 | 0.02357
1799.6 | 360.05 | 19 | 4.3589 | 0.31667 | 0.072648
1999.6 | 400.15 | 30 | 5.4772 | 0.5 | 0.091287
2199.8 | 440.15 | 51 | 7.1414 | 0.85 | 0.11902
2399.8 | 480.25 | 113 | 10.6301 | 1.8833 | 0.17717
2600 | 520.35 | 151 | 12.2882 | 2.5167 | 0.2048
2800 | 560.45 | 224 | 14.9666 | 3.7333 | 0.24944
3000 | 600.5 | 303 | 17.4069 | 5.05 | 0.29011
3100.2 | 620.5 | 337 | 18.3576 | 5.6167 | 0.30596
3200 | 640.5 | 360 | 18.9737 | 6 | 0.31623
3300 | 660.5 | 412 | 20.2978 | 6.8667 | 0.3383
3400 | 680.55 | 422 | 20.5426 | 7.0333 | 0.34238
3500.2 | 700.6 | 475 | 21.7945 | 7.9167 | 0.36324
3600 | 720.6 | 521 | 22.8254 | 8.6833 | 0.38042
3700 | 740.62 | 584 | 24.1661 | 9.7333 | 0.40277
3800 | 760.7 | 582 | 24.1247 | 9.7 | 0.40208
3900.2 | 780.7 | 629 | 25.0799 | 10.4833 | 0.418
4000.2 | 800.85 | 669 | 25.865 | 11.15 | 0.43108
4100.2 | 820.9 | 675 | 25.9808 | 11.25 | 0.43301
4200.2 | 840.9 | 701 | 26.4764 | 11.6833 | 0.44127
4300.2 | 861.05 | 732 | 27.0555 | 12.2 | 0.45092
4400.2 | 881.1 | 763 | 27.6225 | 12.7167 | 0.46037
4500.2 | 901.2 | 824 | 28.7054 | 13.7333 | 0.47842
4600 | 921.3 | 846 | 29.0861 | 14.1 | 0.48477
4700 | 941.4 | 937 | 30.6105 | 15.6167 | 0.51017
4800.2 | 961.6 | 1007 | 31.7333 | 16.7833 | 0.52889
4900.2 | 981.75 | 1317 | 36.2905 | 21.95 | 0.60484

## QC4_Long_1
Test Date: | 2017-02-06 |  |  | Test Date: | 2017-08-06
File: | GE11-X-L-CERN-0001_QC4_20170602.xlsx |  |  | File: | GE11-X-L-CERN-0001_QC4_20170608.xlsx
Elog: | http://cmsonline.cern.ch/cms-elog/988255 |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/989692
Description: |  |  |  | Description: | After QC2 Long
Measured R_Equiv (Mohm) | 5 |  |  | Measured R_Equiv (Mohm) | 5
Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) |  | 140
Status: | Passed |  |  | Status: | Passed
Vmon | Imon | Rate | Rate Err | Vmon | Imon | Rate | Rate Err
(V) | (uA) | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz)
199.2 | 39.65 | 0 | 0 | 198.8 | 39.45 | 0 | 0
399.2 | 79.6 | 0 | 0 | 399 | 79.25 | 0 | 0
599.6 | 119.65 | 0 | 0 | 599.2 | 119.05 | 0 | 0
799.8 | 159.65 | 0 | 0 | 799.4 | 158.85 | 0 | 0
999.8 | 199.7 | 0 | 0 | 999.4 | 198.75 | 0 | 0
1199.8 | 239.65 | 0.066667 | 0.033333 | 1199.6 | 238.6 | 0.016667 | 0.016667
1399.8 | 279.75 | 0.21667 | 0.060093 | 1399.6 | 278.45 | 0.15 | 0.05
1599.8 | 319.8 | 0.41667 | 0.083333 | 1599.6 | 318.4 | 0.33333 | 0.074536
1799.8 | 359.9 | 0.71667 | 0.10929 | 1799.6 | 358.25 | 0.68333 | 0.10672
1999.8 | 400 | 1.0333 | 0.13123 | 1999.6 | 398.15 | 0.7 | 0.10801
2200.2 | 440.15 | 1.3333 | 0.14907 | 2199.8 | 438.15 | 0.81667 | 0.11667
2400.2 | 480.25 | 1.5167 | 0.15899 | 2399.6 | 478.05 | 1.1333 | 0.13744
2600.2 | 520.45 | 1.95 | 0.18028 | 2599.8 | 518 | 1.2167 | 0.1424
2800 | 560.55 | 2.75 | 0.21409 | 2799.8 | 557.95 | 1.7167 | 0.16915
3000.2 | 600.7 | 3.3167 | 0.23511 | 2999.8 | 597.95 | 2.1667 | 0.19003
3100.2 | 620.75 | 3.45 | 0.23979 | 3099.8 | 617.9 | 2.6833 | 0.21148
3200.2 | 640.8 | 4.4833 | 0.27335 | 3199.8 | 637.9 | 2.95 | 0.22174
3300.4 | 660.9 | 4.8333 | 0.28382 | 3299.8 | 657.85 | 3.4167 | 0.23863
3400.6 | 681 | 4.9 | 0.28577 | 3399.8 | 677.85 | 3.6333 | 0.24608
3500.4 | 701.05 | 5.4833 | 0.30231 | 3499.8 | 697.95 | 4.1167 | 0.26194
3600.4 | 721.15 | 5.7167 | 0.30867 | 3600 | 718 | 4.3333 | 0.26874
3700.4 | 741.25 | 6.8167 | 0.33706
3800.4 | 761.35 | 6.4167 | 0.32702
3900.4 | 781.45 | 7.1667 | 0.34561
4000.2 | 801.6 | 7.7833 | 0.36017

## QC4_Long_2
Test Date: | 2017-02-06 |  |  | Test Date: | 2017-02-06 |  |  | Test Date: | 2017-07-06
File: | GE11-X-L-CERN-0002-20170602.xls |  |  | File: | GE11-X-L-CERN-0002-20170607-Test1.xls |  |  | File: | GE11-X-L-CERN-0002-20170607-Test2.xls
Elog: | http://cmsonline.cern.ch/cms-elog/988188 |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/989356 |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/989426
Description: | Initial Test |  |  | Description: | After adding shielding foil on the HV area |  |  | Description: | After powering on for 3h (at 4kV)
Measured R_Equiv (Mohm) | 5 |  |  | Measured R_Equiv (Mohm) | 5 |  |  | Measured R_Equiv (Mohm) | 5
Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) |  | 140
Status: | Passed |  |  | Status: | Passed |  |  | Status: | Passed
Vmon | Imon | Rate | Rate Err | Vmon | Imon | Rate | Rate Err | Vmon | Imon | Rate | Rate Err
(V) | (uA) | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz)
198.8 | 39.5 | 0 | 0 | 198.8 | 39.45 | 0 | 0 | 199.2 | 39.55 | 0 | 0
399 | 79.3 | 0 | 0 | 399 | 79.3 | 0 | 0 | 399 | 79.35 | 0 | 0
599.2 | 119.1 | 0 | 0 | 599.2 | 119.15 | 0 | 0 | 599.2 | 119.2 | 0 | 0
799.2 | 158.95 | 0 | 0 | 799.4 | 159 | 0 | 0 | 799.4 | 159.05 | 0 | 0
999.4 | 198.85 | 0 | 0 | 999.4 | 198.95 | 0 | 0 | 999.4 | 199 | 0 | 0
1199.6 | 238.7 | 0.083333 | 0.037268 | 1199.6 | 238.85 | 0 | 0 | 1199.6 | 238.9 | 0 | 0
1399.6 | 278.55 | 0.46667 | 0.088192 | 1399.6 | 278.75 | 0.083333 | 0.037268 | 1399.6 | 278.8 | 0 | 0
1599.6 | 318.5 | 0.68333 | 0.10672 | 1599.6 | 318.75 | 0.18333 | 0.055277 | 1599.6 | 318.75 | 0 | 0
1799.6 | 358.35 | 1.0667 | 0.13333 | 1799.6 | 358.7 | 0.31667 | 0.072648 | 1799.6 | 358.75 | 0 | 0
1999.6 | 398.3 | 1.9667 | 0.18105 | 1999.6 | 398.65 | 0.58333 | 0.098601 | 1999.6 | 398.7 | 0 | 0
2199.8 | 438.25 | 3.2667 | 0.23333 | 2199.8 | 438.65 | 0.63333 | 0.10274 | 2199.8 | 438.7 | 0 | 0
2399.6 | 478.2 | 4.2333 | 0.26562 | 2399.6 | 478.65 | 0.8 | 0.11547 | 2399.6 | 478.65 | 0.016667 | 0.016667
2599.8 | 518.15 | 5.7667 | 0.31002 | 2599.8 | 518.65 | 1.0667 | 0.13333 | 2599.8 | 518.7 | 0.21667 | 0.060093
2799.8 | 558.05 | 6.25 | 0.32275 | 2799.8 | 558.65 | 1.6 | 0.1633 | 2800 | 558.7 | 0.36667 | 0.078174
2999.8 | 598.1 | 7.6333 | 0.35668 | 2999.8 | 598.65 | 2.4667 | 0.20276 | 2999.8 | 598.7 | 0.56667 | 0.097183
3099.8 | 618.05 | 8.3167 | 0.37231 | 3099.8 | 618.65 | 2.45 | 0.20207 | 3099.8 | 618.7 | 0.61667 | 0.10138
3199.8 | 638.05 | 8.8 | 0.38297 | 3199.8 | 638.6 | 2.6167 | 0.20883 | 3199.8 | 638.65 | 0.93333 | 0.12472
3299.8 | 658 | 9.55 | 0.39896 | 3299.8 | 658.65 | 2.9167 | 0.22048 | 3299.8 | 658.65 | 1.1833 | 0.14044
3399.8 | 678 | 10.2167 | 0.41265 | 3399.8 | 678.7 | 2.5667 | 0.20683 | 3399.8 | 678.7 | 1.3333 | 0.14907
3499.8 | 698.05 | 10.9167 | 0.42655 | 3500 | 698.65 | 3 | 0.22361 | 3500 | 698.7 | 1.8167 | 0.17401
3600 | 718.05 | 11.35 | 0.43493 | 3600 | 718.65 | 2.9333 | 0.22111 | 3600 | 718.75 | 1.9 | 0.17795
3700 | 738.1 | 12.9333 | 0.46428 | 3700 | 738.65 | 3.1167 | 0.22791 | 3700 | 738.7 | 2.2 | 0.19149
3800.2 | 758.1 | 14.5667 | 0.49272 | 3800.2 | 758.75 | 3.6167 | 0.24552 | 3800.2 | 758.75 | 2.4333 | 0.20138
3900.2 | 778.15 | 14.5167 | 0.49188 | 3900.2 | 778.75 | 3.45 | 0.23979 | 3900.2 | 778.8 | 2.8333 | 0.21731
4000.2 | 798.2 | 15.9 | 0.51478 | 4000.2 | 798.8 | 3.7167 | 0.24889 | 4000.2 | 798.8 | 2.9333 | 0.22111
4100.2 | 818.25 | 16.85 | 0.52994 | 4100.2 | 818.85 | 3.9333 | 0.25604 | 4100.2 | 818.85 | 3.25 | 0.23274
4200.2 | 838.35 | 18 | 0.54772 | 4200.2 | 838.95 | 4.3667 | 0.26977 | 4200.2 | 838.85 | 3.5 | 0.24152
4300.4 | 858.45 | 18.8833 | 0.561 | 4300.4 | 858.95 | 4.7833 | 0.28235 | 4300.4 | 858.95 | 3.2167 | 0.23154
4400.2 | 878.6 | 20 | 0.57735 | 4400.4 | 879 | 4.5833 | 0.27639 | 4400.2 | 879 | 4.1333 | 0.26247
4500.6 | 898.75 | 21.5667 | 0.59954 | 4500.6 | 899.2 | 4.9833 | 0.28819 | 4500.6 | 899.05 | 4.15 | 0.263
4600.4 | 918.95 | 23.2667 | 0.62272 | 4600.4 | 919.15 | 5.35 | 0.29861 | 4600.4 | 919.15 | 4.7833 | 0.28235
4700.6 | 939.15 | 20.9 | 0.5902 | 4700.6 | 939.25 | 5.6667 | 0.30732 | 4700.6 | 939.25 | 5.2333 | 0.29533
4800.6 | 959.45 | 27.8833 | 0.68171 | 4800.6 | 959.5 | 8.5167 | 0.37676 | 4800.6 | 959.35 | 7.2833 | 0.34841
4900.6 | 979.75 | 49.7167 | 0.91028 | 4900.2 | 979.55 | 11.0833 | 0.42979 | 4900.6 | 979.45 | 9.7667 | 0.40346

## QC4_Long_3
Test Date: | 2017-12-07 |  |  | Test Date: | 07/13/2017 |  |  | Test Date: | 07/13/2017 |  |  | Test Date: | 2017-01-08 |  |  | Test Date: | 2017-01-08
File: | GE11-X-L-CERN-0004_QC4_20170712_Test1.xlsx |  |  | File: | GE11-X-L-CERN-0004_QC4_20170713_Test2.xlsx |  |  | File: | GE11-X-L-CERN-0003_QC4_20170713_Test2.xlsx |  |  | File: | GE11-X-L-CERN-0003_QC4_20170801_Test1.xlsx |  |  | File: | GE11-X-L-CERN-0003_QC4_20170801_Test2.xlsx
Elog: |  |  |  | Elog: |  |  |  | Elog: |  |  |  | Elog: |  |  |  | Elog:
Description: | Initial Test |  |  | Description: | After powering on the detector at 4kV overnight. |  |  | Description: | After replacing the faulty screws near the short base of the drift board in the clean room |  |  | Description: | After fixing the RO in the clean room |  |  | Description: | After powering on the detector at 4kV for one hour
Measured R_Equiv (Mohm) | 5 |  |  | Measured R_Equiv (Mohm) | 5 |  |  | Measured R_Equiv (Mohm) | 5 |  |  | Measured R_Equiv (Mohm) | 5 |  |  | Measured R_Equiv (Mohm) | 5
Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) |  | 140
Status: | Passed |  |  | Status: | Passed |  |  | Status: | Passed |  |  | Status: | Passed |  |  | Status: | Passed
Vmon | Imon | Rate | Rate Err | Vmon | Imon | Rate | Rate Err | Vmon | Imon | Rate | Rate Err | Vmon | Imon | Rate | Rate Err | Vmon | Imon | Rate | Rate Err
(V) | (uA) | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz)
199.4 | 39.7 | 0 | 0 | 199.6 | 39.8 | 0.016667 | 0.016667 | 199.4 | 39.8 | 0.016667 | 0.016667 | 199.8 | 39.8 | 0 | 0 | 199.6 | 39.8 | 0 | 0
399.6 | 79.7 | 0 | 0 | 399.4 | 79.8 | 0.016667 | 0.016667 | 399.4 | 79.8 | 0 | 0 | 399.4 | 79.75 | 0 | 0 | 399.6 | 79.75 | 0 | 0
599.6 | 119.7 | 0 | 0 | 599.6 | 119.8 | 0 | 0 | 599.6 | 119.8 | 0 | 0 | 599.6 | 119.75 | 0 | 0 | 599.6 | 119.75 | 0 | 0
799.8 | 159.7 | 0.016667 | 0.016667 | 799.8 | 159.85 | 0 | 0 | 799.8 | 159.85 | 0 | 0 | 799.8 | 159.75 | 0 | 0 | 799.8 | 159.8 | 0 | 0
999.8 | 199.8 | 0 | 0 | 999.8 | 199.95 | 0 | 0 | 999.8 | 200 | 0 | 0 | 999.8 | 199.9 | 0 | 0 | 999.8 | 199.95 | 0 | 0
1200 | 239.8 | 0.016667 | 0.016667 | 1199.8 | 240.05 | 0 | 0 | 1199.8 | 240.1 | 0 | 0 | 1199.8 | 240 | 0 | 0 | 1199.8 | 240.05 | 0 | 0
1399.8 | 279.9 | 0.4 | 0.08165 | 1399.8 | 280.15 | 0 | 0 | 1399.8 | 280.2 | 0 | 0 | 1399.8 | 280.05 | 0 | 0 | 1399.8 | 280.15 | 0 | 0
1599.8 | 320 | 1.85 | 0.17559 | 1599.8 | 320.3 | 0.016667 | 0.016667 | 1599.8 | 320.35 | 0 | 0 | 1599.8 | 320.2 | 0 | 0 | 1599.8 | 320.25 | 0 | 0
1799.8 | 360.15 | 3.3 | 0.23452 | 1799.6 | 360.5 | 0 | 0 | 1799.8 | 360.55 | 0 | 0 | 1799.8 | 360.4 | 0.016667 | 0.016667 | 1799.8 | 360.5 | 0 | 0
1999.8 | 400.3 | 4.3 | 0.26771 | 1999.8 | 400.7 | 0 | 0 | 1999.8 | 400.75 | 0 | 0 | 1999.8 | 400.6 | 0.21667 | 0.060093 | 1999.8 | 400.7 | 0.016667 | 0.016667
2199.8 | 440.45 | 5.9333 | 0.31447 | 2199.8 | 440.9 | 0 | 0 | 2199.8 | 440.95 | 0.033333 | 0.02357 | 2199.8 | 440.8 | 0.58333 | 0.098601 | 2199.8 | 440.9 | 0.1 | 0.040825
2399.8 | 480.6 | 7.4667 | 0.35277 | 2399.8 | 481.15 | 0 | 0 | 2399.8 | 481.15 | 0 | 0 | 2399.8 | 481 | 0.95 | 0.12583 | 2399.8 | 481.15 | 0.13333 | 0.04714
2600 | 520.85 | 9.2167 | 0.39193 | 2600 | 521.35 | 0 | 0 | 2600 | 521.4 | 0 | 0 | 2600 | 521.25 | 1.7167 | 0.16915 | 2600 | 521.4 | 0.41667 | 0.083333
2800 | 561.1 | 12.15 | 0.45 | 2800 | 561.7 | 0.083333 | 0.037268 | 2800.2 | 561.7 | 0.1 | 0.040825 | 2800 | 561.6 | 2.4 | 0.2 | 2800.2 | 561.75 | 0.8 | 0.11547
3000 | 601.3 | 14 | 0.48305 | 3000 | 601.95 | 0.23333 | 0.062361 | 3000.2 | 601.95 | 0.26667 | 0.066667 | 3000 | 601.8 | 3.4333 | 0.23921 | 3000.2 | 602 | 1.4 | 0.15275
3100.2 | 621.4 | 13.6667 | 0.47726 | 3100 | 622.05 | 0.4 | 0.08165 | 3100.2 | 622.05 | 0.26667 | 0.066667 | 3100.2 | 621.95 | 3.6833 | 0.24777 | 3100.2 | 622.15 | 1.6667 | 0.16667
3200 | 641.5 | 13.5667 | 0.47551 | 3200 | 642.15 | 0.55 | 0.095743 | 3200 | 642.15 | 0.4 | 0.08165 | 3200 | 642.05 | 3.4667 | 0.24037 | 3200 | 642.25 | 1.9667 | 0.18105
3300 | 661.6 | 13.35 | 0.4717 | 3300 | 662.25 | 0.6 | 0.1 | 3300 | 662.25 | 0.68333 | 0.10672 | 3300 | 662.2 | 3.65 | 0.24664 | 3300 | 662.4 | 2.2 | 0.19149
3400 | 681.75 | 13.2333 | 0.46963 | 3400 | 682.4 | 1.1667 | 0.13944 | 3400 | 682.4 | 0.8 | 0.11547 | 3400 | 682.35 | 4.0833 | 0.26087 | 3400 | 682.6 | 2.5 | 0.20412
3500.2 | 701.95 | 13.1333 | 0.46786 | 3500.2 | 702.6 | 1 | 0.1291 | 3500.2 | 702.6 | 1.1667 | 0.13944 | 3500.2 | 702.5 | 4.7333 | 0.28087 | 3500.2 | 702.75 | 2.65 | 0.21016
3600.2 | 722.05 | 13.2667 | 0.47022 | 3600 | 722.7 | 1.2 | 0.14142 | 3600.2 | 722.7 | 1 | 0.1291 | 3600.2 | 722.65 | 4.85 | 0.28431 | 3600.2 | 722.9 | 2.9167 | 0.22048
3700.2 | 742.2 | 13.7833 | 0.47929 | 3700.2 | 742.85 | 1.4 | 0.15275 | 3700.2 | 742.85 | 1.4833 | 0.15723 | 3700.2 | 742.85 | 5.4167 | 0.30046 | 3700.2 | 743.1 | 2.9667 | 0.22236
3800.2 | 762.35 | 14.5167 | 0.49188 | 3800 | 763.05 | 1.65 | 0.16583 | 3800.2 | 763.05 | 1.3833 | 0.15184 | 3800.2 | 763 | 5.5333 | 0.30368 | 3800 | 763.3 | 3.05 | 0.22546
3900.2 | 782.55 | 15.65 | 0.51072 | 3900.2 | 783.2 | 1.6667 | 0.16667 | 3900.2 | 783.25 | 1.9167 | 0.17873 | 3900.2 | 783.2 | 5.7833 | 0.31047 | 3900.2 | 783.5 | 3.5667 | 0.24381
3999.8 | 802.75 | 18.85 | 0.56051 | 4000.2 | 803.35 | 1.9667 | 0.18105 | 4000.2 | 803.35 | 1.75 | 0.17078 | 4000.2 | 803.4 | 6.1333 | 0.31972 | 4000.2 | 803.65 | 3.85 | 0.25331
4100.2 | 822.85 | 17.4667 | 0.53955 | 4100.2 | 823.55 | 2.1667 | 0.19003 | 4100.2 | 823.55 | 2.1667 | 0.19003 | 4100.2 | 823.55 | 6.1833 | 0.32102 | 4100.2 | 823.9 | 4.2167 | 0.2651
4200.2 | 843.05 | 17.9 | 0.5462 | 4200.2 | 843.75 | 2.3667 | 0.19861 | 4200.2 | 843.75 | 2.15 | 0.1893 | 4200.2 | 843.8 | 6.3333 | 0.32489 | 4200.2 | 844.1 | 4.5667 | 0.27588
4299.8 | 863.3 | 20.7833 | 0.58855 | 4300.2 | 863.95 | 2.3667 | 0.19861 | 4300 | 863.95 | 2.9167 | 0.22048 | 4300.2 | 864.05 | 6.4 | 0.3266 | 4300.2 | 864.3 | 4.8167 | 0.28333
|  |  |  | 4400.2 | 884.15 | 3.0333 | 0.22485 | 4400.2 | 884.15 | 2.9333 | 0.22111 | 4400.2 | 884.3 | 6.7833 | 0.33624 | 4400.2 | 884.55 | 4.6667 | 0.27889
|  |  |  | 4500.2 | 904.4 | 3.2833 | 0.23393 | 4500.2 | 904.4 | 2.9333 | 0.22111 | 4500.2 | 904.6 | 9.9167 | 0.40654 | 4500.2 | 904.8 | 5.1 | 0.29155
|  |  |  | 4600 | 924.55 | 3.7333 | 0.24944 | 4600.2 | 924.6 | 3.6333 | 0.24608 | 4600 | 924.85 | 14.5833 | 0.49301 | 4600.2 | 925.05 | 6.3833 | 0.32617
|  |  |  | 4700 | 944.8 | 6.8667 | 0.3383 | 4700 | 944.85 | 6.05 | 0.31754 | 4700 | 945.2 | 28.8333 | 0.69322 | 4700 | 945.35 | 9.3667 | 0.39511
|  |  |  | 4800.2 | 965.1 | 12.65 | 0.45917 | 4799.8 | 965.15 | 12.1667 | 0.45031 | 4800.2 | 965.55 | 84.5167 | 1.1868 | 4800.2 | 965.65 | 14.85 | 0.49749
|  |  |  | 4900.2 | 985.3 | 19.85 | 0.57518 | 4900.2 | 985.5 | 16.8833 | 0.53046 |  |  |  |  | 4900.2 | 992.4 | 14.45 | 0.49075

## QC4_Long_4
Test Date: | 2017-07-27 |  |  | Test Date: | 2017-07-27
File: | GE11-X-L-CERN-0004-20172707-Test2.xls |  |  | File: | GE11-X-L-CERN-0004-20172707-Test2.xls
Elog: | http://cmsonline.cern.ch/cms-elog/989426 |  |  | Elog: | http://cmsonline.cern.ch/cms-elog/989426
Description: | After fixing a faulty screw Test 1 |  |  | Description: | After fixing a faulty screw Test 2
Measured R_Equiv (Mohm) | 5 |  |  | Measured R_Equiv (Mohm) | 5
Disc Thresh (mV) |  | 140 |  | Disc Thresh (mV) |  | 140
Status: | Passed |  |  | Status: | Passed
Vmon | Imon | Rate | Rate Err | Vmon | Imon | Rate | Rate Err
(V) | (uA) | (Hz) | (Hz) | (V) | (uA) | (Hz) | (Hz)
199.4 | 39.7 | 0 | 0 | 199.4 | 39.95 | 0.016667 | 0.016667
399.6 | 79.7 | 0.016667 | 0.016667 | 399.4 | 80.15 | 0 | 0
599.6 | 119.7 | 0.016667 | 0.016667 | 599.6 | 120.4 | 0 | 0
799.8 | 159.75 | 0.05 | 0.028868 | 799.8 | 160.6 | 0 | 0
999.8 | 199.85 | 0 | 0 | 999.8 | 200.9 | 0 | 0
1119.8 | 239.85 | 0 | 0 | 1199.8 | 241.2 | 0 | 0
1399.8 | 279.95 | 0.25 | 0.06455 | 1399.8 | 281.45 | 0 | 0
1599.8 | 320.05 | 0.46667 | 0.088192 | 1599.8 | 321.7 | 0.016667 | 0.016667
1799.8 | 360.2 | 0.76667 | 0.11304 | 1799.6 | 362 | 0.46667 | 0.088192
1999.8 | 400.35 | 1.1333 | 0.13744 | 1999.8 | 402.3 | 0.75 | 0.1118
2199.8 | 440.5 | 1.6 | 0.1633 | 2199.8 | 442.65 | 1.0833 | 0.13437
2399.8 | 480.7 | 2.3167 | 0.1965 | 2399.8 | 482.95 | 1.7833 | 0.1724
2600 | 520.95 | 2.8333 | 0.21731 | 2600 | 523.3 | 2.3167 | 0.1965
2800.2 | 561.15 | 3.7167 | 0.24889 | 2800 | 563.6 | 3.4833 | 0.24095
3000.2 | 601.4 | 3.1833 | 0.23034 | 3000.2 | 603.9 | 4.5167 | 0.27437
3100.2 | 621.5 | 3.9333 | 0.25604 | 3100.2 | 624.05 | 5.0167 | 0.28916
3200 | 641.6 | 5.0833 | 0.29107 | 3200 | 644.2 | 5.25 | 0.2958
3299.8 | 661.75 | 5.3167 | 0.29768 | 3300 | 664.4 | 5.35 | 0.29861
3400 | 681.85 | 5.75 | 0.30957 | 3400 | 684.55 | 5.7 | 0.30822
3500.2 | 702 | 6.2833 | 0.32361 | 3500.2 | 704.75 | 6.0333 | 0.3171
3600.2 | 722.15 | 6.6833 | 0.33375 | 3600.2 | 724.9 | 6.8 | 0.33665
3700.2 | 742.3 | 7.15 | 0.34521 | 3700.2 | 754.05 | 6.8667 | 0.3383
3800 | 762.45 | 8.95 | 0.38622 | 3800.2 | 765.2 | 6.9167 | 0.33953
3900.2 | 782.6 | 8.45 | 0.37528 | 3900.2 | 785.4 | 7.7167 | 0.35862
4000 | 802.8 | 9.7333 | 0.40277 | 4000.2 | 805.55 | 7.9833 | 0.36477
4100.2 | 822.95 | 9.4 | 0.39581 | 4100.2 | 825.75 | 7.75 | 0.3594
4200.2 | 843.1 | 9.3333 | 0.39441 | 4200.2 | 845.95 | 8.6833 | 0.38042
4300.2 | 863.25 | 9.8167 | 0.40449 | 4300.2 | 866.1 | 8.9833 | 0.38694
4400.2 | 888.3 | 9.8833 | 0.40586 | 4400.2 | 886.3 | 8.2667 | 0.37118
4500.2 | 908.6 | 10.65 | 0.42131 | 4500.2 | 906.6 | 8.8667 | 0.38442
4600 | 928.9 | 10.9833 | 0.42785 | 4599.8 | 927.05 | 10.8667 | 0.42557
4700 | 949.25 | 11.1833 | 0.43173 | 4700 | 947.15 | 10.7333 | 0.42295
4800.2 | 969.65 | 12.1333 | 0.44969 | 4800.2 | 967.4 | 11.7 | 0.44159
4900 | 990.05 | 21 | 0.59161 | 4900 | 988.3 | 21.5667 | 0.59954