## Data Summary
QC5 Effective gain
Username: | Christos Roskas | QC5 Template | V2 |  | Environment |  |  |  |  | For Measuring Rate |  |  |  | For Measuring Current
Amplifier Settings |  | Date | 2018-04-05 |  | Vmon | Imon | Time | Pressure | Temp | Source Off |  | Source On |  | Source Off |  | Source On
[10 empty] | Counts | Err | Counts | Err | Current | Err | Current | Err
Make & Model: | ORTEC 474 |  |  |  | (V) | (uA) | (HH:MM) | (mbar) | (Deg C) | (N) | (N) | (N) | (N) | (A) | (A) | (A) | (A)
Coarse Gain: | 4x |  |  | Ar/CO2:70/30 Range | 3600 | 720 | 10:25 | 993 | 20.4 | 191 | 13.8203 | 108530 | 329.4389 | -2.4088e-11 | 7.6611e-12 | -1.7561e-09 | 1.3439e-10
Fine Gain: | 4.5 |  |  |  | 3550 | 710 | 10:35 | 993 | 20.4 | 159 | 12.6095 | 103565 | 321.8152 | -1.5997e-11 | 4.231e-12 | -1.2262e-09 | 9.5664e-11
Diff Time (ns): | 100 |  |  |  | 3500 | 700 | 10:45 | 993 | 20.3 | 129 | 11.3578 | 97360 | 312.0256 | -1.2225e-11 | 3.4025e-12 | -8.775e-10 | 6.9002e-11
Int Time (ns): | 100 |  |  |  | 3450 | 690 | 10:55 | 993 | 20.3 | 99 | 9.9499 | 85752 | 292.8344 | -1.0238e-11 | 2.545e-12 | -6.1518e-10 | 4.7551e-11
Detector Settings |  |  |  |  | 3400 | 680 | 11:05 | 993 | 20.4 | 73 | 8.544 | 53846 | 232.0474 | -7.9804e-12 | 2.5856e-12 | -4.3111e-10 | 3.3603e-11
[5 empty] | 3350 | 670 | 11:15 | 993 | 20.3 | 37 | 6.0828 | 39482 | 198.7008 | -7.5416e-12 | 2.9761e-12 | -3.0469e-10 | 2.3527e-11
Serial Number: | GE1/1-X-L-GHENT-0006 |  |  |  | 3300 | 660 | 11:25 | 993 | 20.4 | 32 | 5.6569 | 23896 | 154.5833 | -6.1328e-12 | 1.481e-12 | -2.1321e-10 | 1.638e-11
Position (ieta,iphi): | 4,2 |  |  |  | 3250 | 650 | 11:35 | 992 | 20.3 | 13 | 3.6056 | 6751 | 82.1645 | -5.4125e-12 | 1.8298e-12 | -1.5228e-10 | 1.1294e-11
Gas (X/Y/Z): | Ar/CO2 |  |  |  | 3200 | 640 | 11:45 | 993 | 20.3 | 6 | 2.4495 | 506 | 22.4944 | -4.3836e-12 | 1.414e-12 | -1.0966e-10 | 7.9298e-12
Gas Frac (%X/%Y/%Z): | 70/30 |  |  |  | 3150 | 630 | 11:55 | 993 | 20.3 | 2 | 1.4142 | 13 | 3.6056 | -3.7684e-12 | 1.1363e-12 | -7.798e-11 | 6.0572e-12
Flow Rate (L/hr): | 5 |  |  |  | 3100 | 620 | 12:05 | 993 | 20.2 | 0 | 0 | 7 | 2.6458 | -3.5393e-12 | 1.4287e-12 | -5.6599e-11 | 4.5913e-12
HV Circuit R_Equiv (MOhm): | 5 |  |  |  | 3050 | 610 | 12:15 | 993 | 20.2 | 1 | 1 | 8 | 2.8284 | -3.6136e-12 | 1.7347e-12 | -4.1796e-11 | 3.1213e-12
Total HV Divider R (MOhm): | 4.7 |  |  |  | 3000 | 600 | 12:25 | 993 | 20.4 | 1 | 1 | 5 | 2.2361 | -3.2457e-12 | 3.446e-12 | -3.0508e-11 | 4.6055e-12
R_Drift (MOhm): | 1.126 |  |  |  | 2950 | 590 | 12:35 | 993 | 20.4 | 1 | 1 | 5 | 2.2361 | -2.4646e-12 | 1.7135e-12 | -2.2724e-11 | 2.4183e-12
R_GEM1 (MOhm): | 0.562 |  |  |  | 2900 | 580 | 12:45 | 993 | 20.3 | 1 | 1 | 6 | 2.4495 | -1.7291e-12 | 1.5056e-12 | -1.6393e-11 | 1.8856e-12
R_Transfer1 (MOhm): | 0.438 |  |  |  | 2850 | 570 | 13:05 | 99 | 20.3 | 0 | 0 | 2 | 1.4142 | -2.2019e-12 | 1.8807e-12 | -1.2751e-11 | 1.7551e-12
R_GEM2 (MOhm): | 0.55
R_Transfer2 (MOhm): | 0.875 |  |  |  |  | Fe55 | Cd109 | Cu X-Ray
R_GEM3 (MOhm): | 0.525 |  |  | N_primary | 346 | 196 | 322 | 346 | Electron Charge |  | 1.602e-19
R_Induction (MOhm): | 0.625 |  |  | Sigma N_primary | 2.9 | 1.8 | 2.8 | 2.9
Discriminator Settings |  |  |  | After T/P Corrections
Make & Model: | ORTEC 935 |  |  | Imon Equiv | V_mon | V_Drift | Rate | Rate Err | Current | Current Err | Gain | Gain Err
Threshold (mV): | -140 |  |  | (uA) | (V) | (V) | (Hz) | (Hz) | (A) | (A) | (A.U.) | (A.U.)
Walk Adjust (mV): | 23 |  |  | 731.9638 | 3659.819 | 3497.394 | 1805.65 | 5.4955 | -1.7321e-09 | 1.3461e-10 | 17305.7789 | 1353.7925
Width (ns): | 160 |  |  | 721.7976 | 3608.9881 | 3448.8191 | 1723.4333 | 5.3677 | -1.2102e-09 | 9.5757e-11 | 12091.9502 | 962.808
Picoammeter Settings |  |  |  | 711.6315 | 3558.1573 | 3400.2442 | 1620.5167 | 5.2039 | -8.6527e-10 | 6.9086e-11 | 8645.3209 | 694.5561
|  |  |  | 701.4653 | 3507.3265 | 3351.6692 | 1427.55 | 4.8834 | -6.0494e-10 | 4.7619e-11 | 6044.2513 | 478.8251
Make & Model: | KEITHLEY 6487 |  |  | 691.2991 | 3456.4957 | 3303.0943 | 896.2167 | 3.8701 | -4.2313e-10 | 3.3702e-11 | 4227.6992 | 338.8397
Triax Red: | SIG |  |  | 681.133 | 3405.6649 | 3254.5194 | 657.4167 | 3.3132 | -2.9715e-10 | 2.3715e-11 | 2968.968 | 238.4183
Triax Black: | GND |  |  | 670.9668 | 3354.8341 | 3205.9445 | 397.7333 | 2.5781 | -2.0708e-10 | 1.6447e-11 | 2069.0264 | 165.3621
Triax Green: | GND |  |  | 660.8006 | 3304.0032 | 3157.3696 | 112.3 | 1.3707 | -1.4686e-10 | 1.1441e-11 | 1467.3694 | 115.0597
Scalar Settings |  |  |  | 650.6345 | 3253.1724 | 3108.7947 | 8.3333 | 0.37712 | -1.0528e-10 | 8.0549e-12 | 1051.8526 | 81.025
|  |  |  | 640.4683 | 3202.3416 | 3060.2197 | 0.18333 | 0.06455 | -7.4212e-11 | 6.1628e-12 | 741.483 | 61.9294
|  |  |  | 630.3022 | 3151.5108 | 3011.6448 | 0.11667 | 0.044096 | -5.306e-11 | 4.8084e-12 | 530.1479 | 48.2752
Make & Model: | ORTEC 974 |  |  | 620.136 | 3100.68 | 2963.0699 | 0.11667 | 0.05 | -3.8182e-11 | 3.571e-12 | 381.4943 | 35.8409
Acquisition Time (s): | 60 |  |  | 609.9698 | 3049.8491 | 2914.495 | 0.066667 | 0.040825 | -2.7262e-11 | 5.7521e-12 | 272.3857 | 57.5226
Source Settings |  |  |  | 599.8037 | 2999.0183 | 2865.9201 | 0.066667 | 0.040825 | -2.0259e-11 | 2.9638e-12 | 202.4203 | 29.6676
|  |  |  | 589.6375 | 2948.1875 | 2817.3452 | 0.083333 | 0.044096 | -1.4664e-11 | 2.413e-12 | 146.5138 | 24.1445
Source: | Ag (Silver) |  |  | 579.4713 | 2897.3567 | 2768.7702 | 0.033333 | 0.02357 | -1.0549e-11 | 2.5725e-12 | 105.3961 | 25.7199
HV (kV): | 40
Current (uA): | 5 |  |  | Summary |  |  | Correction parameters (P5) |  |  |  | For Ana. Framwork
Activity (MBq): | Enter value |  |  | Avrg temperature (C) | 293.475 |  | P0 | 964.4 |  |  | 731.963793537848,17305.7788666297,1353.79246853117
Filter Status: | 9 Cu tape layers |  |  | Err temperature (C) | 0.066144 |  | T0 | 297.1 |  |  | 721.797629738711,12091.9501840252,962.808033410268
Collimator Status: | Removed |  |  | Avrg pressure (mbar) | 937.0625 | [5 empty] | 711.631465939574,8645.32087207451,694.556060073159
|  |  |  | Err pressure (mbar) | 216.3869 | [5 empty] | 701.465302140438,6044.25127690258,478.825100118735
|  |  |  | Expo_fit_1_param_1 | 3.5847e-07 | [5 empty] | 691.299138341301,4227.69917614315,338.839702330513
|  |  |  | Expo_fit_1_param_2 | 0.033552 | [5 empty] | 681.132974542164,2968.96797009394,238.418291506471
670.966810743027,2069.02636280336,165.362078616375
660.800646943891,1467.36935610685,115.059746102575
650.634483144754,1051.85261767029,81.025017295892
640.468319345617,741.482987503078,61.9294272439265
630.30215554648,530.147858864186,48.2751790581396
620.135991747343,381.494335137337,35.8409471908607
609.969827948207,272.385658746276,57.5225682046221
599.80366414907,202.42032243558,29.6675708912357
589.637500349933,146.513771247524,24.1444760222019
579.471336550796,105.396145808973,25.7198598501872

## 550uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-2.2019e-12 | 1.8807e-12 | -1.2751e-11 | 1.7551e-12
GEMs On |  | GEMs On
-1.3642e-12 | 0.31002 | -1.5689e-11 | 0.31202
-9.0949e-13 | 0.99506 | -1.1141e-11 | 0.99806
-2.0464e-12 | 1.3981 | -9.5497e-12 | 1.4031
-2.2737e-12 | 1.8031 | -1.7508e-11 | 1.8081
-2.5011e-12 | 2.2081 | -1.0687e-11 | 2.2341
-1.819e-12 | 2.6122 | -1.3642e-11 | 2.6452
-2.2737e-12 | 3.0172 | -1.1823e-11 | 3.0532
-1.3642e-12 | 3.4212 | -7.5033e-12 | 3.4592
-2.5011e-12 | 3.8272 | -1.5007e-11 | 3.8652
2.2737e-13 | 4.2312 | -1.5916e-11 | 4.2722
-6.8212e-13 | 4.6353 | -8.4128e-12 | 4.6783
-3.638e-12 | 5.0413 | -1.4779e-11 | 5.0843
-1.1369e-12 | 5.4463 | -1.2051e-11 | 5.4943
-2.2737e-12 | 5.8553 | -1.1369e-11 | 5.9003
-1.3642e-12 | 6.2604 | -1.7053e-11 | 6.3064
-3.4106e-12 | 6.6654 | -1.0914e-11 | 6.7124
-9.0949e-13 | 7.0704 | -7.9581e-12 | 7.1184
-1.1369e-12 | 7.4754 | -1.4325e-11 | 7.5254
-3.1832e-12 | 7.8815 | -1.0004e-11 | 7.9485
-2.2737e-12 | 8.2865 | -1.2051e-11 | 8.3785
-2.2737e-12 | 8.6915 | -1.4552e-11 | 8.7845
-1.1369e-12 | 9.0965 | -1.3642e-11 | 9.1905
-3.8654e-12 | 9.5015 | -1.0687e-11 | 9.6096
-4.5475e-13 | 9.9066 | -1.0687e-11 | 10.0176
-1.819e-12 | 10.3116 | -1.5007e-11 | 10.4236
-3.1832e-12 | 10.7176 | -1.387e-11 | 10.8296
-3.638e-12 | 11.1226 | -1.0914e-11 | 11.2356
0 | 11.5267 | -1.7053e-11 | 11.6427
-6.8212e-13 | 11.9297 | -9.5497e-12 | 12.0507
-9.0949e-13 | 12.3337 | -1.387e-11 | 12.4557
-2.7285e-12 | 12.7387 | -1.5234e-11 | 12.8667
-3.4106e-12 | 13.1438 | -1.4552e-11 | 13.2718
-1.3642e-12 | 13.5478 | -1.0687e-11 | 13.6768
-2.2737e-12 | 13.9528 | -1.3415e-11 | 14.0828
-3.4106e-12 | 14.3578 | -1.1369e-11 | 14.4888
-2.0464e-12 | 14.7628 | -1.3642e-11 | 14.9018
-2.0464e-12 | 15.1679 | -1.3415e-11 | 15.3089
2.2737e-13 | 15.5719 | -1.3415e-11 | 15.7139
2.2737e-13 | 15.9759 | -1.5916e-11 | 16.1199
-4.3201e-12 | 16.3799 | -1.2278e-11 | 16.526
-2.5011e-12 | 16.784 | -1.1141e-11 | 16.932
-2.0464e-12 | 17.189 | -1.0914e-11 | 17.337
-2.5011e-12 | 17.595 | -1.4779e-11 | 17.744
-2.5011e-12 | 18 | -1.0459e-11 | 18.15
-9.0949e-13 | 18.4041 | -1.4325e-11 | 18.5611
-4.5475e-13 | 18.8091 | -8.1855e-12 | 18.9891
2.2737e-13 | 19.2121 | -1.1823e-11 | 19.3931
-4.5475e-12 | 19.6171 | -1.5007e-11 | 19.8001
-4.3201e-12 | 20.0221 | -1.3642e-11 | 20.2052
6.8212e-13 | 20.4282 | -1.1369e-11 | 20.6112
0 | 20.8312 | -1.1141e-11 | 21.0162
-4.7748e-12 | 21.2342 | -1.2506e-11 | 21.4232
-3.4106e-12 | 21.6392 | -1.3415e-11 | 21.8283
-1.819e-12 | 22.0443 | -1.0459e-11 | 22.2343
0 | 22.4483 | -1.3188e-11 | 22.6403
1.1369e-12 | 22.8503 | -1.2051e-11 | 23.0463
-3.638e-12 | 23.2553 | -1.3415e-11 | 23.4623
-4.7748e-12 | 23.6604 | -1.0459e-11 | 23.8704
-1.1369e-12 | 24.0664 | -1.1141e-11 | 24.2764
-1.1369e-12 | 24.4694 | -1.4325e-11 | 24.6834
-3.4106e-12 | 24.8764 | -1.2051e-11 | 25.0894
-3.1832e-12 | 25.2804 | -1.1141e-11 | 25.4965
-1.1369e-12 | 25.6855 | -1.3188e-11 | 25.9015
-9.0949e-13 | 26.0905 | -1.0687e-11 | 26.3105
-5.0022e-12 | 26.4945 | -1.387e-11 | 26.7145
-2.7285e-12 | 26.9005 | -1.1141e-11 | 27.1245
1.5916e-12 | 27.3076 | -1.2051e-11 | 27.5306
-7.0486e-12 | 27.7126 | -1.4325e-11 | 27.9456
-6.1391e-12 | 28.1156 | -1.0687e-11 | 28.3506
-4.5475e-12 | 28.5206 | -1.5007e-11 | 28.7567
0 | 28.9246 | -1.2506e-11 | 29.1627
-6.8212e-13 | 29.3267 | -1.0004e-11 | 29.5667
-3.638e-12 | 29.7317 | -1.2278e-11 | 29.9737
-3.1832e-12 | 30.1377 | -1.5007e-11 | 30.3797
2.2737e-13 | 30.5428 | -1.2506e-11 | 30.7858
2.0464e-12 | 30.9468 | -1.2051e-11 | 31.1908
-4.5475e-12 | 31.3518 | -1.1141e-11 | 31.5968
-4.3201e-12 | 31.7568 | -1.1823e-11 | 32.0048
0 | 32.1608 | -1.4325e-11 | 32.4098
1.1369e-12 | 32.5619 | -1.3415e-11 | 32.8149
1.1369e-12 | 32.9659 | -1.1141e-11 | 33.2219
-3.8654e-12 | 33.3709 | -1.0687e-11 | 33.6469
-5.2296e-12 | 33.7759 | -1.6826e-11 | 34.0549
-6.8212e-13 | 34.181 | -1.296e-11 | 34.463
-6.8212e-13 | 34.586 | -1.2051e-11 | 34.87
-3.638e-12 | 34.99 | -1.1141e-11 | 35.276
-5.2296e-12 | 35.394 | -1.0687e-11 | 35.682
6.8212e-13 | 35.7991 | -1.0687e-11 | 36.0881
-6.8212e-13 | 36.2031 | -1.2506e-11 | 36.4931
-3.638e-12 | 36.6081 | -1.3188e-11 | 36.8991
-4.5475e-12 | 37.0131 | -1.4325e-11 | 37.3051
-9.0949e-13 | 37.4191 | -1.3415e-11 | 37.7172
-6.8212e-13 | 37.8232 | -1.3642e-11 | 38.1222
-2.7285e-12 | 38.2272 | -1.0459e-11 | 38.5272
-3.8654e-12 | 38.6332 | -1.296e-11 | 38.9332
-1.1369e-12 | 39.0382 | -1.4779e-11 | 39.3392
-1.1369e-12 | 39.4433 | -1.2051e-11 | 39.7483
-5.2296e-12 | 39.8473 | -1.387e-11 | 40.1543
-2.5011e-12 | 40.2533 | -1.4552e-11 | 40.5593
-9.0949e-13 | 40.6583 | -1.3188e-11 | 40.9653
1.3642e-12 | 41.0624 | -1.1823e-11 | 41.3704
-1.1369e-12 | 41.4654 | -1.3642e-11 | 41.7754
-5.2296e-12 | 41.8694 | -1.296e-11 | 42.1804
-1.3642e-12 | 42.2744 | -1.2506e-11 | 42.5844
1.1369e-12 | 42.6794 | -1.296e-11 | 42.9895
-2.0464e-12 | 43.0835 | -1.3642e-11 | 43.3945
-3.1832e-12 | 43.4885 | -1.2506e-11 | 43.8005
-4.5475e-12 | 43.8935 | -1.0459e-11 | 44.2095
-9.0949e-13 | 44.2965 | -1.0687e-11 | 44.6305
-4.5475e-13 | 44.6996 | -1.1823e-11 | 45.0356
-9.0949e-13 | 45.1036 | -1.4325e-11 | 45.4396
-4.3201e-12 | 45.5066 | -9.7771e-12 | 45.8446
-2.7285e-12 | 45.9106 | -1.0459e-11 | 46.2507
-1.3642e-12 | 46.3156 | -1.5007e-11 | 46.6567
-4.5475e-13 | 46.7207 | -1.1823e-11 | 47.0647
-2.0464e-12 | 47.1257 | -1.1141e-11 | 47.4707
-5.0022e-12 | 47.5307 | -1.387e-11 | 47.8767
-3.4106e-12 | 47.9347 | -1.3642e-11 | 48.2818
0 | 48.3408 | -1.1141e-11 | 48.6898
0 | 48.7438 | -1.3188e-11 | 49.0948
-1.3642e-12 | 49.1438 | -1.3188e-11 | 49.5018
-2.2737e-12 | 49.5488 | -1.1141e-11 | 49.9079
-1.819e-12 | 49.9529 | -1.2051e-11 | 50.3119
-6.8212e-13 | 50.3579 | -1.5007e-11 | 50.7179
-3.1832e-12 | 50.7619 | -1.3642e-11 | 51.1219
-1.819e-12 | 51.1669 | -1.0914e-11 | 51.526
-3.1832e-12 | 51.572 | -9.3223e-12 | 51.931
-2.2737e-12 | 51.977 | -1.3415e-11 | 52.337
-1.819e-12 | 52.382 | -1.387e-11 | 52.742
-1.819e-12 | 52.786 | -1.2278e-11 | 53.148
-1.819e-12 | 53.191 | -1.3188e-11 | 53.5531
-6.8212e-13 | 53.5951 | -1.5234e-11 | 53.9581
-9.0949e-13 | 53.9991 | -1.3188e-11 | 54.3641
-1.819e-12 | 54.4031 | -1.4552e-11 | 54.7741
0 | 54.8101 | -1.2506e-11 | 55.1802
-1.3642e-12 | 55.2152 | -1.296e-11 | 55.5852
-6.8212e-13 | 55.6192 | -1.296e-11 | 55.9942
-3.4106e-12 | 56.0242 | -1.3642e-11 | 56.3992
-3.1832e-12 | 56.4292 | -1.1141e-11 | 56.8043
-4.3201e-12 | 56.8353 | -1.2278e-11 | 57.2083
2.2737e-13 | 57.2403 | -1.296e-11 | 57.6133
-3.638e-12 | 57.6443 | -1.387e-11 | 58.0183
-2.2737e-12 | 58.0503 | -1.2506e-11 | 58.4233
2.2737e-13 | 58.4553 | -1.2278e-11 | 58.8284
-3.1832e-12 | 58.8584 | -1.4552e-11 | 59.2334
-3.4106e-12 | 59.2644 | -1.5689e-11 | 59.6374
-3.4106e-12 | 59.6694 | -1.2051e-11 | 60.0424
-3.1832e-12 | 60.0744 | -1.1823e-11 | 60.4475
2.2737e-13 | 60.4785 | -1.0459e-11 | 60.8515
-9.0949e-13 | 60.8825 | -1.1823e-11 | 61.2585
-5.9117e-12 | 61.2865 | -1.2278e-11 | 61.6635
-1.3642e-12 | 61.6905 | -1.3642e-11 | 62.0695
-9.0949e-13 | 62.0956 | -1.3188e-11 | 62.4746
2.2737e-13 | 62.4986 | -1.387e-11 | 62.8806
-5.6843e-12 | 62.9026 | -1.387e-11 | 63.2856
-1.819e-12 | 63.3086 | -1.1369e-11 | 63.6906
-1.3642e-12 | 63.7136 | -1.4779e-11 | 64.0957
-1.1369e-12 | 64.1187 | -1.387e-11 | 64.5007
-2.7285e-12 | 64.5227 | -1.5007e-11 | 64.9067
-2.7285e-12 | 64.9267 | -1.387e-11 | 65.3117
-2.5011e-12 | 65.3307 | -1.2051e-11 | 65.7168
-2.7285e-12 | 65.7348 | -1.2506e-11 | 66.1228
-2.0464e-12 | 66.1388 | -1.296e-11 | 66.5278
-4.3201e-12 | 66.5448 | -1.5007e-11 | 66.9328
-2.2737e-12 | 66.9508 | -1.1823e-11 | 67.3368
-2.7285e-12 | 67.3559 | -1.0459e-11 | 67.7439
-3.8654e-12 | 67.7599 | -1.1823e-11 | 68.1489
-4.5475e-12 | 68.1659 | -1.3188e-11 | 68.5539
-9.0949e-13 | 68.5709 | -1.5234e-11 | 68.9589
-4.5475e-12 | 68.975 | -1.2051e-11 | 69.365
-2.5011e-12 | 69.378 | -1.4779e-11 | 69.77
-4.5475e-12 | 69.783 | -1.0459e-11 | 70.176
-4.5475e-12 | 70.187 | -1.387e-11 | 70.581
-3.1832e-12 | 70.592 | -1.0459e-11 | 70.9861
-9.0949e-13 | 70.9971 | -1.4552e-11 | 71.3911
-2.0464e-12 | 71.4001 | -1.4779e-11 | 71.7971
-6.1391e-12 | 71.8061 | -1.296e-11 | 72.2021
-2.0464e-12 | 72.2101 | -1.296e-11 | 72.6072
-3.1832e-12 | 72.6151 | -1.2278e-11 | 73.0132
-3.1832e-12 | 73.0192 | -1.4325e-11 | 73.4182
-5.0022e-12 | 73.4242 | -1.296e-11 | 73.8232
4.5475e-13 | 73.8292 | -9.7771e-12 | 74.2292
-9.0949e-13 | 74.2332 | -1.3642e-11 | 74.6333
-9.0949e-13 | 74.6373 | -1.4779e-11 | 75.0393
-3.4106e-12 | 75.0413 | -1.4779e-11 | 75.4443
-5.0022e-12 | 75.4463 | -1.3642e-11 | 75.8503
-2.7285e-12 | 75.8503 | -1.5234e-11 | 76.2544
4.5475e-13 | 76.2544 | -1.1369e-11 | 76.6594
-2.7285e-12 | 76.6584 | -1.4552e-11 | 77.0634
-8.6402e-12 | 77.0634 | -1.2051e-11 | 77.4704
-4.7748e-12 | 77.4684 | -1.5007e-11 | 77.8755
6.8212e-13 | 77.8735 | -1.3642e-11 | 78.2805
-9.0949e-13 | 78.2775 | -1.1141e-11 | 78.6865
-3.1832e-12 | 78.6805 | -1.2051e-11 | 79.0915
-5.9117e-12 | 79.0855 | -1.296e-11 | 79.4965
-1.3642e-12 | 79.4895 | -1.2506e-11 | 79.9016
4.5475e-13 | 79.8936 | -1.2278e-11 | 80.3076
-2.5011e-12 | 80.2976 | -1.2506e-11 | 80.7126
-4.5475e-12 | 80.7026 | -1.3415e-11 | 81.1176
-2.2737e-12 | 81.1076 | -1.2278e-11 | 81.5227
-9.0949e-13 | 81.5137 | -1.3415e-11 | 81.9287
-1.1369e-12 | 81.9177 | -1.4325e-11 | 82.3347
-8.4128e-12 | 82.3217 | -1.2278e-11 | 82.7397
-2.0464e-12 | 82.7277 | -1.296e-11 | 83.1448
-1.3642e-12 | 83.1317 | -1.4552e-11 | 83.5508
-1.1369e-12 | 83.5358 | -1.3188e-11 | 83.9548
-4.5475e-12 | 83.9418
-3.1832e-12 | 84.3468
1.1369e-12 | 84.7519

## 560uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.7291e-12 | 1.5056e-12 | -1.6393e-11 | 1.8856e-12
GEMs On |  | GEMs On
-2.5011e-12 | 0.31102 | -1.728e-11 | 0.31302
-1.3642e-12 | 0.99406 | -1.2278e-11 | 0.99806
-1.819e-12 | 1.4041 | -1.6598e-11 | 1.4041
-2.0464e-12 | 1.8111 | -1.6144e-11 | 1.8091
-2.0464e-12 | 2.2191 | -1.8872e-11 | 2.2151
0 | 2.6242 | -1.387e-11 | 2.6191
-2.2737e-12 | 3.0282 | -1.3415e-11 | 3.0242
-9.0949e-13 | 3.4332 | -1.8417e-11 | 3.4302
-1.3642e-12 | 3.8422 | -1.5916e-11 | 3.8352
-3.4106e-12 | 4.2462 | -1.5461e-11 | 4.2402
-1.819e-12 | 4.6563 | -1.6371e-11 | 4.6453
-3.1832e-12 | 5.0633 | -1.6826e-11 | 5.0513
-2.2737e-12 | 5.4673 | -1.4097e-11 | 5.4573
-4.3201e-12 | 5.8713 | -1.6371e-11 | 5.8633
-4.5475e-13 | 6.2784 | -1.7508e-11 | 6.2684
0 | 6.6844 | -1.8645e-11 | 6.6744
-6.8212e-13 | 7.0864 | -1.6826e-11 | 7.0794
-2.0464e-12 | 7.4904 | -1.9099e-11 | 7.4864
6.8212e-13 | 7.8945 | -1.4097e-11 | 7.9175
-2.5011e-12 | 8.2985 | -1.7508e-11 | 8.3245
-2.0464e-12 | 8.7045 | -1.8645e-11 | 8.7295
-6.8212e-13 | 9.1095 | -1.6826e-11 | 9.1355
-1.819e-12 | 9.5125 | -1.7735e-11 | 9.5435
-1.819e-12 | 9.9176 | -2.0236e-11 | 9.9496
-2.0464e-12 | 10.3216 | -1.387e-11 | 10.3566
-4.5475e-13 | 10.7266 | -1.3642e-11 | 10.7666
2.2737e-13 | 11.1316 | -1.9782e-11 | 11.1716
-1.3642e-12 | 11.5357 | -1.6598e-11 | 11.5767
-1.819e-12 | 11.9407 | -1.4097e-11 | 11.9807
-1.1369e-12 | 12.3457 | -1.728e-11 | 12.3867
-6.8212e-13 | 12.7507 | -2.0009e-11 | 12.7927
-6.8212e-13 | 13.1547 | -1.2733e-11 | 13.1988
-3.1832e-12 | 13.5588 | -1.9782e-11 | 13.6098
-1.819e-12 | 13.9628 | -1.2278e-11 | 14.0568
6.8212e-13 | 14.3688 | -1.9099e-11 | 14.4648
-4.5475e-13 | 14.7718 | -1.6598e-11 | 14.8768
-3.638e-12 | 15.1769 | -1.6826e-11 | 15.2889
-4.3201e-12 | 15.5819 | -1.3642e-11 | 15.6959
-2.0464e-12 | 15.9859 | -1.7735e-11 | 16.1019
-1.3642e-12 | 16.3909 | -1.9782e-11 | 16.5059
-2.0464e-12 | 16.796 | -1.8645e-11 | 16.916
-3.1832e-12 | 17.202 | -1.9327e-11 | 17.322
-9.0949e-13 | 17.608 | -1.5007e-11 | 17.732
1.3642e-12 | 18.012 | -1.296e-11 | 18.14
-4.7748e-12 | 18.417 | -1.8872e-11 | 18.5451
-2.5011e-12 | 18.8221 | -1.6144e-11 | 18.9581
-1.1369e-12 | 19.2261 | -1.728e-11 | 19.3651
-3.8654e-12 | 19.6301 | -1.6371e-11 | 19.7701
-5.2296e-12 | 20.0402 | -1.4552e-11 | 20.1742
-2.2737e-12 | 20.4462 | -1.6144e-11 | 20.5802
-1.1369e-12 | 20.8492 | -1.7508e-11 | 20.9862
-2.0464e-12 | 21.2682 | -1.5916e-11 | 21.3922
-2.0464e-12 | 21.6722 | -1.296e-11 | 21.7972
-3.1832e-12 | 22.0783 | -1.8417e-11 | 22.2033
-1.819e-12 | 22.4823 | -1.728e-11 | 22.6083
-3.1832e-12 | 22.8973 | -1.7508e-11 | 23.0383
-3.4106e-12 | 23.3013 | -1.3642e-11 | 23.4493
0 | 23.7064 | -1.6598e-11 | 23.8574
1.819e-12 | 24.1094 | -1.5234e-11 | 24.2754
-2.2737e-12 | 24.5144 | -1.728e-11 | 24.6824
-3.638e-12 | 24.9184 | -1.7508e-11 | 25.0884
-2.2737e-12 | 25.3224 | -1.6598e-11 | 25.4945
2.2737e-13 | 25.7275 | -1.5234e-11 | 25.9085
0 | 26.1305 | -1.5916e-11 | 26.3145
-9.0949e-13 | 26.5325 | -1.6826e-11 | 26.7205
-3.4106e-12 | 26.9375 | -1.387e-11 | 27.1266
-1.819e-12 | 27.3426 | -1.8645e-11 | 27.5456
1.3642e-12 | 27.7476 | -1.5916e-11 | 27.9526
-2.2737e-12 | 28.1516 | -1.728e-11 | 28.3576
-4.7748e-12 | 28.5566 | -1.8872e-11 | 28.7636
-3.4106e-12 | 28.9617 | -1.4552e-11 | 29.1697
1.3642e-12 | 29.3657 | -1.2051e-11 | 29.5747
-1.3642e-12 | 29.7707 | -1.6826e-11 | 29.9807
-5.9117e-12 | 30.1757 | -1.5007e-11 | 30.3907
-6.8212e-13 | 30.5817 | -1.6371e-11 | 30.7988
-4.3201e-12 | 30.9918 | -1.9782e-11 | 31.2038
-4.3201e-12 | 31.3978 | -1.4779e-11 | 31.6078
-1.1369e-12 | 31.8018 | -1.5234e-11 | 32.0138
-9.0949e-13 | 32.2068 | -1.8645e-11 | 32.4198
-3.1832e-12 | 32.6119 | -1.6144e-11 | 32.8259
-4.3201e-12 | 33.0169 | -1.4552e-11 | 33.2319
-1.819e-12 | 33.4209 | -1.6826e-11 | 33.6369
6.8212e-13 | 33.8249 | -1.9099e-11 | 34.044
-3.4106e-12 | 34.232 | -1.296e-11 | 34.449
-2.0464e-12 | 34.639 | -1.728e-11 | 34.855
-1.819e-12 | 35.046 | -1.5916e-11 | 35.261
-2.0464e-12 | 35.453 | -1.6598e-11 | 35.667
-9.0949e-13 | 35.8591 | -1.6826e-11 | 36.0951
-6.8212e-13 | 36.2671 | -1.6371e-11 | 36.5041
-1.1369e-12 | 36.6741 | -1.6826e-11 | 36.9101
-4.5475e-13 | 37.0791 | -1.6144e-11 | 37.3151
-6.8212e-13 | 37.4871 | -1.6598e-11 | 37.7262
-2.0464e-12 | 37.8912 | -1.6826e-11 | 38.1382
-2.2737e-12 | 38.2972 | -1.6371e-11 | 38.5432
-1.3642e-12 | 38.7032 | -1.8872e-11 | 38.9502
2.0464e-12 | 39.1082 | -1.6598e-11 | 39.3783
-2.0464e-12 | 39.5143 | -1.6826e-11 | 39.7853
-2.2737e-12 | 39.9193 | -1.5461e-11 | 40.1903
0 | 40.3253 | -1.7508e-11 | 40.5963
2.2737e-12 | 40.7263 | -1.4779e-11 | 41.0014
-3.1832e-12 | 41.1324 | -1.6371e-11 | 41.4084
-2.0464e-12 | 41.5374 | -1.8417e-11 | 41.8134
-6.8212e-13 | 41.9434 | -1.4097e-11 | 42.2194
-1.819e-12 | 42.3474 | -1.5234e-11 | 42.6254
-3.4106e-12 | 42.7535 | -1.6371e-11 | 43.0315
-2.9559e-12 | 43.1595 | -1.8645e-11 | 43.4365
-1.819e-12 | 43.5685 | -1.5007e-11 | 43.8415
-1.1369e-12 | 43.9725 | -1.3642e-11 | 44.2515
-1.819e-12 | 44.3775 | -1.5461e-11 | 44.6576
-2.0464e-12 | 44.7846 | -1.5234e-11 | 45.0646
-1.3642e-12 | 45.1916 | -1.6826e-11 | 45.4716
-1.1369e-12 | 45.5966 | -1.5234e-11 | 45.8776
2.2737e-13 | 46.0046 | -1.4779e-11 | 46.2826
-3.8654e-12 | 46.4117 | -1.5234e-11 | 46.6897
-2.5011e-12 | 46.8167 | -1.5916e-11 | 47.0957
-1.1369e-12 | 47.2217 | -1.6371e-11 | 47.5007
-9.0949e-13 | 47.6297 | -1.7508e-11 | 47.9067
-4.3201e-12 | 48.0348 | -1.5461e-11 | 48.3128
-4.5475e-13 | 48.4388 | -1.5916e-11 | 48.8828
0 | 48.8438 | -1.7963e-11 | 49.3438
0 | 49.2448 | -1.5461e-11 | 49.7529
-9.0949e-13 | 49.6468 | -1.8417e-11 | 50.1659
-2.0464e-12 | 50.0519 | -1.7735e-11 | 50.5719
-2.5011e-12 | 50.4559 | -1.7963e-11 | 50.9779
-9.0949e-13 | 51.1169 | -1.5461e-11 | 51.3869
-2.5011e-12 | 51.5209 | -1.7963e-11 | 51.792
-2.5011e-12 | 51.932 | -1.9327e-11 | 52.218
-6.8212e-13 | 52.337 | -1.5916e-11 | 52.624
0 | 52.741 | -1.5461e-11 | 53.029
-9.0949e-13 | 53.144 | -1.7963e-11 | 53.4351
-2.5011e-12 | 53.5481 | -1.4097e-11 | 53.8401
-2.5011e-12 | 53.9531 | -1.728e-11 | 54.2461
-4.5475e-13 | 54.3591 | -1.7963e-11 | 54.6511
-1.3642e-12 | 54.7631 | -1.6371e-11 | 55.0562
-1.3642e-12 | 55.1672 | -1.6598e-11 | 55.4722
-4.5475e-12 | 55.5712 | -1.4097e-11 | 55.8772
-2.9559e-12 | 55.9762 | -1.6144e-11 | 56.2832
-9.0949e-13 | 56.3822 | -1.7508e-11 | 56.6912
-2.5011e-12 | 56.7863 | -1.7963e-11 | 57.0983
-3.8654e-12 | 57.1913 | -1.3642e-11 | 57.5033
-2.0464e-12 | 57.5943 | -1.5007e-11 | 57.9093
0 | 58.0003 | -1.3415e-11 | 58.3163
-4.5475e-13 | 58.4013 | -1.7508e-11 | 58.7344
-2.5011e-12 | 58.8074 | -1.728e-11 | 59.1394
-5.0022e-12 | 59.2124 | -1.7963e-11 | 59.5484
-1.819e-12 | 59.6164 | -1.8417e-11 | 59.9534
9.0949e-13 | 60.0214 | -1.6826e-11 | 60.3595
0 | 60.4245 | -1.8645e-11 | 60.7665
-1.3642e-12 | 60.8265 | -1.5916e-11 | 61.1725
-4.3201e-12 | 61.2315 | -2.0009e-11 | 61.5785
4.5475e-13 | 61.6365 | -1.6826e-11 | 61.9845
1.3642e-12 | 62.0395 | -1.6826e-11 | 62.3946
-3.638e-12 | 62.4446 | -1.2506e-11 | 62.8016
-4.3201e-12 | 62.8476 | -1.5007e-11 | 63.2116
-3.4106e-12 | 63.2526 | -1.6826e-11 | 63.6176
2.2737e-13 | 63.6566 | -1.4097e-11 | 64.0237
4.5475e-13 | 64.0607 | -1.4779e-11 | 64.4277
-3.638e-12 | 64.4657 | -1.8417e-11 | 64.8347
-3.638e-12 | 64.8707 | -1.5916e-11 | 65.2417
-9.0949e-13 | 65.2767 | -1.728e-11 | 65.6478
-9.0949e-13 | 65.6808 | -1.9327e-11 | 66.0548
-4.5475e-13 | 66.0848 | -1.8645e-11 | 66.4608
-3.8654e-12 | 66.4878 | -1.387e-11 | 66.8668
-1.3642e-12 | 66.8928 | -1.728e-11 | 67.2749
0 | 67.2969 | -1.3415e-11 | 67.6819
-6.8212e-13 | 67.6989 | -1.8872e-11 | 68.0939
-2.2737e-12 | 68.1039 | -1.6826e-11 | 68.4999
-3.4106e-12 | 68.5089 | -1.3642e-11 | 68.9069
-1.3642e-12 | 68.9139 | -1.9099e-11 | 69.313
-6.8212e-13 | 69.32 | -2.1146e-11 | 69.717
-2.5011e-12 | 69.724 | -1.7508e-11 | 70.123
-1.3642e-12 | 70.13 | -1.4097e-11 | 70.529
2.2737e-13 | 70.535 | -1.296e-11 | 70.9481
-1.1369e-12 | 70.9391 | -1.4552e-11 | 71.4921
-4.3201e-12 | 71.3441 | -1.8417e-11 | 71.8971
-2.5011e-12 | 71.7501 | -1.4097e-11 | 72.3041
-2.9559e-12 | 72.1561 | -1.4552e-11 | 72.7092
0 | 72.5601 | -1.9782e-11 | 73.1152
0 | 72.9622 | -1.5007e-11 | 73.5202
-2.9559e-12 | 73.3652 | -1.4779e-11 | 73.9312
-2.2737e-12 | 73.7702 | -1.5461e-11 | 74.3363
-6.8212e-13 | 74.1752 | -1.4552e-11 | 74.7923
-9.0949e-13 | 74.5793 | -1.6826e-11 | 75.1973
0 | 74.9833 | -2.0918e-11 | 75.6033
-9.0949e-13 | 75.3853 | -1.3642e-11 | 76.0093
-3.638e-12 | 75.7883 | -1.6144e-11 | 76.4154
-4.5475e-13 | 76.1934 | -1.728e-11 | 76.8224
-2.9559e-12 | 76.5974 | -1.5916e-11 | 77.2294
-3.1832e-12 | 77.0024 | -1.5007e-11 | 77.6354
-3.4106e-12 | 77.4074 | -1.6371e-11 | 78.0405
-4.5475e-13 | 77.8124 | -1.6826e-11 | 78.4465
-3.4106e-12 | 78.2165 | -1.4779e-11 | 78.8515
-3.1832e-12 | 78.6205 | -1.6826e-11 | 79.2585
-9.0949e-13 | 79.0245 | -1.6144e-11 | 79.6636
-2.0464e-12 | 79.4285 | -1.5007e-11 | 80.0696
-6.8212e-13 | 79.8336 | -1.7735e-11 | 80.4786
-1.1369e-12 | 80.2366 | -1.7735e-11 | 80.8866
-2.0464e-12 | 80.6406 | -1.4779e-11 | 81.2926
-2.5011e-12 | 81.0476 | -1.728e-11 | 81.6977
-1.3642e-12 | 81.4527 | -1.6144e-11 | 82.1037
-9.0949e-13 | 81.8587 | -1.5916e-11 | 82.5077
-4.5475e-13 | 82.2627 | -1.7508e-11 | 82.9147
-3.1832e-12 | 82.6667 | -1.7508e-11 | 83.3198
-1.1369e-12 | 83.0708 | -1.6144e-11 | 83.7258
-1.819e-12 | 83.4768 | -1.2051e-11 | 84.1308

## 570uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-2.4646e-12 | 1.7135e-12 | -2.2724e-11 | 2.4183e-12
GEMs On |  | GEMs On
-6.8212e-13 | 0.31102 | -2.2965e-11 | 0.31202
-6.8212e-13 | 0.99506 | -1.9782e-11 | 0.99706
-3.4106e-12 | 1.3981 | -2.6148e-11 | 1.4031
-3.638e-12 | 1.8041 | -2.0464e-11 | 1.8081
-4.5475e-12 | 2.2091 | -2.0918e-11 | 2.2131
-6.8212e-13 | 2.6141 | -1.9782e-11 | 2.6181
1.5916e-12 | 3.0182 | -2.5921e-11 | 3.0242
-1.819e-12 | 3.4232 | -2.0236e-11 | 3.4292
-2.5011e-12 | 3.8282 | -2.0464e-11 | 3.8352
-9.0949e-13 | 4.2332 | -2.6148e-11 | 4.2412
-2.5011e-12 | 4.6373 | -2.5011e-11 | 4.6453
-3.4106e-12 | 5.0413 | -2.2965e-11 | 5.0503
-2.7285e-12 | 5.4463 | -2.2055e-11 | 5.4553
-9.0949e-13 | 5.8503 | -2.4102e-11 | 5.8603
-3.1832e-12 | 6.2534 | -2.1828e-11 | 6.2664
-2.5011e-12 | 6.6584 | -2.2055e-11 | 6.6724
-2.0464e-12 | 7.0674 | -2.0236e-11 | 7.0774
-6.8212e-13 | 7.4724 | -2.5921e-11 | 7.4824
-1.3642e-12 | 7.8755 | -2.4329e-11 | 7.8885
-3.8654e-12 | 8.2795 | -2.3647e-11 | 8.2935
-2.7285e-12 | 8.6845 | -2.2055e-11 | 8.6995
-1.3642e-12 | 9.0905 | -2.3192e-11 | 9.1075
-3.4106e-12 | 9.4965 | -2.0236e-11 | 9.5125
-2.0464e-12 | 9.9016 | -2.251e-11 | 9.9186
-1.819e-12 | 10.3066 | -2.2283e-11 | 10.3256
-2.2737e-12 | 10.7106 | -2.4329e-11 | 10.7386
-2.2737e-12 | 11.1156 | -2.3647e-11 | 11.1426
-3.638e-12 | 11.5207 | -2.0236e-11 | 11.5557
-1.1369e-12 | 11.9247 | -2.4329e-11 | 11.9617
-1.819e-12 | 12.3307 | -2.3647e-11 | 12.3677
-1.3642e-12 | 12.7347 | -2.2055e-11 | 12.7727
-1.819e-12 | 13.1387 | -1.9099e-11 | 13.1777
-3.1832e-12 | 13.5438 | -1.9782e-11 | 13.6088
-3.1832e-12 | 13.9488 | -2.1828e-11 | 14.0418
-2.5011e-12 | 14.3538 | -2.5011e-11 | 14.4468
-3.8654e-12 | 14.7588 | -2.7967e-11 | 14.8529
-3.638e-12 | 15.1629 | -2.2283e-11 | 15.2589
-2.2737e-12 | 15.5679 | -2.1828e-11 | 15.6639
-2.2737e-12 | 15.9719 | -2.2055e-11 | 16.0679
-2.7285e-12 | 16.3769 | -2.0691e-11 | 16.4739
-3.4106e-12 | 16.781 | -2.0236e-11 | 16.879
-1.819e-12 | 17.186 | -2.5693e-11 | 17.284
-1.1369e-12 | 17.59 | -2.7512e-11 | 17.691
-2.0464e-12 | 17.995 | -2.4102e-11 | 18.096
-1.819e-12 | 18.4011 | -2.2055e-11 | 18.5011
-9.0949e-13 | 18.8051 | -2.3647e-11 | 18.9061
-3.8654e-12 | 19.2091 | -2.1828e-11 | 19.3111
-2.2737e-12 | 19.6141 | -2.3419e-11 | 19.7161
-2.0464e-12 | 20.0181 | -2.2055e-11 | 20.1222
-6.8212e-13 | 20.4242 | -2.2965e-11 | 20.5282
-3.638e-12 | 20.8282 | -2.16e-11 | 20.9332
-2.2737e-12 | 21.2332 | -2.3192e-11 | 21.3382
-1.1369e-12 | 21.6392 | -2.4784e-11 | 21.7442
-9.0949e-13 | 22.0443 | -2.2283e-11 | 22.1493
-2.0464e-12 | 22.4483 | -2.251e-11 | 22.5563
-4.3201e-12 | 22.8523 | -1.9782e-11 | 22.9623
-4.5475e-12 | 23.2573 | -2.251e-11 | 23.3693
-1.1369e-12 | 23.6623 | -2.0918e-11 | 23.7744
-1.819e-12 | 24.0674 | -1.9099e-11 | 24.1804
-4.7748e-12 | 24.4714 | -2.16e-11 | 24.5854
-3.4106e-12 | 24.8754 | -2.16e-11 | 24.9914
-1.3642e-12 | 25.2804 | -2.1828e-11 | 25.3965
-1.819e-12 | 25.6865 | -2.0918e-11 | 25.8015
-2.7285e-12 | 26.0905 | -1.9554e-11 | 26.2065
-4.5475e-12 | 26.4955 | -2.2283e-11 | 26.6115
-1.3642e-12 | 26.9005 | -2.3192e-11 | 27.0165
-9.0949e-13 | 27.3056 | -2.0464e-11 | 27.4216
-1.1369e-12 | 27.7086 | -2.0918e-11 | 27.8276
-3.8654e-12 | 28.1136 | -2.5011e-11 | 28.2336
-3.1832e-12 | 28.5186 | -2.2283e-11 | 28.6396
-1.1369e-12 | 28.9216 | -2.0691e-11 | 29.0457
6.8212e-13 | 29.3257 | -2.0236e-11 | 29.4517
-2.2737e-12 | 29.7287 | -2.3192e-11 | 29.8577
-6.1391e-12 | 30.1337 | -2.16e-11 | 30.2637
-2.7285e-12 | 30.5388 | -1.9099e-11 | 30.6687
6.8212e-13 | 30.9448 | -2.2055e-11 | 31.0738
-6.8212e-13 | 31.3488 | -2.6375e-11 | 31.4778
-6.3665e-12 | 31.7528 | -2.0236e-11 | 31.8848
-6.3665e-12 | 32.1578 | -2.0464e-11 | 32.2888
1.1369e-12 | 32.5629 | -2.3192e-11 | 32.6949
-1.1369e-12 | 32.9689 | -2.2965e-11 | 33.0999
-7.5033e-12 | 33.3739 | -2.2965e-11 | 33.5049
-3.8654e-12 | 33.7789 | -1.9554e-11 | 33.9099
0 | 34.183 | -2.2283e-11 | 34.315
2.2737e-13 | 34.585 | -2.3647e-11 | 34.72
2.2737e-13 | 34.989 | -1.8645e-11 | 35.126
-4.3201e-12 | 35.393 | -2.3647e-11 | 35.531
-4.5475e-12 | 35.7981 | -2.3419e-11 | 35.9351
-1.1369e-12 | 36.2051 | -2.683e-11 | 36.3421
-6.8212e-13 | 36.6091 | -1.9327e-11 | 36.7471
-3.638e-12 | 37.0121 | -2.7285e-11 | 37.1531
-7.0486e-12 | 37.4181 | -2.5466e-11 | 37.5572
-1.819e-12 | 37.8232 | -2.2283e-11 | 37.9642
-4.5475e-13 | 38.2272 | -2.3192e-11 | 38.3752
-3.1832e-12 | 38.6312 | -2.1146e-11 | 38.7792
-3.8654e-12 | 39.0352 | -2.3647e-11 | 39.1842
-5.2296e-12 | 39.4403 | -2.2283e-11 | 39.5993
-2.0464e-12 | 39.8443 | -1.7735e-11 | 40.0043
-1.3642e-12 | 40.2493 | -2.5921e-11 | 40.4103
-2.0464e-12 | 40.6543 | -2.7057e-11 | 40.8163
-3.8654e-12 | 41.0603 | -1.9554e-11 | 41.2224
-1.3642e-12 | 41.4644 | -2.0691e-11 | 41.6274
-6.8212e-13 | 41.8684 | -2.7512e-11 | 42.0334
-1.1369e-12 | 42.2734 | -2.3419e-11 | 42.4394
-5.9117e-12 | 42.6784 | -2.16e-11 | 42.8464
-5.0022e-12 | 43.0825 | -2.4329e-11 | 43.2515
-1.1369e-12 | 43.4875 | -2.5466e-11 | 43.6565
-2.7285e-12 | 43.8925 | -1.5461e-11 | 44.0625
-4.3201e-12 | 44.2975 | -2.2055e-11 | 44.4675
-4.3201e-12 | 44.7026 | -2.9559e-11 | 44.8736
-6.8212e-13 | 45.1076 | -2.3192e-11 | 45.2776
-3.638e-12 | 45.5116 | -2.2283e-11 | 45.6826
-3.4106e-12 | 45.9166 | -2.4784e-11 | 46.0886
-5.2296e-12 | 46.3216 | -2.5693e-11 | 46.4937
-6.8212e-13 | 46.7257 | -1.9099e-11 | 46.8997
-9.0949e-13 | 47.1297 | -2.0236e-11 | 47.3047
0 | 47.5327 | -2.5466e-11 | 47.7097
-3.1832e-12 | 47.9347 | -2.4329e-11 | 48.1157
-5.0022e-12 | 48.3398 | -1.7053e-11 | 48.5218
-4.7748e-12 | 48.7448 | -2.2965e-11 | 48.9268
4.5475e-13 | 49.1498 | -2.7057e-11 | 49.3318
-2.0464e-12 | 49.5548 | -2.1828e-11 | 49.7358
-4.7748e-12 | 49.9589 | -2.2055e-11 | 50.1409
-4.5475e-12 | 50.3639 | -2.0691e-11 | 50.5469
2.2737e-13 | 50.7679 | -2.5693e-11 | 50.9519
2.2737e-13 | 51.1729 | -2.251e-11 | 51.3579
-1.3642e-12 | 51.5769 | -2.3419e-11 | 51.764
-5.0022e-12 | 51.981 | -2.7057e-11 | 52.17
-3.1832e-12 | 52.387 | -2.0691e-11 | 52.575
-2.2737e-12 | 52.791 | -2.16e-11 | 52.981
-4.5475e-13 | 53.195 | -2.1146e-11 | 53.386
-3.1832e-12 | 53.6001 | -2.3647e-11 | 53.7921
-3.4106e-12 | 54.0041 | -1.9099e-11 | 54.1981
2.2737e-13 | 54.4101 | -2.0691e-11 | 54.6031
-4.5475e-13 | 54.8141 | -2.8194e-11 | 55.0091
-1.3642e-12 | 55.2182 | -2.4329e-11 | 55.4142
-6.1391e-12 | 55.6232 | -1.9782e-11 | 55.8202
-5.2296e-12 | 56.0272 | -2.7285e-11 | 56.2262
-9.0949e-13 | 56.4312 | -2.5693e-11 | 56.6312
-1.3642e-12 | 56.8353 | -1.9099e-11 | 57.0373
-3.638e-12 | 57.2403 | -2.251e-11 | 57.4433
-3.4106e-12 | 57.6453 | -2.5011e-11 | 57.8483
-2.7285e-12 | 58.0523 | -2.1828e-11 | 58.2543
-1.819e-12 | 58.4573 | -1.9554e-11 | 58.6594
-3.638e-12 | 58.8634 | -2.5921e-11 | 59.0654
-5.0022e-12 | 59.2674 | -2.5011e-11 | 59.4704
-2.2737e-12 | 59.6724 | -2.0464e-11 | 59.8774
-1.3642e-12 | 60.0774 | -2.3647e-11 | 60.2844
-2.5011e-12 | 60.4815 | -2.2965e-11 | 60.6895
-3.638e-12 | 60.8865 | -2.1146e-11 | 61.0945
-1.819e-12 | 61.2915 | -2.3192e-11 | 61.5005
-4.5475e-13 | 61.6965 | -2.4329e-11 | 61.9065
-2.7285e-12 | 62.1016 | -2.16e-11 | 62.3116
-4.5475e-12 | 62.5056 | -2.0236e-11 | 62.7166
-3.8654e-12 | 62.9096 | -2.683e-11 | 63.1226
-1.1369e-12 | 63.3136 | -2.0236e-11 | 63.5286
-1.819e-12 | 63.7186 | -2.16e-11 | 63.9357
-3.638e-12 | 64.1237 | -2.16e-11 | 64.3417
-2.2737e-12 | 64.5277 | -2.2283e-11 | 64.7477
-1.3642e-12 | 64.9327 | -2.1146e-11 | 65.1537
-1.3642e-12 | 65.3377 | -2.6148e-11 | 65.5588
-3.8654e-12 | 65.7428 | -2.5693e-11 | 65.9638
-2.5011e-12 | 66.1468 | -2.4102e-11 | 66.3698
-1.3642e-12 | 66.5528 | -2.0464e-11 | 66.7748
-1.819e-12 | 66.9578 | -2.5466e-11 | 67.1808
-2.0464e-12 | 67.3619 | -2.6148e-11 | 67.5869
-5.9117e-12 | 67.7659 | -2.16e-11 | 67.9919
-4.3201e-12 | 68.1709 | -2.4784e-11 | 68.3969
-1.819e-12 | 68.5759 | -2.7057e-11 | 68.8029
-2.7285e-12 | 68.9819 | -2.1146e-11 | 69.207
-5.0022e-12 | 69.386 | -2.0464e-11 | 69.613
-3.1832e-12 | 69.791 | -2.6375e-11 | 70.019
-2.5011e-12 | 70.196 | -2.2055e-11 | 70.425
-1.1369e-12 | 70.601 | -2.0918e-11 | 70.83
-4.5475e-12 | 71.0051 | -2.0691e-11 | 71.2351
-3.1832e-12 | 71.4101 | -2.5011e-11 | 71.6401
-1.3642e-12 | 71.8141 | -2.3647e-11 | 72.0461
-1.1369e-12 | 72.2201 | -1.8645e-11 | 72.4521
-2.2737e-12 | 72.6261 | -2.5921e-11 | 72.8592
-4.3201e-12 | 73.0312 | -2.4329e-11 | 73.2652
-6.8212e-13 | 73.4362 | -2.4784e-11 | 73.6702
-2.2737e-12 | 73.8402 | -2.0464e-11 | 74.0752
-5.0022e-12 | 74.2452 | -2.4102e-11 | 74.4803
-3.638e-12 | 74.6493 | -1.9327e-11 | 74.8863
-6.8212e-13 | 75.0543 | -2.251e-11 | 75.2903
-1.1369e-12 | 75.4583 | -2.251e-11 | 75.6943
-4.3201e-12 | 75.8633 | -2.2055e-11 | 76.0994
-3.8654e-12 | 76.2674 | -2.251e-11 | 76.5054
-3.8654e-12 | 76.6724 | -2.16e-11 | 76.9094
-9.0949e-13 | 77.0774 | -2.6375e-11 | 77.3144
-2.0464e-12 | 77.4824 | -2.6375e-11 | 77.7194
-4.5475e-12 | 77.8864 | -2.251e-11 | 78.1235
-1.819e-12 | 78.2905 | -2.4329e-11 | 78.5285
-6.8212e-13 | 78.6955 | -2.3192e-11 | 78.9335
-6.8212e-13 | 79.0995 | -2.4556e-11 | 79.3405
-1.819e-12 | 79.5036 | -2.2283e-11 | 79.7446
-5.2296e-12 | 79.9076 | -2.3192e-11 | 80.1506
-3.4106e-12 | 80.3126 | -2.3647e-11 | 80.5556
0 | 80.7176 | -2.0464e-11 | 80.9606
0 | 81.1196 | -2.4102e-11 | 81.3646
-4.5475e-13 | 81.5217 | -2.2055e-11 | 81.7707
-1.3642e-12 | 81.9247 | -2.4556e-11 | 82.1767
-5.9117e-12 | 82.3297 | -1.6598e-11 | 82.5827
-1.1369e-12 | 82.7347 | -2.3192e-11 | 82.9868
-4.5475e-13 | 83.1408
-2.7285e-12 | 83.5448
-5.2296e-12 | 83.9488
-3.638e-12 | 84.3538
-6.8212e-13 | 84.7579
-4.5475e-13 | 85.1619
-3.8654e-12 | 85.5659
-5.0022e-12 | 85.9719
-4.5475e-13 | 86.3759
-9.0949e-13 | 86.78
-2.7285e-12 | 87.185
-4.3201e-12 | 87.59
-1.3642e-12 | 87.995
0 | 88.4011

## 580uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-3.2457e-12 | 3.446e-12 | -3.0508e-11 | 4.6055e-12
GEMs On |  | GEMs On
-5.6843e-12 | 0.31202 | -2.8876e-11 | 0.31002
-6.8212e-12 | 0.99606 | -3.0468e-11 | 0.99406
6.8212e-13 | 1.4011 | -2.5238e-11 | 1.4031
-4.5475e-13 | 1.8051 | -2.7967e-11 | 1.8091
-7.9581e-12 | 2.2091 | -3.4106e-11 | 2.2151
-6.1391e-12 | 2.6141 | -2.9331e-11 | 2.6212
3.4106e-12 | 3.0202 | -2.9104e-11 | 3.0262
-1.3642e-12 | 3.4242 | -4.0018e-11 | 3.4332
-8.1855e-12 | 3.8292 | -3.2287e-11 | 3.8372
2.7285e-12 | 4.2552 | -2.6148e-11 | 4.2422
-1.0459e-11 | 4.6683 | -3.3197e-11 | 4.6473
-2.2737e-12 | 5.0733 | -3.9563e-11 | 5.0533
5.457e-12 | 5.4783 | -2.7512e-11 | 5.4583
-6.3665e-12 | 5.8833 | -2.7285e-11 | 5.8633
-1.2051e-11 | 6.2884 | -3.5016e-11 | 6.2684
2.2737e-12 | 6.6944 | -3.3879e-11 | 6.6734
-6.1391e-12 | 7.1024 | -3.0013e-11 | 7.0784
-7.0486e-12 | 7.5084 | -2.6148e-11 | 7.4834
3.4106e-12 | 7.9145 | -3.6607e-11 | 7.8885
-2.5011e-12 | 8.3195 | -3.0695e-11 | 8.2955
-7.7307e-12 | 8.7245 | -2.5921e-11 | 8.7015
-6.3665e-12 | 9.1295 | -3.2969e-11 | 9.1085
1.5916e-12 | 9.5335 | -3.3197e-11 | 9.5135
-1.5916e-12 | 9.9386 | -2.4102e-11 | 9.9196
-8.6402e-12 | 10.3436 | -3.1378e-11 | 10.3256
1.5916e-12 | 10.7496 | -3.5016e-11 | 10.7306
-2.7285e-12 | 11.1576 | -2.5011e-11 | 11.1366
-7.5033e-12 | 11.5617 | -3.0468e-11 | 11.5427
-2.4329e-11 | 11.9667 | -3.5698e-11 | 11.9487
-1.5916e-12 | 12.3707 | -2.7967e-11 | 12.3537
-3.8654e-12 | 12.7767 | -2.774e-11 | 12.7597
-1.1823e-11 | 13.1807 | -3.5016e-11 | 13.1647
-5.0022e-12 | 13.5858 | -3.3197e-11 | 13.5698
3.4106e-12 | 13.9908 | -3.0695e-11 | 13.9748
-7.0486e-12 | 14.3958 | -2.5466e-11 | 14.3798
-7.7307e-12 | 14.8009 | -3.4106e-11 | 14.7849
7.276e-12 | 15.2049 | -2.5921e-11 | 15.1909
1.3642e-12 | 15.6099 | -2.4329e-11 | 15.5969
-6.3665e-12 | 16.0149 | -3.5243e-11 | 16.0019
-8.6402e-12 | 16.4189 | -2.9786e-11 | 16.4079
6.8212e-13 | 16.825 | -2.4102e-11 | 16.812
3.8654e-12 | 17.228 | -2.9786e-11 | 17.218
-9.0949e-13 | 17.632 | -2.8649e-11 | 17.623
-6.3665e-12 | 18.036 | -3.1605e-11 | 18.028
-7.7307e-12 | 18.4411 | -2.5238e-11 | 18.434
2.2737e-12 | 18.8451 | -3.1378e-11 | 18.8391
3.1832e-12 | 19.2501 | -3.6607e-11 | 19.2441
-6.8212e-12 | 19.6551 | -2.6603e-11 | 19.6501
-6.3665e-12 | 20.0591 | -2.5466e-11 | 20.0552
6.8212e-13 | 20.4642 | -3.547e-11 | 20.4602
-4.5475e-13 | 20.8692 | -3.0468e-11 | 20.8662
-7.5033e-12 | 21.2732 | -2.8649e-11 | 21.2712
-5.2296e-12 | 21.6782 | -2.8649e-11 | 21.6752
9.0949e-13 | 22.0833 | -3.4561e-11 | 22.0813
-1.3642e-12 | 22.4863 | -3.2514e-11 | 22.4863
-5.457e-12 | 22.8923 | -2.3419e-11 | 22.8923
-3.8654e-12 | 23.2973 | -2.7285e-11 | 23.2963
0 | 23.7024 | -3.8426e-11 | 23.7014
0 | 24.1054 | -2.9786e-11 | 24.1064
-3.1832e-12 | 24.5084 | -2.5466e-11 | 24.5114
-7.7307e-12 | 24.9134 | -3.3197e-11 | 24.9174
-3.638e-12 | 25.3175 | -3.1605e-11 | 25.3235
0 | 25.7225 | -2.4784e-11 | 25.7295
0 | 26.1235 | -3.2514e-11 | 26.1345
-2.2737e-12 | 26.5265 | -3.5243e-11 | 26.5385
-8.1855e-12 | 26.9325 | -2.7285e-11 | 26.9435
-1.1369e-12 | 27.3376 | -2.9331e-11 | 27.3496
0 | 27.7426 | -3.1605e-11 | 27.7546
-2.2737e-12 | 28.1446 | -3.2969e-11 | 28.1606
-7.5033e-12 | 28.5496 | -2.3647e-11 | 28.5666
-3.638e-12 | 28.9557 | -3.0695e-11 | 28.9727
2.9559e-12 | 29.3607 | -3.8881e-11 | 29.3777
6.8212e-13 | 29.7637 | -2.251e-11 | 29.7827
-5.6843e-12 | 30.1677 | -2.6603e-11 | 30.1877
-8.1855e-12 | 30.5727 | -3.3651e-11 | 30.5927
2.2737e-12 | 30.9778 | -3.7971e-11 | 30.9988
-9.0949e-13 | 31.3838 | -2.8649e-11 | 31.4018
-9.5497e-12 | 31.7878 | -2.2965e-11 | 31.8078
-3.1832e-12 | 32.1938 | -3.5698e-11 | 32.2138
2.2737e-12 | 32.5999 | -3.2514e-11 | 32.6189
-8.1855e-12 | 33.0039 | -2.4102e-11 | 33.0239
-5.6843e-12 | 33.4099 | -3.0013e-11 | 33.4289
1.1369e-12 | 33.8149 | -3.8199e-11 | 33.8339
1.1369e-12 | 34.22 | -2.5466e-11 | 34.238
-5.6843e-12 | 34.625 | -2.0464e-11 | 34.643
-6.8212e-12 | 35.03 | -3.7971e-11 | 35.049
3.4106e-12 | 35.435 | -3.4333e-11 | 35.455
3.4106e-12 | 35.8391 | -2.5921e-11 | 35.8601
-3.638e-12 | 36.2431 | -2.7285e-11 | 36.2651
-6.1391e-12 | 36.6471 | -3.4333e-11 | 36.6721
2.2737e-12 | 37.0521 | -2.9786e-11 | 37.0771
1.3642e-12 | 37.4571 | -2.7512e-11 | 37.4831
-6.3665e-12 | 37.8622 | -3.547e-11 | 37.8882
-5.2296e-12 | 38.2662 | -3.3879e-11 | 38.2942
9.0949e-13 | 38.6702 | -2.683e-11 | 38.7002
1.5916e-12 | 39.0752 | -3.0695e-11 | 39.1052
-6.1391e-12 | 39.4803 | -3.7744e-11 | 39.5103
-3.8654e-12 | 39.8853 | -2.6375e-11 | 39.9163
-4.5475e-13 | 40.2913 | -2.8194e-11 | 40.3213
-4.5475e-13 | 40.6953 | -3.4106e-11 | 40.7273
-5.457e-12 | 41.0994 | -3.4333e-11 | 41.1334
-5.6843e-12 | 41.5034 | -2.5921e-11 | 41.5394
-4.0927e-12 | 41.9074 | -3.2287e-11 | 41.9444
6.8212e-13 | 42.3124 | -3.7744e-11 | 42.3494
-2.2737e-13 | 42.7164 | -2.2737e-11 | 42.7555
-3.638e-12 | 43.1195 | -3.1605e-11 | 43.1615
-5.6843e-12 | 43.5245 | -3.5698e-11 | 43.5675
-4.5475e-13 | 43.9295 | -2.5466e-11 | 43.9725
2.2737e-13 | 44.3325 | -2.5011e-11 | 44.3775
-4.3201e-12 | 44.7366 | -2.9786e-11 | 44.7816
-4.7748e-12 | 45.1406 | -3.4561e-11 | 45.1866
-5.6843e-12 | 45.5456 | -2.2965e-11 | 45.5916
-1.3642e-12 | 45.9496 | -3.1378e-11 | 45.9976
-2.9559e-12 | 46.3537 | -3.638e-11 | 46.4027
-8.8676e-12 | 46.7597 | -3.4106e-11 | 46.8077
-5.6843e-12 | 47.1637 | -2.2283e-11 | 47.2137
-1.3642e-12 | 47.5687 | -2.774e-11 | 47.6187
-2.9559e-12 | 47.9747 | -3.6607e-11 | 48.0237
-2.9559e-12 | 48.3798 | -2.251e-11 | 48.4298
-3.1832e-12 | 48.7858 | -3.0695e-11 | 48.8348
-9.0949e-13 | 49.1898 | -3.4333e-11 | 49.2398
-3.638e-12 | 49.5938 | -3.6607e-11 | 49.6458
-6.1391e-12 | 49.9999 | -2.4784e-11 | 50.0509
-3.8654e-12 | 50.4039 | -2.8876e-11 | 50.4559
6.8212e-13 | 50.8089 | -3.8199e-11 | 50.8609
-2.7285e-12 | 51.2129 | -2.7967e-11 | 51.2669
-7.5033e-12 | 51.618 | -2.5238e-11 | 51.673
2.2737e-13 | 52.025 | -3.5698e-11 | 52.078
-1.5916e-12 | 52.428 | -3.3651e-11 | 52.484
-3.8654e-12 | 52.834 | -2.251e-11 | 52.889
-5.6843e-12 | 53.239 | -2.9786e-11 | 53.294
1.1369e-12 | 53.6441 | -3.1605e-11 | 53.6991
1.1369e-12 | 54.0481 | -3.0013e-11 | 54.1041
-3.1832e-12 | 54.4531 | -2.3874e-11 | 54.5101
2.2737e-13 | 54.8581 | -3.3879e-11 | 54.9171
-2.7285e-12 | 55.2622 | -3.0013e-11 | 55.3232
-9.0949e-13 | 55.6662 | -2.9786e-11 | 55.7292
-1.1369e-12 | 56.0702 | -3.4561e-11 | 56.1362
-4.3201e-12 | 56.4752 | -3.2742e-11 | 56.5412
-5.0022e-12 | 56.8813 | -2.4329e-11 | 56.9463
0 | 57.2863 | -2.8649e-11 | 57.3513
-2.2737e-12 | 57.6893 | -3.547e-11 | 57.7573
-7.7307e-12 | 58.0943 | -2.7512e-11 | 58.1623
-5.0022e-12 | 58.4993 | -2.5011e-11 | 58.5673
1.3642e-12 | 58.9044 | -2.9104e-11 | 58.9724
6.8212e-13 | 59.3084 | -3.7517e-11 | 59.3774
-4.3201e-12 | 59.7114 | -2.9331e-11 | 59.7824
-5.0022e-12 | 60.1164 | -2.5921e-11 | 60.1884
-2.2737e-12 | 60.5215 | -3.3879e-11 | 60.5945
2.2737e-13 | 60.9265 | -3.3651e-11 | 61.0005
-3.8654e-12 | 61.3305 | -2.2055e-11 | 61.4055
-5.457e-12 | 61.7345 | -3.2742e-11 | 61.8125
-2.7285e-12 | 62.1395 | -3.3879e-11 | 62.2166
9.0949e-13 | 62.5446 | -2.6148e-11 | 62.6236
-4.3201e-12 | 62.9496 | -2.8194e-11 | 63.0286
-5.6843e-12 | 63.3546 | -3.8426e-11 | 63.4336
-5.0022e-12 | 63.7597 | -2.6148e-11 | 63.8387
-2.2737e-13 | 64.1637 | -2.8876e-11 | 64.2437
-1.819e-12 | 64.5687 | -3.4106e-11 | 64.6487
-6.1391e-12 | 64.9737 | -3.2287e-11 | 65.0547
-5.0022e-12 | 65.3777 | -2.16e-11 | 65.4607
-9.0949e-13 | 65.7828 | -2.6603e-11 | 65.8648
-1.3642e-12 | 66.1858 | -3.3651e-11 | 66.2708
-5.6843e-12 | 66.5908 | -3.6152e-11 | 66.6758
-5.2296e-12 | 66.9948 | -2.1146e-11 | 67.0808
-1.3642e-12 | 67.3999 | -3.3879e-11 | 67.4859
-1.819e-12 | 67.8059 | -3.9108e-11 | 67.8919
-4.7748e-12 | 68.2109 | -2.7285e-11 | 68.2969
-3.638e-12 | 68.6159 | -2.8649e-11 | 68.7019
-1.5916e-12 | 69.022 | -3.5243e-11 | 69.1069
-2.9559e-12 | 69.425 | -2.9786e-11 | 69.513
-4.7748e-12 | 69.83 | -2.4329e-11 | 69.918
-5.2296e-12 | 70.234 | -3.1378e-11 | 70.323
-2.2737e-13 | 70.639 | -3.2742e-11 | 70.7291
-3.1832e-12 | 71.0431 | -3.2287e-11 | 71.1331
-4.7748e-12 | 71.4481 | -2.6603e-11 | 71.5391
-6.5938e-12 | 71.8521 | -3.6152e-11 | 71.9451
-2.2737e-12 | 72.2581 | -3.5698e-11 | 72.3511
-9.0949e-13 | 72.6612 | -2.7285e-11 | 72.7572
-2.7285e-12 | 73.0662 | -3.2969e-11 | 73.1622
-3.8654e-12 | 73.4702 | -3.7744e-11 | 73.5672
-3.638e-12 | 73.8742 | -3.0013e-11 | 73.9722
-1.819e-12 | 74.2782 | -2.8649e-11 | 74.3782
-1.1369e-12 | 74.6833 | -3.3879e-11 | 74.7843
-6.1391e-12 | 75.0883 | -2.3419e-11 | 75.1893
-4.3201e-12 | 75.4933 | -2.8876e-11 | 75.5963
-2.7285e-12 | 75.8983 | -3.547e-11 | 76.0014
-2.2737e-13 | 76.3034 | -3.547e-11 | 76.4064
-4.0927e-12 | 76.7074 | -2.683e-11 | 76.8114
-6.1391e-12 | 77.1104 | -2.8194e-11 | 77.2174
-1.819e-12 | 77.5144 | -3.9108e-11 | 77.6224
-2.2737e-12 | 77.9195 | -3.0013e-11 | 78.0285
-2.7285e-12 | 78.3225 | -2.8876e-11 | 78.4325
-4.7748e-12 | 78.7275 | -3.4561e-11 | 78.8385
-5.457e-12 | 79.1315 | -3.638e-11 | 79.2435
-2.2737e-12 | 79.5356 | -2.8194e-11 | 79.6486
-3.1832e-12 | 79.9406 | -2.6148e-11 | 80.0536
-2.9559e-12 | 80.3456 | -3.547e-11 | 80.4586
-3.8654e-12 | 80.7506 | -2.8194e-11 | 80.8646
-1.5916e-12 | 81.1546 | -2.5921e-11 | 81.2707
2.2737e-13 | 81.5587 | -3.2514e-11 | 81.6757
-2.2737e-13 | 81.9617 | -4.1609e-11 | 82.0817
-3.638e-12 | 82.3667 | -2.4102e-11 | 82.4867
-5.0022e-12 | 82.7697 | -2.9104e-11 | 82.8927
-2.9559e-12 | 83.1748 | -3.6835e-11 | 83.2978
-2.7285e-12 | 83.5798 | -3.115e-11 | 83.7038
-4.7748e-12 | 83.9848 | -2.683e-11 | 84.1098
-4.0927e-12 | 84.3898
-1.819e-12 | 84.7948
6.8212e-13 | 85.1989
-2.5011e-12 | 85.6019
-6.8212e-12 | 86.0069
-3.8654e-12 | 86.4119
-1.5916e-12 | 86.816
-1.5916e-12 | 87.221
-6.5938e-12 | 87.627
-2.7285e-12 | 88.031
-1.3642e-12 | 88.4361
-9.0949e-13 | 88.8411
-6.5938e-12 | 89.2451
-3.638e-12 | 89.6501
-1.5916e-12 | 90.0551
6.8212e-13 | 90.4602
-6.8212e-12 | 90.8642
-4.7748e-12 | 91.2682
-2.2737e-13 | 91.6732
0 | 92.0763
-3.1832e-12 | 92.4793
-6.5938e-12 | 92.8833
-5.2296e-12 | 93.2883
9.0949e-13 | 93.6934
-2.7285e-12 | 94.0974
-5.457e-12 | 94.5014
-6.8212e-12 | 94.9064
-3.1832e-12 | 95.3114
2.0464e-12 | 95.7155
-4.7748e-12 | 96.1195
-5.0022e-12 | 96.5245
-4.5475e-13 | 96.9305
0 | 97.3346
-1.3642e-12 | 97.7356
-7.9581e-12 | 98.1416
-5.6843e-12 | 98.5466
1.1369e-12 | 98.9517
-3.638e-12 | 99.3567
-5.6843e-12 | 99.7627
-2.9559e-12 | 100.1687
-9.0949e-13 | 100.5738
-6.3665e-12 | 100.9778
-5.0022e-12 | 101.3838
-2.5011e-12 | 101.7888
-4.5475e-13 | 102.1928
-4.0927e-12 | 102.5959
-7.5033e-12 | 103.0009
-5.0022e-12 | 103.4049
0 | 103.8099
-4.0927e-12 | 104.212
-5.6843e-12 | 104.618
-5.2296e-12 | 105.023
-9.0949e-13 | 105.427
-1.1369e-12 | 105.8301
-1.819e-12 | 106.2341
-6.3665e-12 | 106.6391
-1.1369e-12 | 107.0441
1.5916e-12 | 107.4501
-7.9581e-12 | 107.8552
-6.5938e-12 | 108.2592
-1.3642e-12 | 108.6652
6.8212e-13 | 109.0682
-2.9559e-12 | 109.4713
-6.8212e-12 | 109.8753
-7.7307e-12 | 110.2803
-9.0949e-13 | 110.6863
-9.0949e-13 | 111.0904
-6.8212e-12 | 111.4944
-8.8676e-12 | 111.8994
-1.1369e-12 | 112.3044
-1.3642e-12 | 112.7094
-6.1391e-12 | 113.1145
-4.7748e-12 | 113.5185
2.2737e-12 | 113.9235
-1.5916e-12 | 114.3275
-5.0022e-12 | 114.7316
-7.7307e-12 | 115.1366
-9.0949e-13 | 115.5416
0 | 115.9456
-1.819e-12 | 116.3477
-6.3665e-12 | 116.7537
-1.3642e-12 | 117.1597
6.8212e-13 | 117.5637
-2.9559e-12 | 117.9677
-8.8676e-12 | 118.3728
-5.2296e-12 | 118.7768
-4.5475e-13 | 119.1818
-5.2296e-12 | 119.5868
-6.5938e-12 | 119.9939
-2.5011e-12 | 120.3989
-2.9559e-12 | 120.8039
-5.0022e-12 | 121.2079
-7.7307e-12 | 121.613
-4.0927e-12 | 122.017
-1.3642e-12 | 122.423
-4.3201e-12 | 122.828
-6.1391e-12 | 123.233
-1.3642e-12 | 123.6381
-3.1832e-12 | 124.0441
-5.457e-12 | 124.4491
-3.8654e-12 | 124.8531
-2.9559e-12 | 125.2572
-3.1832e-12 | 125.6622
-3.638e-12 | 126.0662
-4.7748e-12 | 126.4702
-1.3642e-12 | 126.8753
-2.7285e-12 | 127.2803
-5.0022e-12 | 127.6853
-5.6843e-12 | 128.0893
-2.7285e-12 | 128.4943
-3.1832e-12 | 128.9004
-2.2737e-12 | 129.3044
-5.6843e-12 | 129.7084
-4.7748e-12 | 130.1114
-4.3201e-12 | 130.5155
-2.9559e-12 | 130.9205
-3.638e-12 | 131.3245
-4.0927e-12 | 131.7295
-2.7285e-12 | 132.1336
-3.638e-12 | 132.5376
-3.638e-12 | 132.9426
-3.1832e-12 | 133.3486
-4.7748e-12 | 133.7537
-2.7285e-12 | 134.1577
-4.3201e-12 | 134.5637
-3.8654e-12 | 134.9677
-3.8654e-12 | 135.3737
-6.1391e-12 | 135.7788
-2.7285e-12 | 136.1848
-1.3642e-12 | 136.5888
-3.1832e-12 | 136.9928
-5.2296e-12 | 137.3979
-4.7748e-12 | 137.8019
-2.9559e-12 | 138.2079
-3.638e-12 | 138.6119
-4.3201e-12 | 139.017
-2.9559e-12 | 139.422
-1.819e-12 | 139.827
-2.1828e-11 | 140.232
-5.0022e-12 | 140.637
-3.638e-12 | 141.0431
-2.7285e-12 | 141.4481
-2.5011e-12 | 141.8541
-5.0022e-12 | 142.2591
-5.6843e-12 | 142.6632
-4.3201e-12 | 143.0692
-3.8654e-12 | 143.4742
1.3642e-11 | 143.8792
-1.5916e-12 | 144.2843
-2.9559e-12 | 144.6893
-5.457e-12 | 145.0943
-4.3201e-12 | 145.4993
-4.0927e-12 | 145.9043
-2.2737e-12 | 146.3094
-4.3201e-12 | 146.7134
-5.457e-12 | 147.1184
-2.9559e-12 | 147.5224

## 590uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-3.6136e-12 | 1.7347e-12 | -4.1796e-11 | 3.1213e-12
GEMs On |  | GEMs On
-1.819e-12 | 0.31002 | -4.7521e-11 | 0.31402
-4.3201e-12 | 0.99406 | -3.7289e-11 | 0.99806
-3.8654e-12 | 1.3991 | -4.502e-11 | 1.4051
-4.0927e-12 | 1.8031 | -4.07e-11 | 1.8111
-3.8654e-12 | 2.2081 | -4.0018e-11 | 2.2171
-4.5475e-12 | 2.6191 | -3.9336e-11 | 2.6221
-2.5011e-12 | 3.0252 | -4.2974e-11 | 3.0282
-3.1832e-12 | 3.4312 | -3.7289e-11 | 3.4342
-2.5011e-12 | 3.8372 | -3.5698e-11 | 3.8392
-3.4106e-12 | 4.2412 | -4.7976e-11 | 4.2442
-1.819e-12 | 4.6483 | -4.5702e-11 | 4.6493
-3.8654e-12 | 5.0543 | -4.1155e-11 | 5.0533
-4.3201e-12 | 5.4633 | -3.9108e-11 | 5.4593
-3.1832e-12 | 5.8703 | -3.638e-11 | 5.8643
-4.5475e-12 | 6.2764 | -4.3656e-11 | 6.2694
-4.5475e-12 | 6.6834 | -4.1837e-11 | 6.6754
-1.1369e-12 | 7.0894 | -4.1382e-11 | 7.0814
-3.1832e-12 | 7.4974 | -4.4565e-11 | 7.4874
-2.5011e-12 | 7.9025 | -4.2064e-11 | 7.8915
-4.5475e-12 | 8.3105 | -3.5243e-11 | 8.2975
-4.5475e-12 | 8.7165 | -4.5475e-11 | 8.7025
-4.0927e-12 | 9.1205 | -4.4338e-11 | 9.1075
-3.1832e-12 | 9.5275 | -4.0245e-11 | 9.5115
-1.5916e-12 | 9.9326 | -4.3883e-11 | 9.9176
-2.5011e-12 | 10.3426 | -3.8199e-11 | 10.3226
-2.7285e-12 | 10.7476 | -4.411e-11 | 10.7286
-1.5916e-12 | 11.1526 | -4.1837e-11 | 11.1356
-3.1832e-12 | 11.5577 | -4.4338e-11 | 11.5417
-3.4106e-12 | 11.9657 | -3.9108e-11 | 11.9457
-3.8654e-12 | 12.3717 | -4.5247e-11 | 12.3507
-2.7285e-12 | 12.7757 | -4.1837e-11 | 12.7557
-2.9559e-12 | 13.1817 | -4.4338e-11 | 13.1617
-4.3201e-12 | 13.5858 | -4.411e-11 | 13.5688
-1.819e-12 | 14.0018 | -4.3883e-11 | 13.9738
-3.1832e-12 | 14.4048 | -4.411e-11 | 14.3808
-2.5011e-12 | 14.8099 | -3.979e-11 | 14.7849
-2.7285e-12 | 15.2139 | -4.4565e-11 | 15.1919
-2.7285e-12 | 15.6199 | -3.7744e-11 | 15.5979
-3.8654e-12 | 16.0259 | -3.9336e-11 | 16.0039
-5.6843e-12 | 16.4359 | -4.411e-11 | 16.4089
-1.3642e-12 | 16.842 | -4.07e-11 | 16.814
-2.5011e-12 | 17.247 | -3.7971e-11 | 17.218
-1.819e-12 | 17.653 | -4.07e-11 | 17.623
-5.2296e-12 | 18.058 | -4.5475e-11 | 18.029
-2.9559e-12 | 18.4631 | -4.4338e-11 | 18.434
-2.7285e-12 | 18.8671 | -4.1155e-11 | 18.8391
-2.9559e-12 | 19.2981 | -4.5247e-11 | 19.2451
-4.5475e-12 | 19.7031 | -3.7744e-11 | 19.6511
-3.4106e-12 | 20.1081 | -4.5702e-11 | 20.0561
-4.3201e-12 | 20.5132 | -4.2974e-11 | 20.4632
-2.9559e-12 | 20.9232 | -4.5702e-11 | 20.8682
-2.9559e-12 | 21.3282 | -4.5247e-11 | 21.2732
-3.8654e-12 | 21.7362 | -4.2519e-11 | 21.6782
-3.4106e-12 | 22.1423 | -4.0245e-11 | 22.0843
-3.4106e-12 | 22.5503 | -4.2519e-11 | 22.4893
-3.8654e-12 | 22.9553 | -4.1155e-11 | 22.8943
-4.0927e-12 | 23.3633 | -4.0018e-11 | 23.3013
-2.7285e-12 | 23.7704 | -4.0018e-11 | 23.7064
-4.5475e-12 | 24.1754 | -4.1382e-11 | 24.1124
-2.9559e-12 | 24.5834 | -3.7517e-11 | 24.5194
-2.7285e-12 | 25.0154 | -3.6835e-11 | 24.9244
-5.9117e-12 | 25.4204 | -3.7517e-11 | 25.3304
-5.0022e-12 | 25.8255 | -4.2064e-11 | 25.7365
-3.4106e-12 | 26.2305 | -4.7976e-11 | 26.1415
-5.0022e-12 | 26.6345 | -4.5929e-11 | 26.5475
-5.2296e-12 | 27.0386 | -4.0018e-11 | 26.9535
-3.8654e-12 | 27.4436 | -4.1155e-11 | 27.3586
-4.5475e-12 | 27.8626 | -4.2064e-11 | 27.7646
-2.7285e-12 | 28.2676 | -3.9108e-11 | 28.1706
-5.0022e-12 | 28.6746 | -4.7976e-11 | 28.5756
-5.9117e-12 | 29.0787 | -4.0245e-11 | 28.9827
-3.1832e-12 | 29.4837 | -4.1155e-11 | 29.3877
-3.4106e-12 | 29.9137 | -4.2519e-11 | 29.7927
-4.3201e-12 | 30.3227 | -4.2974e-11 | 30.1987
-3.1832e-12 | 30.7278 | -3.9336e-11 | 30.6038
-4.0927e-12 | 31.1578 | -4.0473e-11 | 31.0088
-4.0927e-12 | 31.5628 | -3.6152e-11 | 31.4148
-1.3642e-12 | 31.9678 | -4.1382e-11 | 31.8198
-3.1832e-12 | 32.3749 | -3.9108e-11 | 32.2258
-3.4106e-12 | 32.7799 | -4.2519e-11 | 32.6299
-2.0464e-12 | 33.1879 | -4.2064e-11 | 33.0349
-5.457e-12 | 33.5979 | -3.7971e-11 | 33.4389
-2.5011e-12 | 34.0039 | -4.3656e-11 | 33.8439
-1.819e-12 | 34.412 | -4.5929e-11 | 34.25
-5.457e-12 | 34.818 | -3.6835e-11 | 34.654
-3.1832e-12 | 35.226 | -3.8199e-11 | 35.06
-2.7285e-12 | 35.632 | -3.9336e-11 | 35.465
-5.457e-12 | 36.0361 | -4.1609e-11 | 35.8721
-5.0022e-12 | 36.4421 | -3.9108e-11 | 36.2771
-4.0927e-12 | 36.8471 | -4.1837e-11 | 36.6831
-3.8654e-12 | 37.2521 | -3.7971e-11 | 37.0881
-5.9117e-12 | 37.6572 | -3.3424e-11 | 37.4941
-2.0464e-12 | 38.0632 | -4.2974e-11 | 37.8992
-4.5475e-13 | 38.4682 | -4.0473e-11 | 38.3052
-5.2296e-12 | 38.8732 | -4.1155e-11 | 38.7102
-6.5938e-12 | 39.2792 | -4.2746e-11 | 39.1162
-3.8654e-12 | 39.6843 | -4.8431e-11 | 39.5213
-2.5011e-12 | 40.0893 | -4.1609e-11 | 39.9273
-1.5916e-12 | 40.4943 | -4.2519e-11 | 40.3333
-2.0464e-12 | 40.9113 | -3.8199e-11 | 40.7383
-6.5938e-12 | 41.3164 | -4.8203e-11 | 41.1433
-2.7285e-12 | 41.7214 | -4.3201e-11 | 41.5474
-1.5916e-12 | 42.1274 | -4.1155e-11 | 41.9524
-4.0927e-12 | 42.5334 | -4.1837e-11 | 42.3574
-1.2051e-11 | 42.9395 | -4.07e-11 | 42.7625
-3.4106e-12 | 43.3445 | -4.5702e-11 | 43.1685
-4.5475e-12 | 43.7495 | -4.0245e-11 | 43.5735
-4.5475e-13 | 44.1645 | -4.2974e-11 | 43.9785
-1.819e-12 | 44.5686 | -4.5475e-11 | 44.3845
-5.0022e-12 | 44.9726 | -4.1837e-11 | 44.7896
-5.2296e-12 | 45.3786 | -3.8654e-11 | 45.1946
-3.1832e-12 | 45.7836 | -4.2974e-11 | 45.5996
0 | 46.1876 | -5.0477e-11 | 46.0056
-3.1832e-12 | 46.5907 | -3.9108e-11 | 46.4106
-8.1855e-12 | 46.9947 | -3.8654e-11 | 46.8167
-4.0927e-12 | 47.3997 | -4.2519e-11 | 47.2217
4.5475e-13 | 47.8047 | -4.4338e-11 | 47.6267
-5.6843e-12 | 48.2108 | -4.0018e-11 | 48.0328
-5.457e-12 | 48.6158 | -4.4565e-11 | 48.4378
-1.3642e-12 | 49.0208 | -4.4565e-11 | 48.8428
-1.5916e-12 | 49.4258 | -4.4565e-11 | 49.2478
-3.1832e-12 | 49.8319 | -3.9108e-11 | 49.6548
-1.3642e-12 | 50.2369 | -4.0245e-11 | 50.0589
-2.9559e-12 | 50.6459 | -4.6839e-11 | 50.4639
-5.6843e-12 | 51.0519 | -3.979e-11 | 50.8689
-1.819e-12 | 51.4579 | -3.9108e-11 | 51.2729
-2.0464e-12 | 51.865 | -4.6612e-11 | 51.678
-6.5938e-12 | 52.27 | -4.5475e-11 | 52.084
-5.6843e-12 | 52.675 | -4.3201e-11 | 52.489
-6.8212e-12 | 53.2951 | -4.1382e-11 | 52.895
-3.8654e-12 | 53.7591 | -3.8881e-11 | 53.3021
-3.8654e-12 | 54.1731 | -4.0245e-11 | 53.7071
-2.7285e-12 | 54.5781 | -4.3656e-11 | 54.1121
-2.5011e-12 | 54.9851 | -4.3656e-11 | 54.5181
-6.3665e-12 | 55.3902 | -4.0018e-11 | 54.9241
-5.9117e-12 | 55.8042 | -4.1609e-11 | 55.3292
-2.5011e-12 | 56.2182 | -3.7289e-11 | 55.7362
-1.819e-12 | 56.6222 | -4.2519e-11 | 56.1402
-3.4106e-12 | 57.0273 | -4.0018e-11 | 56.5452
-3.4106e-12 | 57.4323 | -4.1609e-11 | 56.9503
-1.819e-12 | 57.8393 | -4.1382e-11 | 57.3553
-3.4106e-12 | 58.2473 | -4.1837e-11 | 57.7603
-4.3201e-12 | 58.6534 | -3.979e-11 | 58.1653
-3.1832e-12 | 59.0594 | -4.1609e-11 | 58.5694
-2.2737e-13 | 59.4644 | -4.0018e-11 | 58.9744
-4.3201e-12 | 59.8684 | -4.4565e-11 | 59.3804
-4.0927e-12 | 60.2745 | -4.1382e-11 | 59.7854
-1.819e-12 | 60.6805 | -4.2064e-11 | 60.1914
-2.5011e-12 | 61.1045 | -4.2974e-11 | 60.5965
-1.5916e-12 | 61.5195 | -4.9113e-11 | 61.0015
-2.0464e-12 | 61.9245 | -3.6835e-11 | 61.4055
-5.0022e-12 | 62.3296 | -3.8654e-11 | 61.8115
-2.7285e-12 | 62.7356 | -4.4338e-11 | 62.2166
-7.276e-12 | 63.1496 | -3.8881e-11 | 62.6226
-6.8212e-13 | 63.5556 | -4.0018e-11 | 63.0286
-5.457e-12 | 63.9667 | -4.5247e-11 | 63.4336
-5.0022e-12 | 64.3707 | -4.07e-11 | 63.8387
-4.0927e-12 | 64.7767 | -4.2519e-11 | 64.2437
-3.4106e-12 | 65.1957 | -4.2519e-11 | 64.6497
-1.5916e-12 | 65.6008 | -4.9113e-11 | 65.0537
-4.5475e-12 | 66.0068 | -4.2064e-11 | 65.4587
-4.5475e-12 | 66.4118 | -4.411e-11 | 65.8638
-2.5011e-12 | 66.8168 | -4.6839e-11 | 66.2688
-3.1832e-12 | 67.2218 | -4.3656e-11 | 66.6738
-6.5938e-12 | 67.6279 | -4.5929e-11 | 67.0788
-2.5011e-12 | 68.0339 | -4.3201e-11 | 67.4839
-2.0464e-12 | 68.4389 | -3.9108e-11 | 67.8899
-5.2296e-12 | 68.8449 | -4.0018e-11 | 68.2959
-7.9581e-12 | 69.249 | -4.0018e-11 | 68.6999
-1.3642e-12 | 69.654 | -3.8654e-11 | 69.106
-5.0022e-12 | 70.083 | -4.411e-11 | 69.512
-6.3665e-12 | 70.489 | -4.2519e-11 | 69.917
-2.0464e-12 | 70.894 | -3.638e-11 | 70.322
-1.819e-12 | 71.2991 | -3.7289e-11 | 70.728
-6.3665e-12 | 71.7221 | -4.6612e-11 | 71.1341
-6.3665e-12 | 72.1551 | -4.2746e-11 | 71.5391
-5.9117e-12 | 72.6091 | -4.1155e-11 | 71.9441
-3.4106e-12 | 73.0132 | -4.0473e-11 | 72.3491
-1.5916e-12 | 73.4162 | -4.5929e-11 | 72.7542
-2.5011e-12 | 73.8222 | -4.0473e-11 | 73.1592
-5.2296e-12 | 74.2272 | -3.5243e-11 | 73.5642
-1.3642e-12 | 74.6323 | -3.8881e-11 | 73.9692
-4.5475e-12 | 75.0453 | -4.07e-11 | 74.3762
-4.0927e-12 | 75.4493 | -4.07e-11 | 74.7813
-2.0464e-12 | 75.8563 | -4.1837e-11 | 75.1863
-4.0927e-12 | 76.2614 | -4.0245e-11 | 75.5913
-2.9559e-12 | 76.6754 | -4.411e-11 | 75.9973
-3.4106e-12 | 77.0864 | -4.0245e-11 | 76.4034
-2.0464e-12 | 77.4954 | -4.0473e-11 | 76.8084
-4.0927e-12 | 77.9015 | -4.07e-11 | 77.2154
-6.3665e-12 | 78.3065 | -4.1382e-11 | 77.6204
-4.5475e-12 | 78.7115 | -3.4106e-11 | 78.0255
0 | 79.1165 | -4.07e-11 | 78.4295
-2.0464e-12 | 79.5186 | -4.411e-11 | 78.8355
-6.3665e-12 | 79.9236 | -4.0245e-11 | 79.2415
-4.5475e-12 | 80.3296 | -4.0245e-11 | 79.6476
-2.7285e-12 | 80.7346 | -3.9108e-11 | 80.0526
-5.457e-12 | 81.1426 | -4.5702e-11 | 80.4576
-6.3665e-12 | 81.5477 | -4.2064e-11 | 80.8626
-6.8212e-13 | 81.9527 | -4.502e-11 | 81.2686
-4.3201e-12 | 82.3587 | -4.9113e-11 | 81.6737
-6.3665e-12 | 82.7657 | -4.5702e-11 | 82.0787
-4.5475e-12 | 83.1718 | -3.7517e-11 | 82.4827
-2.2737e-13 | 83.5768 | -4.7066e-11 | 82.8897
-4.5475e-12 | 83.9818 | -3.8654e-11 | 83.2958

## 600uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-3.5393e-12 | 1.4287e-12 | -5.6599e-11 | 4.5913e-12
GEMs On |  | GEMs On
-2.7285e-12 | 0.31302 | -5.4797e-11 | 0.31502
-2.9559e-12 | 0.99806 | -5.6843e-11 | 1.0001
-2.7285e-12 | 1.4041 | -5.6843e-11 | 1.4051
-1.5916e-12 | 1.8141 | -5.2523e-11 | 1.8121
-3.1832e-12 | 2.2201 | -6.753e-11 | 2.2221
-5.2296e-12 | 2.6242 | -5.7298e-11 | 2.6292
-1.819e-12 | 3.0322 | -5.2523e-11 | 3.0352
-5.9117e-12 | 3.4412 | -5.1841e-11 | 3.4422
-5.9117e-12 | 3.8452 | -5.4797e-11 | 3.8492
-1.819e-12 | 4.2522 | -5.5479e-11 | 4.2542
-4.0927e-12 | 4.6573 | -5.6843e-11 | 4.6613
-6.5938e-12 | 5.0643 | -6.2073e-11 | 5.0663
-3.1832e-12 | 5.4703 | -5.3888e-11 | 5.4723
-1.3642e-12 | 5.8763 | -4.9567e-11 | 5.8783
-7.7307e-12 | 6.2834 | -6.0254e-11 | 6.2844
-1.1369e-12 | 6.6884 | -5.2523e-11 | 6.6904
-1.3642e-12 | 7.0954 | -6.0709e-11 | 7.0964
-5.2296e-12 | 7.5034 | -6.23e-11 | 7.5054
-1.819e-12 | 7.9075 | -5.8435e-11 | 7.9115
-1.1369e-12 | 8.3155 | -5.0477e-11 | 8.3175
-5.0022e-12 | 8.7215 | -5.6389e-11 | 8.7395
-5.6843e-12 | 9.1255 | -6.23e-11 | 9.1445
-6.8212e-13 | 9.5305 | -5.9345e-11 | 9.5495
-1.1369e-12 | 9.9376 | -5.2069e-11 | 9.9556
-7.9581e-12 | 10.3416 | -4.8431e-11 | 10.3616
-1.819e-12 | 10.7486 | -5.3205e-11 | 10.7656
-2.5011e-12 | 11.1556 | -5.8208e-11 | 11.1736
-5.0022e-12 | 11.5627 | -5.5479e-11 | 11.5787
-3.1832e-12 | 11.9677 | -5.8662e-11 | 11.9857
-5.0022e-12 | 12.3737 | -6.2073e-11 | 12.3917
-5.457e-12 | 12.7847 | -5.7298e-11 | 12.7967
-2.5011e-12 | 13.1907 | -6.3437e-11 | 13.2027
-1.5916e-12 | 13.5958 | -5.0932e-11 | 13.6088
-2.5011e-12 | 14.0138 | -5.6843e-11 | 14.0148
-5.0022e-12 | 14.4198 | -5.5024e-11 | 14.4198
-5.9117e-12 | 14.8239 | -5.5024e-11 | 14.8248
-1.819e-12 | 15.2279 | -5.2978e-11 | 15.2489
-6.8212e-13 | 15.6339 | -5.8208e-11 | 15.6559
-5.457e-12 | 16.0379 | -5.8435e-11 | 16.0639
-4.3201e-12 | 16.4429 | -5.9345e-11 | 16.4689
-1.1369e-12 | 16.848 | -4.9113e-11 | 16.875
-1.5916e-12 | 17.252 | -5.5934e-11 | 17.28
-2.0464e-12 | 17.657 | -5.5707e-11 | 17.69
-4.5475e-12 | 18.065 | -5.2978e-11 | 18.1
-1.3642e-12 | 18.4691 | -6.662e-11 | 18.5051
-4.5475e-12 | 18.8751 | -5.5707e-11 | 18.9111
-7.0486e-12 | 19.2801 | -5.457e-11 | 19.3371
-3.4106e-12 | 19.6841 | -4.7976e-11 | 19.7441
-1.5916e-12 | 20.0911 | -5.5479e-11 | 20.1491
-4.5475e-12 | 20.4962 | -5.2296e-11 | 20.5542
-2.7285e-12 | 20.9102 | -5.8208e-11 | 20.9632
-4.3201e-12 | 21.3152 | -5.366e-11 | 21.3702
-1.1369e-12 | 21.7252 | -4.6612e-11 | 21.7752
-5.0022e-12 | 22.1343 | -5.2978e-11 | 22.1903
-4.5475e-12 | 22.5413 | -6.3437e-11 | 22.6013
-2.7285e-12 | 22.9453 | -6.0254e-11 | 23.0063
-1.5916e-12 | 23.3513 | -5.4342e-11 | 23.4113
-5.457e-12 | 23.7564 | -5.457e-11 | 23.8184
-4.5475e-12 | 24.1634 | -5.6389e-11 | 24.2254
-1.5916e-12 | 24.5684 | -5.0477e-11 | 24.6334
-4.5475e-12 | 24.9744 | -6.321e-11 | 25.0414
-5.6843e-12 | 25.3804 | -6.3665e-11 | 25.4485
-4.0927e-12 | 25.7845 | -5.7298e-11 | 25.8545
-2.7285e-12 | 26.1885 | -4.0927e-11 | 26.2615
-3.1832e-12 | 26.5985 | -6.23e-11 | 26.6675
-3.4106e-12 | 27.0175 | -5.4342e-11 | 27.0846
-1.819e-12 | 27.4236 | -5.1614e-11 | 27.5116
-2.7285e-12 | 27.8326 | -5.5479e-11 | 27.9266
-2.7285e-12 | 28.2376 | -5.3205e-11 | 28.3446
-5.0022e-12 | 28.6426 | -5.8208e-11 | 28.7506
-5.2296e-12 | 29.0647 | -5.5479e-11 | 29.1547
-4.0927e-12 | 29.4697 | -5.8662e-11 | 29.5607
-2.0464e-12 | 29.8747 | -5.3433e-11 | 29.9667
-3.1832e-12 | 30.2917 | -6.3892e-11 | 30.3737
-3.8654e-12 | 30.6968 | -6.2755e-11 | 30.7818
-3.4106e-12 | 31.1028 | -5.1841e-11 | 31.1998
-2.9559e-12 | 31.5098 | -5.0704e-11 | 31.6058
-2.9559e-12 | 31.9158 | -6.1846e-11 | 32.0118
-3.8654e-12 | 32.3338 | -5.3888e-11 | 32.4158
-5.0022e-12 | 32.7389 | -5.4797e-11 | 32.8219
-3.8654e-12 | 33.1439 | -5.2069e-11 | 33.2279
-2.7285e-12 | 33.5479 | -5.1841e-11 | 33.6339
-3.8654e-12 | 33.9539 | -4.9567e-11 | 34.041
-5.2296e-12 | 34.359 | -5.3205e-11 | 34.448
-2.5011e-12 | 34.763 | -5.6161e-11 | 34.854
-5.2296e-12 | 35.172 | -6.0936e-11 | 35.258
-4.0927e-12 | 35.58 | -5.8435e-11 | 35.665
-4.5475e-12 | 35.9861 | -6.3437e-11 | 36.0731
-5.0022e-12 | 36.3911 | -6.7757e-11 | 36.4791
-3.4106e-12 | 36.7971 | -5.2523e-11 | 36.8861
-3.4106e-12 | 37.2021 | -6.7075e-11 | 37.2921
-2.7285e-12 | 37.6082 | -5.3888e-11 | 37.6982
-2.9559e-12 | 38.0142 | -5.7071e-11 | 38.1122
-4.5475e-12 | 38.4272 | -5.0477e-11 | 38.5182
-2.7285e-12 | 38.8352 | -5.6389e-11 | 38.9232
-2.0464e-12 | 39.2442 | -5.7526e-11 | 39.3332
-3.4106e-12 | 39.6513 | -5.5479e-11 | 39.7423
-3.4106e-12 | 40.0583 | -5.2978e-11 | 40.1553
-3.4106e-12 | 40.4753 | -5.6389e-11 | 40.5653
-3.4106e-12 | 40.8803 | -5.6843e-11 | 40.9703
-2.7285e-12 | 41.2844 | -5.5934e-11 | 41.3874
-5.0022e-12 | 41.6884 | -5.8662e-11 | 41.7934
-4.5475e-12 | 42.0944 | -5.366e-11 | 42.2554
-3.1832e-12 | 42.5054 | -5.7071e-11 | 42.6624
-4.5475e-12 | 42.9145 | -5.4342e-11 | 43.0685
-4.3201e-12 | 43.3205 | -4.6612e-11 | 43.4735
-4.3201e-12 | 43.7255 | -5.8435e-11 | 43.8815
-2.9559e-12 | 44.1295 | -5.5707e-11 | 44.2915
-5.9117e-12 | 44.5346 | -4.8203e-11 | 44.6986
-4.3201e-12 | 44.9406 | -5.1614e-11 | 45.1036
-1.5916e-12 | 45.3466 | -5.5479e-11 | 45.5106
-2.7285e-12 | 45.7506 | -5.5934e-11 | 45.9166
-2.9559e-12 | 46.1546 | -5.7753e-11 | 46.3227
-2.9559e-12 | 46.5597 | -5.457e-11 | 46.7337
-3.1832e-12 | 46.9647 | -6.321e-11 | 47.1407
-3.1832e-12 | 47.3707 | -6.5256e-11 | 47.5507
-2.9559e-12 | 47.7767 | -5.7071e-11 | 47.9617
-2.7285e-12 | 48.1818 | -5.9799e-11 | 48.3688
-4.0927e-12 | 48.5868 | -5.457e-11 | 48.7818
-4.3201e-12 | 48.9908 | -6.4574e-11 | 49.1878
-4.0927e-12 | 49.3968 | -4.9567e-11 | 49.6088
-2.9559e-12 | 49.8019 | -5.5934e-11 | 50.0149
-3.4106e-12 | 50.2089 | -6.2528e-11 | 50.4199
-2.7285e-12 | 50.6159 | -5.2069e-11 | 50.8249
-2.9559e-12 | 51.0359 | -4.9795e-11 | 51.2309
-2.7285e-12 | 51.4419 | -6.8439e-11 | 51.636
-2.9559e-12 | 51.847 | -5.7526e-11 | 52.041
-5.2296e-12 | 52.336 | -5.0932e-11 | 52.447
-4.3201e-12 | 53.018 | -4.7748e-11 | 52.853
-3.8654e-12 | 53.4251 | -6.662e-11 | 53.258
-3.8654e-12 | 53.8381 | -6.1164e-11 | 53.6651
-3.4106e-12 | 54.2451 | -5.889e-11 | 54.0711
-4.5475e-12 | 54.6511 | -6.0709e-11 | 54.4781
-3.4106e-12 | 55.0551 | -5.3888e-11 | 54.8931
-1.819e-12 | 55.5172 | -6.23e-11 | 55.3152
-1.5916e-12 | 55.9232 | -6.321e-11 | 55.7222
-5.2296e-12 | 56.3292 | -5.5024e-11 | 56.1272
-3.1832e-12 | 56.7332 | -5.6389e-11 | 56.5332
-1.1369e-12 | 57.1403 | -6.2755e-11 | 56.9393
-2.5011e-12 | 57.5443 | -5.2069e-11 | 57.3453
-3.4106e-12 | 57.9493 | -5.8662e-11 | 57.7503
-2.7285e-12 | 58.3543 | -5.9799e-11 | 58.1563
-4.5475e-12 | 58.7614 | -5.8662e-11 | 58.5643
-4.0927e-12 | 59.1704 | -5.9799e-11 | 58.9704
-2.9559e-12 | 59.5754 | -5.6389e-11 | 59.3774
-4.0927e-12 | 59.9804 | -5.5024e-11 | 59.7824
-2.5011e-12 | 60.3865 | -5.1614e-11 | 60.1884
-4.3201e-12 | 60.8065 | -6.2755e-11 | 60.5935
-2.5011e-12 | 61.2125 | -5.8435e-11 | 60.9985
-3.1832e-12 | 61.6185 | -5.6843e-11 | 61.4035
-4.5475e-12 | 62.0226 | -5.2069e-11 | 61.8115
-5.0022e-12 | 62.4276 | -6.0254e-11 | 62.2166
-2.9559e-12 | 62.8336 | -5.7753e-11 | 62.6216
-3.8654e-12 | 63.2416 | -5.9345e-11 | 63.0266
-2.0464e-12 | 63.6516 | -5.6161e-11 | 63.4316
-2.0464e-12 | 64.0557 | -5.7298e-11 | 63.8387
-1.1369e-12 | 64.4607 | -5.9572e-11 | 64.2547
-2.7285e-12 | 64.8667 | -5.3205e-11 | 64.6617
-4.3201e-12 | 65.2737 | -5.0477e-11 | 65.0677
-2.7285e-12 | 65.6778 | -5.7753e-11 | 65.4747
-1.1369e-12 | 66.0848 | -5.5024e-11 | 65.8808
-4.0927e-12 | 66.4908 | -5.7526e-11 | 66.2878
-3.4106e-12 | 66.8978 | -5.5479e-11 | 66.7088
-2.7285e-12 | 67.3029 | -5.2523e-11 | 67.1148
-3.1832e-12 | 67.7279 | -5.1614e-11 | 67.5199
-6.8212e-12 | 68.1489 | -5.889e-11 | 67.9299
-4.0927e-12 | 68.5549 | -5.9572e-11 | 68.3369
-1.1369e-12 | 68.9609 | -5.5934e-11 | 68.7419
-5.2296e-12 | 69.367 | -5.4797e-11 | 69.1489
-4.0927e-12 | 69.772 | -5.9799e-11 | 69.557
-2.9559e-12 | 70.177 | -5.8662e-11 | 69.963
-5.2296e-12 | 70.583 | -5.0704e-11 | 70.368
-6.5938e-12 | 70.9901 | -5.5479e-11 | 70.773
-1.819e-12 | 71.3971 | -5.9572e-11 | 71.1791
-1.3642e-12 | 71.8031 | -5.0704e-11 | 71.5851
-4.0927e-12 | 72.2081 | -5.8208e-11 | 71.9891
-5.457e-12 | 72.6132 | -5.2978e-11 | 72.3991
-1.5916e-12 | 73.0182 | -5.3888e-11 | 72.8042
-1.5916e-12 | 73.4242 | -5.889e-11 | 73.2102
-3.8654e-12 | 73.8462 | -5.4342e-11 | 73.6162
-4.0927e-12 | 74.2502 | -5.2296e-11 | 74.0212
-4.0927e-12 | 74.6593 | -6.1164e-11 | 74.4313
-5.457e-12 | 75.0703 | -5.9799e-11 | 74.8373
-4.3201e-12 | 75.4763 | -5.5024e-11 | 75.2433
-1.819e-12 | 76.2834 | -6.3892e-11 | 75.6503
-3.4106e-12 | 76.6874 | -6.23e-11 | 76.0563
-3.1832e-12 | 77.1094 | -5.8208e-11 | 76.4614
-3.1832e-12 | 77.5454 | -5.9799e-11 | 76.8754
-7.0486e-12 | 77.9505 | -6.3892e-11 | 77.2804
-3.8654e-12 | 78.3555 | -6.4119e-11 | 77.6864
-3.8654e-12 | 78.7605 | -5.9799e-11 | 78.0935
-3.1832e-12 | 79.1665 | -5.889e-11 | 78.4975
-6.8212e-12 | 79.5716 | -5.8208e-11 | 78.9045
-2.7285e-12 | 79.9766 | -6.0254e-11 | 79.3105
-2.5011e-12 | 80.3806 | -5.8435e-11 | 79.7156
-2.9559e-12 | 80.7856 | -5.2978e-11 | 80.1206
-5.6843e-12 | 81.1896 | -5.5707e-11 | 80.5266
-4.3201e-12 | 81.5947 | -5.2069e-11 | 80.9316
-3.4106e-12 | 81.9997 | -6.4119e-11 | 81.3367
-4.0927e-12 | 82.4117 | -5.1841e-11 | 81.7407
-4.3201e-12 | 82.8157 | -4.7976e-11 | 82.1467
-2.7285e-12 | 83.2218 | -5.6389e-11 | 82.5517
-1.819e-12 | 83.6258 | -6.0709e-11 | 82.9567
-4.5475e-12 | 84.0308 | -6.2073e-11 | 83.3618
-5.2296e-12 | 84.4338 | -5.5024e-11 | 83.7668
-5.2296e-12 | 84.8388 | -5.5934e-11 | 84.1718
-3.4106e-12 | 85.2449
-1.5916e-12 | 85.6499
-2.9559e-12 | 86.0549
-3.8654e-12 | 86.459
-1.819e-12 | 86.863
-4.3201e-12 | 87.267
-5.457e-12 | 87.672
-3.8654e-12 | 88.077
-2.0464e-12 | 88.4821
-4.0927e-12 | 88.8861
-1.819e-12 | 89.2891
-2.5011e-12 | 89.6931
-5.6843e-12 | 90.0982
-4.3201e-12 | 90.5032
-2.9559e-12 | 90.9082
-2.9559e-12 | 91.3152
-4.5475e-12 | 91.7202
-3.1832e-12 | 92.1253
-4.3201e-12 | 92.5303
-2.0464e-12 | 92.9343
-5.2296e-12 | 93.3403
-4.0927e-12 | 93.7444

## 610uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-3.7684e-12 | 1.1363e-12 | -7.798e-11 | 6.0572e-12
GEMs On |  | GEMs On
-2.7285e-12 | 0.30902 | -7.9808e-11 | 0.30902
-4.7748e-12 | 0.99406 | -7.8444e-11 | 0.99406
-4.7748e-12 | 1.4011 | -7.8217e-11 | 1.3991
-4.5475e-12 | 1.8061 | -7.8899e-11 | 1.8041
-3.4106e-12 | 2.2121 | -7.8217e-11 | 2.2121
-5.6843e-12 | 2.6181 | -7.6625e-11 | 2.6191
-3.8654e-12 | 3.0222 | -7.3442e-11 | 3.0262
-3.1832e-12 | 3.4282 | -7.8444e-11 | 3.4352
-3.638e-12 | 3.8332 | -7.3214e-11 | 3.8422
-3.638e-12 | 4.2392 | -8.049e-11 | 4.2482
-3.638e-12 | 4.6443 | -7.5715e-11 | 4.6533
-2.2737e-12 | 5.0483 | -7.5715e-11 | 5.0663
-2.2737e-12 | 5.4543 | -7.7762e-11 | 5.4743
-3.8654e-12 | 5.8593 | -8.4128e-11 | 5.8803
-4.0927e-12 | 6.2634 | -7.4579e-11 | 6.2864
-3.1832e-12 | 6.6684 | -7.6625e-11 | 6.6914
-4.0927e-12 | 7.0724 | -8.2764e-11 | 7.0974
-5.6843e-12 | 7.4774 | -8.2764e-11 | 7.5034
-5.6843e-12 | 7.8825 | -7.8671e-11 | 7.9095
-3.1832e-12 | 8.2875 | -7.2532e-11 | 8.3165
-3.638e-12 | 8.6935 | -9.3905e-11 | 8.7225
-2.7285e-12 | 9.0975 | -7.5943e-11 | 9.1295
-4.5475e-12 | 9.5015 | -7.7307e-11 | 9.5335
-2.5011e-12 | 9.9066 | -6.9349e-11 | 9.9386
-2.2737e-12 | 10.3116 | -7.7762e-11 | 10.3456
-4.0927e-12 | 10.7166 | -8.2537e-11 | 10.7566
-4.7748e-12 | 11.1206 | -8.2764e-11 | 11.1936
-3.638e-12 | 11.5257 | -8.1627e-11 | 11.6087
-4.0927e-12 | 11.9307 | -9.1859e-11 | 12.0147
-5.0022e-12 | 12.3357 | -7.5943e-11 | 12.4207
-5.0022e-12 | 12.7407 | -7.708e-11 | 12.8267
-4.0927e-12 | 13.1448 | -8.6402e-11 | 13.2418
-3.638e-12 | 13.5498 | -7.2532e-11 | 13.6538
-2.7285e-12 | 13.9548 | -7.8217e-11 | 14.0598
-5.2296e-12 | 14.3588 | -7.708e-11 | 14.4858
-2.7285e-12 | 14.7628 | -8.0036e-11 | 14.8909
-3.4106e-12 | 15.1679 | -7.7307e-11 | 15.3129
-4.5475e-12 | 15.5719 | -7.5715e-11 | 15.7179
-4.5475e-12 | 15.9779 | -6.753e-11 | 16.1259
-5.9117e-12 | 16.3819 | -8.3901e-11 | 16.5329
-2.5011e-12 | 16.786 | -8.1627e-11 | 16.938
-2.7285e-12 | 17.19 | -6.7075e-11 | 17.345
-3.8654e-12 | 17.595 | -8.3901e-11 | 17.754
-3.1832e-12 | 17.999 | -7.4579e-11 | 18.16
-4.0927e-12 | 18.4041 | -8.049e-11 | 18.5671
-2.7285e-12 | 18.8091 | -7.6398e-11 | 18.9741
-2.2737e-12 | 19.2141 | -7.9126e-11 | 19.3791
-6.1391e-12 | 19.6181 | -8.5038e-11 | 19.7841
-2.2737e-12 | 20.0221 | -7.6625e-11 | 20.1912
-3.638e-12 | 20.4262 | -8.6629e-11 | 20.6002
-2.2737e-12 | 20.8312 | -7.3442e-11 | 21.0062
-4.0927e-12 | 21.2352 | -7.3214e-11 | 21.4142
-2.7285e-12 | 21.6402 | -7.2532e-11 | 21.8203
-2.5011e-12 | 22.0453 | -7.5715e-11 | 22.2323
-2.2737e-12 | 22.4503 | -8.0945e-11 | 22.6373
-3.1832e-12 | 22.8553 | -7.4806e-11 | 23.0513
-4.0927e-12 | 23.2603 | -7.7534e-11 | 23.4653
-4.0927e-12 | 23.6654 | -6.2528e-11 | 23.8704
-3.1832e-12 | 24.0694 | -8.1172e-11 | 24.2794
-4.0927e-12 | 24.4734 | -7.1168e-11 | 24.7004
-2.7285e-12 | 24.8774 | -7.2532e-11 | 25.1074
-3.4106e-12 | 25.2834 | -7.276e-11 | 25.5135
-2.7285e-12 | 25.6885 | -7.2305e-11 | 25.9245
-2.7285e-12 | 26.0925 | -7.9126e-11 | 26.3575
-4.5475e-12 | 26.4965 | -7.1395e-11 | 26.7635
-5.6843e-12 | 26.9015 | -8.5947e-11 | 27.1685
-4.5475e-12 | 27.3066 | -7.2305e-11 | 27.5746
-3.1832e-12 | 27.7116 | -8.2537e-11 | 27.9796
-3.4106e-12 | 28.1156 | -8.0036e-11 | 28.3846
-4.0927e-12 | 28.5216 | -8.0036e-11 | 28.7936
-2.2737e-12 | 28.9246 | -8.14e-11 | 29.1987
-3.8654e-12 | 29.3297 | -8.3901e-11 | 29.6067
-1.819e-12 | 29.7347 | -7.7534e-11 | 30.0127
-1.1369e-12 | 30.1397 | -8.7539e-11 | 30.4177
-2.7285e-12 | 30.5448 | -7.2305e-11 | 30.8238
-5.0022e-12 | 30.9488 | -7.0031e-11 | 31.2308
-5.0022e-12 | 31.3528 | -6.5938e-11 | 31.6538
-4.5475e-12 | 31.7588 | -7.6398e-11 | 32.0598
-3.638e-12 | 32.1638 | -8.0945e-11 | 32.4679
-2.2737e-12 | 32.5679 | -7.2532e-11 | 32.8739
-4.7748e-12 | 32.9729 | -7.9581e-11 | 33.2789
-4.7748e-12 | 33.3779 | -8.14e-11 | 33.6859
-3.638e-12 | 33.7819 | -7.185e-11 | 34.0909
-3.638e-12 | 34.187 | -7.7534e-11 | 34.497
-2.0464e-12 | 34.592 | -8.14e-11 | 34.913
-3.4106e-12 | 34.997 | -8.5947e-11 | 35.319
-3.638e-12 | 35.401 | -7.5715e-11 | 35.725
-2.2737e-12 | 35.806 | -8.6857e-11 | 36.1311
-4.7748e-12 | 36.2111 | -7.617e-11 | 36.5361
-2.5011e-12 | 36.6161 | -8.4128e-11 | 36.9421
-2.5011e-12 | 37.0211 | -8.2991e-11 | 37.3481
-3.8654e-12 | 37.4261 | -7.7762e-11 | 37.7532
-3.8654e-12 | 37.8322 | -6.9804e-11 | 38.1592
-2.7285e-12 | 38.2362 | -9.1404e-11 | 38.5642
-4.5475e-12 | 38.6402 | -7.3669e-11 | 38.9692
-5.0022e-12 | 39.0442 | -7.7307e-11 | 39.3753
-5.0022e-12 | 39.4503 | -7.708e-11 | 39.7803
-3.8654e-12 | 39.8553 | -7.3214e-11 | 40.2053
-3.8654e-12 | 40.2603 | -7.9581e-11 | 40.6173
-3.4106e-12 | 40.6653 | -7.8899e-11 | 41.0243
-2.5011e-12 | 41.0694 | -7.5715e-11 | 41.4324
-6.1391e-12 | 41.4744 | -8.1172e-11 | 41.8414
-4.7748e-12 | 41.8794 | -7.0258e-11 | 42.2484
-4.0927e-12 | 42.2834 | -7.0031e-11 | 42.6574
-2.2737e-12 | 42.6884 | -9.3678e-11 | 43.0625
-3.638e-12 | 43.0935 | -7.3669e-11 | 43.4685
-3.8654e-12 | 43.4975 | -7.708e-11 | 43.8735
-3.4106e-12 | 43.9025 | -8.14e-11 | 44.2785
-3.1832e-12 | 44.3075 | -8.3446e-11 | 44.6846
-2.2737e-12 | 44.7116 | -7.4806e-11 | 45.0916
-3.4106e-12 | 45.1166 | -8.6857e-11 | 45.4996
-3.4106e-12 | 45.5216 | -8.049e-11 | 45.9056
-3.8654e-12 | 45.9266 | -7.2077e-11 | 46.3117
-3.4106e-12 | 46.3317 | -8.3901e-11 | 46.7177
-2.5011e-12 | 46.7367 | -8.049e-11 | 47.1237
-3.4106e-12 | 47.1427 | -6.9576e-11 | 47.5277
-3.4106e-12 | 47.5477 | -7.9581e-11 | 47.9347
-3.8654e-12 | 47.9527 | -6.8212e-11 | 48.3418
-3.4106e-12 | 48.3568 | -7.617e-11 | 48.7478
-1.819e-12 | 48.7618 | -6.5711e-11 | 49.1528
-2.0464e-12 | 49.1668 | -7.3442e-11 | 49.5588
-3.1832e-12 | 49.5718 | -8.3446e-11 | 49.9639
-2.7285e-12 | 49.9759 | -8.0036e-11 | 50.3679
-4.0927e-12 | 50.3799 | -6.3437e-11 | 50.7739
-3.8654e-12 | 50.7849 | -8.6402e-11 | 51.1789
-3.4106e-12 | 51.1879 | -7.8899e-11 | 51.5849
-5.9117e-12 | 51.5939 | -8.2309e-11 | 51.99
-2.7285e-12 | 51.999 | -6.7075e-11 | 52.403
-3.638e-12 | 52.404 | -7.4806e-11 | 52.807
-4.5475e-12 | 52.807 | -7.2532e-11 | 53.217
-5.2296e-12 | 53.212 | -8.6402e-11 | 53.6231
-4.7748e-12 | 53.6161 | -7.2077e-11 | 54.0301
-3.1832e-12 | 54.0211 | -7.5943e-11 | 54.4361
-3.638e-12 | 54.4251 | -8.0945e-11 | 54.8481
-5.6843e-12 | 54.8311 | -7.276e-11 | 55.2582
-3.638e-12 | 55.2372 | -8.0036e-11 | 55.6662
-3.8654e-12 | 55.6422 | -7.1395e-11 | 56.0742
-2.2737e-12 | 56.0462 | -7.4351e-11 | 56.4902
-7.5033e-12 | 56.4512 | -8.2991e-11 | 56.8972
-3.4106e-12 | 56.8563 | -7.617e-11 | 57.3043
-3.4106e-12 | 57.2613 | -9.0267e-11 | 57.7123
-4.7748e-12 | 57.6643 | -6.9349e-11 | 58.1183
-4.5475e-12 | 58.0693 | -7.8444e-11 | 58.5254
-3.4106e-12 | 58.4743 | -7.9808e-11 | 58.9324
-1.819e-12 | 58.8784 | -7.708e-11 | 59.3384
-3.638e-12 | 59.2824 | -8.0945e-11 | 59.7444
-3.1832e-12 | 59.6874 | -7.7307e-11 | 60.1484
-4.5475e-12 | 60.0934 | -7.9126e-11 | 60.5545
-3.8654e-12 | 60.4975 | -7.8671e-11 | 60.9605
-3.1832e-12 | 60.9025 | -7.8899e-11 | 61.3675
-4.5475e-12 | 61.3065 | -6.5484e-11 | 61.7725
-3.4106e-12 | 61.7115 | -7.2532e-11 | 62.1766
-2.7285e-12 | 62.1155 | -7.8671e-11 | 62.5806
-3.638e-12 | 62.5216 | -8.5493e-11 | 62.9856
-5.0022e-12 | 62.9266 | -8.5038e-11 | 63.3896
-4.5475e-12 | 63.3316 | -7.7307e-11 | 63.7936
-3.1832e-12 | 63.7366 | -7.4351e-11 | 64.1987
-2.2737e-12 | 64.1407 | -9.0495e-11 | 64.6047
-2.5011e-12 | 64.5447 | -7.3214e-11 | 65.0097
-3.638e-12 | 64.9507 | -7.4351e-11 | 65.4147
-4.5475e-12 | 65.3557 | -7.5033e-11 | 65.8198
-4.5475e-12 | 65.7608 | -8.6857e-11 | 66.2258
-3.8654e-12 | 66.1668 | -8.1627e-11 | 66.6318
-2.7285e-12 | 66.5708 | -8.5947e-11 | 67.0368
-5.6843e-12 | 66.9758 | -8.1172e-11 | 67.4429
-5.2296e-12 | 67.3819 | -9.004e-11 | 67.8469
-4.5475e-12 | 67.7859 | -7.5943e-11 | 68.2529
-4.5475e-12 | 68.1929 | -8.0263e-11 | 68.6579
-1.819e-12 | 68.5969 | -7.0258e-11 | 69.076
-4.0927e-12 | 69.0029 | -7.5943e-11 | 69.485
-4.5475e-12 | 69.408 | -6.8894e-11 | 69.89
-2.5011e-12 | 69.814 | -8.2537e-11 | 70.295
-5.0022e-12 | 70.219 | -7.1168e-11 | 70.704
-4.0927e-12 | 70.623 | -7.2305e-11 | 71.1091
-5.0022e-12 | 71.0281 | -8.0036e-11 | 71.5151
-2.7285e-12 | 71.4321 | -8.2082e-11 | 71.9211
-4.7748e-12 | 71.8381 | -9.004e-11 | 72.3261
-3.1832e-12 | 72.2421 | -6.9576e-11 | 72.7312
-3.4106e-12 | 72.6471 | -8.5265e-11 | 73.1392
-3.638e-12 | 73.0522 | -6.4347e-11 | 73.5442
-4.7748e-12 | 73.4562 | -7.8217e-11 | 73.9502
-4.7748e-12 | 73.8612 | -8.4128e-11 | 74.3552
-3.638e-12 | 74.2652 | -8.6402e-11 | 74.7613
-3.1832e-12 | 74.6703 | -7.1168e-11 | 75.1663
-4.0927e-12 | 75.0743 | -8.4356e-11 | 75.5713
-5.0022e-12 | 75.4783 | -8.1172e-11 | 75.9763
-6.3665e-12 | 75.8833 | -6.5711e-11 | 76.3814
-3.638e-12 | 76.2884 | -8.0036e-11 | 76.7854
-3.8654e-12 | 76.6924 | -7.3442e-11 | 77.1904
-2.7285e-12 | 77.0974 | -8.3674e-11 | 77.5954
-5.0022e-12 | 77.5034 | -7.708e-11 | 78.0005
-2.5011e-12 | 77.9095 | -6.8439e-11 | 78.4045
-2.0464e-12 | 78.3145 | -8.2764e-11 | 78.8105
-5.2296e-12 | 78.7205 | -7.3669e-11 | 79.2155
-4.7748e-12 | 79.1255 | -8.7766e-11 | 79.6205
-2.2737e-12 | 79.5306 | -8.2537e-11 | 80.0266
-2.5011e-12 | 79.9346 | -8.2991e-11 | 80.4316
-7.5033e-12 | 80.3406 | -8.5947e-11 | 80.8366
-5.6843e-12 | 80.7446 | -7.2305e-11 | 81.2426
-3.8654e-12 | 81.1496 | -7.185e-11 | 81.6477
-1.819e-12 | 81.5547 | -7.0941e-11 | 82.0537
-3.8654e-12 | 81.9587 | -8.3674e-11 | 82.4597
-7.7307e-12 | 82.3627 | -8.481e-11 | 82.8657
-4.7748e-12 | 82.7677 | -8.2764e-11 | 83.2718

## 620uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-4.3836e-12 | 1.414e-12 | -1.0966e-10 | 7.9298e-12
GEMs On |  | GEMs On
-4.0927e-12 | 0.31002 | -1.0482e-10 | 0.31202
-9.0949e-13 | 0.99406 | -1.0255e-10 | 0.99706
-2.9559e-12 | 1.3971 | -1.1732e-10 | 1.4041
-6.8212e-12 | 1.8021 | -1.1323e-10 | 1.8081
-4.7748e-12 | 2.2071 | -1.1755e-10 | 2.2131
-2.5011e-12 | 2.6122 | -1.0618e-10 | 2.6181
-4.0927e-12 | 3.0182 | -1.1073e-10 | 3.0232
-5.457e-12 | 3.4232 | -1.1505e-10 | 3.4272
-4.0927e-12 | 3.8282 | -1.0982e-10 | 3.8322
-4.0927e-12 | 4.2322 | -1.0732e-10 | 4.2362
-4.0927e-12 | 4.6363 | -1.0959e-10 | 4.6423
-2.9559e-12 | 5.0403 | -1.0004e-10 | 5.0483
-5.457e-12 | 5.4463 | -1.1596e-10 | 5.4533
-1.5916e-12 | 5.8503 | -1.03e-10 | 5.8573
-4.3201e-12 | 6.2554 | -1.2164e-10 | 6.2634
-4.0927e-12 | 6.6594 | -1.1187e-10 | 6.6684
-5.457e-12 | 7.0634 | -1.0732e-10 | 7.0724
-5.457e-12 | 7.4704 | -1.1846e-10 | 7.4784
-5.9117e-12 | 7.8755 | -1.1096e-10 | 7.8825
-6.1391e-12 | 8.2805 | -1.0209e-10 | 8.2885
-4.3201e-12 | 8.6875 | -1.1914e-10 | 8.6925
-2.5011e-12 | 9.0915 | -1.2369e-10 | 9.0975
-3.1832e-12 | 9.4965 | -1.0004e-10 | 9.5005
-4.5475e-12 | 9.9016 | -1.2528e-10 | 9.9066
-4.0927e-12 | 10.3056 | -9.9817e-11 | 10.3116
-3.4106e-12 | 10.7106 | -1.2119e-10 | 10.7166
-3.638e-12 | 11.1156 | -1.1732e-10 | 11.1226
-4.5475e-12 | 11.5197 | -1.2164e-10 | 11.5287
-5.457e-12 | 11.9247 | -1.1073e-10 | 11.9317
-4.7748e-12 | 12.3287 | -1.0164e-10 | 12.3367
-4.0927e-12 | 12.7337 | -1.1346e-10 | 12.7417
-4.3201e-12 | 13.1387 | -9.436e-11 | 13.1457
-6.3665e-12 | 13.5438 | -1.1505e-10 | 13.5508
-6.8212e-12 | 13.9498 | -1.0687e-10 | 13.9558
-3.638e-12 | 14.3538 | -1.1073e-10 | 14.3608
-4.0927e-12 | 14.7578 | -1.2415e-10 | 14.7658
-4.7748e-12 | 15.1629 | -1.0095e-10 | 15.1699
-3.4106e-12 | 15.5679 | -9.5952e-11 | 15.5749
-3.4106e-12 | 15.9729 | -1.1119e-10 | 15.9799
-3.1832e-12 | 16.3789 | -1.0436e-10 | 16.3849
-5.457e-12 | 16.783 | -1.0209e-10 | 16.789
-3.638e-12 | 17.187 | -9.4133e-11 | 17.194
-4.5475e-12 | 17.593 | -1.1005e-10 | 17.599
-4.3201e-12 | 17.997 | -1.1482e-10 | 18.003
-1.819e-12 | 18.402 | -1.0505e-10 | 18.408
-3.4106e-12 | 18.8071 | -1.005e-10 | 18.8131
-5.9117e-12 | 19.2121 | -1.1778e-10 | 19.2181
-5.6843e-12 | 19.6171 | -1.0186e-10 | 19.6221
-3.638e-12 | 20.0221 | -1.3119e-10 | 20.0262
-3.4106e-12 | 20.4262 | -1.0596e-10 | 20.4322
-6.1391e-12 | 20.8312 | -1.0732e-10 | 20.8362
-6.3665e-12 | 21.2362 | -1.0186e-10 | 21.2402
-3.4106e-12 | 21.6402 | -1.2005e-10 | 21.6452
-3.638e-12 | 22.0453 | -1.0982e-10 | 22.0493
-4.0927e-12 | 22.4493 | -1.0709e-10 | 22.4553
-6.8212e-12 | 22.8543 | -1.1892e-10 | 22.8593
-8.6402e-12 | 23.2603 | -1.1596e-10 | 23.2633
-4.7748e-12 | 23.6654 | -1.0846e-10 | 23.6673
-3.638e-12 | 24.0704 | -1.0346e-10 | 24.0734
-6.1391e-12 | 24.4744 | -1.1369e-10 | 24.4774
-5.6843e-12 | 24.8794 | -1.196e-10 | 24.8824
-3.4106e-12 | 25.2855 | -1.0232e-10 | 25.2864
-5.0022e-12 | 25.6915 | -1.0323e-10 | 25.6915
-4.0927e-12 | 26.0955 | -1.0846e-10 | 26.0965
-4.5475e-12 | 26.5015 | -1.2233e-10 | 26.5005
-4.0927e-12 | 26.9065 | -9.2086e-11 | 26.9055
-4.5475e-12 | 27.3106 | -1.0505e-10 | 27.3106
-6.1391e-12 | 27.7156 | -1.1914e-10 | 27.7156
-2.2737e-12 | 28.1206 | -1.1005e-10 | 28.1206
-2.2737e-12 | 28.5246 | -1.1141e-10 | 28.5256
-3.638e-12 | 28.9297 | -1.0641e-10 | 28.9296
-4.7748e-12 | 29.3357 | -1.0073e-10 | 29.3337
-4.5475e-12 | 29.7397 | -1.1642e-10 | 29.7397
-4.5475e-12 | 30.1447 | -1.3301e-10 | 30.1447
-3.4106e-12 | 30.5487 | -1.196e-10 | 30.5497
-3.4106e-12 | 30.9538 | -1.1482e-10 | 30.9538
-3.638e-12 | 31.3588 | -1.0868e-10 | 31.3588
-4.5475e-12 | 31.7628 | -9.7998e-11 | 31.7628
-4.0927e-12 | 32.1678 | -1.0186e-10 | 32.1678
-6.1391e-12 | 32.5729 | -1.2142e-10 | 32.5719
-3.4106e-12 | 32.9779 | -1.0618e-10 | 32.9769
-6.8212e-12 | 33.3829 | -1.1369e-10 | 33.3829
-4.0927e-12 | 33.7889 | -1.0414e-10 | 33.7879
-3.638e-12 | 34.195 | -1.1664e-10 | 34.193
-4.3201e-12 | 34.599 | -1.1278e-10 | 34.598
-3.1832e-12 | 35.003 | -1.0346e-10 | 35.003
-4.0927e-12 | 35.408 | -9.7089e-11 | 35.408
-3.4106e-12 | 35.813 | -1.0982e-10 | 35.8141
-2.9559e-12 | 36.2191 | -1.0982e-10 | 36.2181
-2.0464e-12 | 36.6231 | -1.0709e-10 | 36.6231
-2.5011e-12 | 37.0281 | -1.03e-10 | 37.0281
-2.9559e-12 | 37.4341 | -1.1005e-10 | 37.4331
-4.0927e-12 | 37.8382 | -1.005e-10 | 37.8372
-5.6843e-12 | 38.2422 | -1.146e-10 | 38.2412
-5.9117e-12 | 38.6462 | -9.1632e-11 | 38.6462
-3.638e-12 | 39.0512 | -1.121e-10 | 39.0522
-2.5011e-12 | 39.4573 | -9.9135e-11 | 39.4573
-4.5475e-12 | 39.8623 | -1.1505e-10 | 39.8623
-5.6843e-12 | 40.2663 | -1.0732e-10 | 40.2673
-4.5475e-12 | 40.6703 | -1.1414e-10 | 40.6723
-6.8212e-13 | 41.0744 | -1.1801e-10 | 41.0763
-3.4106e-12 | 41.4774 | -8.913e-11 | 41.4834
-4.5475e-12 | 41.8814 | -9.868e-11 | 41.8884
-4.5475e-12 | 42.2864 | -1.1664e-10 | 42.2934
-4.0927e-12 | 42.6914 | -1.1119e-10 | 42.6974
-4.3201e-12 | 43.0965 | -1.0505e-10 | 43.1025
-5.0022e-12 | 43.5015 | -1.0414e-10 | 43.5075
-4.5475e-12 | 43.9055 | -1.1778e-10 | 43.9125
-4.0927e-12 | 44.3095 | -1.2164e-10 | 44.3155
-5.0022e-12 | 44.7146 | -1.005e-10 | 44.7206
-2.5011e-12 | 45.1186 | -9.8453e-11 | 45.1256
-4.3201e-12 | 45.5236 | -1.1141e-10 | 45.5316
-4.0927e-12 | 45.9286 | -9.7771e-11 | 45.9366
-3.638e-12 | 46.3336 | -1.1369e-10 | 46.3426
-3.638e-12 | 46.7377 | -1.1846e-10 | 46.7477
-4.3201e-12 | 47.1437 | -1.1482e-10 | 47.1527
-5.9117e-12 | 47.5487 | -1.2255e-10 | 47.5577
-4.0927e-12 | 47.9527 | -9.959e-11 | 47.9627
-3.4106e-12 | 48.3568 | -1.1391e-10 | 48.3678
-4.5475e-12 | 48.7618 | -1.0959e-10 | 48.7728
-4.7748e-12 | 49.1668 | -1.03e-10 | 49.1778
-2.9559e-12 | 49.5718 | -1.1369e-10 | 49.5828
-4.0927e-12 | 49.9769 | -9.2996e-11 | 49.9869
-3.638e-12 | 50.3809 | -1.0959e-10 | 50.3919
-5.457e-12 | 50.7859 | -1.0346e-10 | 50.7969
-3.638e-12 | 51.1899 | -1.1255e-10 | 51.2009
-3.4106e-12 | 51.596 | -9.1404e-11 | 51.6059
-2.5011e-12 | 52.001 | -1.1664e-10 | 52.011
-3.1832e-12 | 52.406 | -1.2142e-10 | 52.416
-4.3201e-12 | 52.812 | -1.0505e-10 | 52.82
-1.5916e-12 | 53.217 | -1.0868e-10 | 53.224
-4.3201e-12 | 53.6221 | -1.0868e-10 | 53.6291
-4.0927e-12 | 54.0261 | -1.2233e-10 | 54.0341
-4.7748e-12 | 54.4311 | -1.0459e-10 | 54.4391
-5.9117e-12 | 54.8351 | -1.0732e-10 | 54.8441
-4.0927e-12 | 55.2402 | -1.1073e-10 | 55.2502
-4.3201e-12 | 55.6452 | -1.1232e-10 | 55.6542
-6.1391e-12 | 56.0502 | -1.0618e-10 | 56.0582
-7.0486e-12 | 56.4562 | -1.1323e-10 | 56.4642
-3.4106e-12 | 56.8612 | -1.0755e-10 | 56.8693
-4.3201e-12 | 57.2663 | -1.0732e-10 | 57.2743
-5.9117e-12 | 57.6703 | -1.1073e-10 | 57.6793
-4.7748e-12 | 58.0753 | -1.055e-10 | 58.0843
-5.0022e-12 | 58.4793 | -1.0937e-10 | 58.4894
-4.3201e-12 | 58.8854 | -1.0937e-10 | 58.8944
-6.3665e-12 | 59.2894 | -1.1573e-10 | 59.2974
-7.276e-12 | 59.6944 | -1.0982e-10 | 59.7014
-3.4106e-12 | 60.0994 | -1.1914e-10 | 60.1064
-2.9559e-12 | 60.5045 | -1.1073e-10 | 60.5105
-5.457e-12 | 60.9105 | -1.0618e-10 | 60.9155
-8.1855e-12 | 61.3135 | -1.1642e-10 | 61.3195
-4.0927e-12 | 61.7175 | -9.7998e-11 | 61.7245
-4.3201e-12 | 62.1215 | -1.2778e-10 | 62.1285
-3.4106e-12 | 62.5266 | -1.0618e-10 | 62.5336
-5.457e-12 | 62.9316 | -1.146e-10 | 62.9386
-7.0486e-12 | 63.3376 | -1.0164e-10 | 63.3436
-3.4106e-12 | 63.7416 | -1.1323e-10 | 63.7477
-3.638e-12 | 64.1477 | -1.0095e-10 | 64.1527
-5.6843e-12 | 64.5527 | -1.1482e-10 | 64.5577
-1.819e-12 | 64.9577 | -1.2687e-10 | 64.9627
-2.2737e-12 | 65.3627 | -1.0209e-10 | 65.3677
-3.638e-12 | 65.7668 | -1.0118e-10 | 65.7718
-5.0022e-12 | 66.1708 | -1.1619e-10 | 66.1778
-6.8212e-12 | 66.5748 | -1.0982e-10 | 66.5818
-3.1832e-12 | 66.9798 | -1.055e-10 | 66.9858
-3.4106e-12 | 67.3838 | -1.1141e-10 | 67.3919
-7.5033e-12 | 67.7879 | -9.868e-11 | 67.7949
-5.457e-12 | 68.1919 | -1.1778e-10 | 68.1999
-3.1832e-12 | 68.5979 | -1.1141e-10 | 68.6059
-1.819e-12 | 69.0019 | -1.0709e-10 | 69.0109
-5.0022e-12 | 69.408 | -1.1255e-10 | 69.416
-5.6843e-12 | 69.813 | -1.0505e-10 | 69.821
-4.5475e-12 | 70.217 | -1.03e-10 | 70.226
-2.9559e-12 | 70.622 | -1.1119e-10 | 70.631
-7.5033e-12 | 71.0281 | -1.0732e-10 | 71.0341
-6.3665e-12 | 71.4341 | -1.1528e-10 | 71.4391
-3.4106e-12 | 71.8381 | -1.1573e-10 | 71.8451
-3.638e-12 | 72.2421 | -1.0436e-10 | 72.2491
-6.3665e-12 | 72.6462 | -1.0232e-10 | 72.6542
-7.9581e-12 | 73.0522 | -1.0846e-10 | 73.0592
-4.0927e-12 | 73.4572 | -1.0891e-10 | 73.4622
-6.8212e-13 | 73.8612 | -1.1869e-10 | 73.8672
-5.0022e-12 | 74.2643 | -1.2142e-10 | 74.2712
-5.6843e-12 | 74.6683 | -1.0004e-10 | 74.6753
-7.0486e-12 | 75.0733 | -1.1073e-10 | 75.0803
-4.5475e-12 | 75.4783 | -1.171e-10 | 75.4853
-3.4106e-12 | 75.8823 | -1.0846e-10 | 75.8893
-6.1391e-12 | 76.2874 | -1.1187e-10 | 76.2944
-4.7748e-12 | 76.6934 | -1.2028e-10 | 76.7004
-2.5011e-12 | 77.0974 | -1.08e-10 | 77.1034
-2.2737e-12 | 77.5014 | -1.0164e-10 | 77.5094
-5.9117e-12 | 77.9065 | -1.055e-10 | 77.9145
-5.457e-12 | 78.3115 | -1.1323e-10 | 78.3185
-5.0022e-12 | 78.7165 | -1.1073e-10 | 78.7245
-5.0022e-12 | 79.1205 | -9.527e-11 | 79.1305
-5.6843e-12 | 79.5255 | -1.0482e-10 | 79.5375
-6.8212e-12 | 79.9306 | -1.0255e-10 | 79.9416
-2.9559e-12 | 80.3366 | -1.1778e-10 | 80.3456
-3.4106e-12 | 80.7406 | -1.0755e-10 | 80.7506
-3.638e-12 | 81.1456 | -1.1005e-10 | 81.1556
-5.9117e-12 | 81.5507 | -1.2142e-10 | 81.5607
-4.3201e-12 | 81.9547 | -1.0709e-10 | 81.9657
-2.5011e-12 | 82.3587 | -1.1642e-10 | 82.3697
-4.3201e-12 | 82.7627

## 630uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-5.4125e-12 | 1.8298e-12 | -1.5228e-10 | 1.1294e-11
GEMs On |  | GEMs On
-5.9117e-12 | 0.31002 | -1.5302e-10 | 0.31502
-7.7307e-12 | 0.99506 | -1.5712e-10 | 1.0001
-5.2296e-12 | 1.4001 | -1.6507e-10 | 1.4061
-2.9559e-12 | 1.8061 | -1.4552e-10 | 1.8101
-5.2296e-12 | 2.2101 | -1.6848e-10 | 2.2161
-7.9581e-12 | 2.6151 | -1.653e-10 | 2.6202
-6.3665e-12 | 3.0192 | -1.5689e-10 | 3.0242
-4.5475e-12 | 3.4242 | -1.5211e-10 | 3.4302
-3.638e-12 | 3.8282 | -1.5029e-10 | 3.8342
-5.0022e-12 | 4.2332 | -1.7508e-10 | 4.2392
-6.8212e-12 | 4.6373 | -1.703e-10 | 4.6443
-4.5475e-12 | 5.0423 | -1.4461e-10 | 5.0493
-5.2296e-12 | 5.4473 | -1.5939e-10 | 5.4533
-5.0022e-12 | 5.8523 | -1.5348e-10 | 5.8583
-5.0022e-12 | 6.2574 | -1.6144e-10 | 6.2624
-6.5938e-12 | 6.6614 | -1.528e-10 | 6.6674
-6.8212e-12 | 7.0664 | -1.4529e-10 | 7.0724
-1.819e-12 | 7.4714 | -1.5984e-10 | 7.4764
-7.5033e-12 | 7.8765 | -1.4461e-10 | 7.8825
-6.5938e-12 | 8.2825 | -1.7371e-10 | 8.2885
-3.8654e-12 | 8.6875 | -1.6053e-10 | 8.6985
-5.2296e-12 | 9.0925 | -1.6576e-10 | 9.1035
-8.1855e-12 | 9.4955 | -1.537e-10 | 9.5095
-8.1855e-12 | 9.9006 | -1.6507e-10 | 9.9156
-4.0927e-12 | 10.3056 | -1.4734e-10 | 10.3196
-3.638e-12 | 10.7106 | -1.5325e-10 | 10.7246
-1.0232e-11 | 11.1146 | -1.5871e-10 | 11.1296
-7.9581e-12 | 11.5197 | -1.3688e-10 | 11.5337
-5.2296e-12 | 11.9257 | -1.653e-10 | 11.9387
-3.8654e-12 | 12.3297 | -1.5166e-10 | 12.3467
-7.7307e-12 | 12.7347 | -1.3802e-10 | 12.7527
-8.8676e-12 | 13.1397 | -1.678e-10 | 13.1578
-5.9117e-12 | 13.5458 | -1.5302e-10 | 13.5618
2.7285e-12 | 13.9498 | -1.5075e-10 | 13.9688
-7.9581e-12 | 14.3558 | -1.4802e-10 | 14.3728
-7.9581e-12 | 14.7598 | -1.5234e-10 | 14.7779
-6.8212e-12 | 15.1649 | -1.3756e-10 | 15.1829
-2.0464e-12 | 15.5689 | -1.4711e-10 | 15.6339
-4.5475e-12 | 15.9739 | -1.5984e-10 | 16.0399
-7.7307e-12 | 16.3799 | -1.5598e-10 | 16.4449
-9.0949e-12 | 16.785 | -1.3688e-10 | 16.852
-2.7285e-12 | 17.19 | -1.7349e-10 | 17.265
-4.0927e-12 | 17.595 | -1.5507e-10 | 17.669
-6.3665e-12 | 18 | -1.3756e-10 | 18.073
-5.9117e-12 | 18.405 | -1.5052e-10 | 18.4781
-3.8654e-12 | 18.8101 | -1.4666e-10 | 18.8851
-5.0022e-12 | 19.2141 | -1.5962e-10 | 19.2911
-7.5033e-12 | 19.6191 | -1.487e-10 | 19.6961
-5.9117e-12 | 20.0241 | -1.4779e-10 | 20.1001
-4.5475e-12 | 20.4292 | -1.3779e-10 | 20.5052
-2.9559e-12 | 20.8332 | -1.4256e-10 | 20.9102
-4.0927e-12 | 21.2382 | -1.6621e-10 | 21.3172
-6.8212e-12 | 21.6422 | -1.6144e-10 | 21.7222
-6.8212e-12 | 22.0473 | -1.2551e-10 | 22.1263
6.8212e-13 | 22.4513 | -1.6848e-10 | 22.5313
-4.5475e-12 | 22.8563 | -1.3893e-10 | 22.9363
-7.5033e-12 | 23.2603 | -1.628e-10 | 23.3423
-6.5938e-12 | 23.6643 | -1.3165e-10 | 23.7524
-5.457e-12 | 24.0694 | -1.5552e-10 | 24.1594
4.5475e-13 | 24.4734 | -1.5848e-10 | 24.5654
-5.6843e-12 | 24.8784 | -1.5689e-10 | 24.9704
-6.3665e-12 | 25.2834 | -1.5552e-10 | 25.3744
-5.6843e-12 | 25.6885 | -1.3256e-10 | 25.7815
-6.5938e-12 | 26.0935 | -1.4415e-10 | 26.1985
-2.9559e-12 | 26.4985 | -1.7008e-10 | 26.6035
-6.8212e-12 | 26.9025 | -1.5552e-10 | 27.0095
-6.5938e-12 | 27.3076 | -1.4802e-10 | 27.4136
-5.0022e-12 | 27.7126 | -1.6826e-10 | 27.8196
-8.4128e-12 | 28.1166 | -1.6075e-10 | 28.2246
-5.457e-12 | 28.5206 | -1.4188e-10 | 28.6656
-1.0459e-11 | 28.9246 | -1.3802e-10 | 29.0717
-8.1855e-12 | 29.3307 | -1.5802e-10 | 29.4777
-5.457e-12 | 29.7357 | -1.4552e-10 | 29.8827
-6.5938e-12 | 30.1397 | -1.5939e-10 | 30.2867
-6.3665e-12 | 30.5448 | -1.2415e-10 | 30.6908
-5.0022e-12 | 30.9508 | -1.653e-10 | 31.0938
-5.6843e-12 | 31.3558 | -1.5416e-10 | 31.4988
-3.638e-12 | 31.7608 | -1.4961e-10 | 31.9028
-6.8212e-12 | 32.1668 | -1.5757e-10 | 32.3079
-4.0927e-12 | 32.5709 | -1.653e-10 | 32.7119
-5.9117e-12 | 32.9749 | -1.5848e-10 | 33.1159
-3.638e-12 | 33.3809 | -1.3665e-10 | 33.5209
-6.3665e-12 | 33.7849 | -1.528e-10 | 33.9249
-5.6843e-12 | 34.19 | -1.4074e-10 | 34.33
-5.0022e-12 | 34.595 | -1.5734e-10 | 34.736
-6.8212e-12 | 34.999 | -1.4688e-10 | 35.14
-5.2296e-12 | 35.403 | -1.5348e-10 | 35.545
-6.8212e-12 | 35.8081 | -1.6917e-10 | 35.9501
-4.5475e-12 | 36.2131 | -1.4734e-10 | 36.3561
-6.3665e-12 | 36.6181 | -1.5712e-10 | 36.7621
-5.0022e-12 | 37.0231 | -1.5757e-10 | 37.1671
-8.1855e-12 | 37.4291 | -1.5848e-10 | 37.5722
-7.7307e-12 | 37.8342 | -1.6075e-10 | 37.9782
-5.2296e-12 | 38.2402 | -1.3506e-10 | 38.3832
-5.6843e-12 | 38.6452 | -1.6053e-10 | 38.7882
-4.5475e-12 | 39.0492 | -1.2847e-10 | 39.1922
-5.457e-12 | 39.4543 | -1.4688e-10 | 39.5963
-5.0022e-12 | 39.8583 | -1.5075e-10 | 40.0003
-3.638e-12 | 40.2633 | -1.4847e-10 | 40.4043
-5.6843e-12 | 40.6683 | -1.5484e-10 | 40.8093
-5.9117e-12 | 41.0733 | -1.487e-10 | 41.2154
-2.7285e-12 | 41.4784 | -1.5712e-10 | 41.6194
-3.638e-12 | 41.8824 | -1.5893e-10 | 42.0244
-4.5475e-12 | 42.2874 | -1.3688e-10 | 42.4334
-6.3665e-12 | 42.6934 | -1.4916e-10 | 42.8385
-3.638e-12 | 43.0985 | -1.5712e-10 | 43.2435
-5.2296e-12 | 43.5035 | -1.6257e-10 | 43.6485
-4.3201e-12 | 43.9085 | -1.5325e-10 | 44.0525
-7.0486e-12 | 44.3135 | -1.5166e-10 | 44.4575
-2.5011e-12 | 44.7196 | -1.5075e-10 | 44.8626
-4.3201e-12 | 45.1236 | -1.2869e-10 | 45.2676
-6.8212e-12 | 45.5296 | -1.3551e-10 | 45.6726
-6.5938e-12 | 45.9366 | -1.4211e-10 | 46.0776
-3.638e-12 | 46.3406 | -1.5757e-10 | 46.4827
-4.0927e-12 | 46.7457 | -1.5984e-10 | 46.8877
-5.6843e-12 | 47.1497 | -1.4097e-10 | 47.2917
-4.5475e-12 | 47.5547 | -1.4074e-10 | 47.6957
-4.3201e-12 | 47.9587 | -1.6689e-10 | 48.1007
-4.0927e-12 | 48.3648 | -1.4143e-10 | 48.5058
-5.457e-12 | 48.7698 | -1.3893e-10 | 48.9118
-5.6843e-12 | 49.1748 | -1.7144e-10 | 49.3158
-2.9559e-12 | 49.5798 | -1.6075e-10 | 49.7198
-5.457e-12 | 49.9859 | -1.3529e-10 | 50.1249
-9.5497e-12 | 50.3909 | -1.4211e-10 | 50.5289
-5.457e-12 | 50.7949 | -1.4552e-10 | 50.9329
-5.2296e-12 | 51.1989 | -1.5825e-10 | 51.3389
-5.6843e-12 | 51.6039 | -1.6075e-10 | 51.745
-6.5938e-12 | 52.008 | -1.4802e-10 | 52.149
-5.6843e-12 | 52.414 | -1.387e-10 | 52.555
-6.3665e-12 | 52.818 | -1.653e-10 | 52.96
-5.2296e-12 | 53.224 | -1.728e-10 | 53.364
-7.7307e-12 | 53.6291 | -1.5075e-10 | 53.7691
-6.8212e-12 | 54.0351 | -1.5484e-10 | 54.1741
-2.5011e-12 | 54.4401 | -1.7985e-10 | 54.5781
-4.5475e-12 | 54.8441 | -1.5348e-10 | 54.9821
-5.0022e-12 | 55.2482 | -1.4779e-10 | 55.3892
-9.7771e-12 | 55.6532 | -1.4688e-10 | 55.7932
-4.5475e-12 | 56.0572 | -1.6075e-10 | 56.1982
-5.9117e-12 | 56.4622 | -1.3165e-10 | 56.6032
-5.0022e-12 | 56.8663 | -1.6053e-10 | 57.0083
-3.638e-12 | 57.2713 | -1.5348e-10 | 57.4123
-7.0486e-12 | 57.6753 | -1.5234e-10 | 57.8183
-4.3201e-12 | 58.0783 | -1.528e-10 | 58.2233
-5.9117e-12 | 58.4823 | -1.5484e-10 | 58.6294
-6.8212e-12 | 58.8874 | -1.4461e-10 | 59.0334
-1.819e-12 | 59.2914 | -1.6439e-10 | 59.4394
-5.0022e-12 | 59.6964 | -1.4438e-10 | 59.8444
-4.5475e-12 | 60.1004 | -1.6371e-10 | 60.2495
-4.3201e-12 | 60.5055 | -1.5461e-10 | 60.6545
-4.3201e-12 | 60.9095 | -1.6962e-10 | 61.0595
-7.0486e-12 | 61.3135 | -1.3415e-10 | 61.4645
-5.9117e-12 | 61.7165 | -1.5848e-10 | 61.8695
-5.9117e-12 | 62.1225 | -1.8053e-10 | 62.2736
-3.8654e-12 | 62.5276 | -1.4961e-10 | 62.6786
-5.2296e-12 | 62.9316 | -1.3688e-10 | 63.0846
-4.3201e-12 | 63.3356 | -1.3074e-10 | 63.4886
-5.2296e-12 | 63.7407 | -1.5552e-10 | 63.8937
-5.2296e-12 | 64.1447 | -1.5666e-10 | 64.2977
-4.5475e-12 | 64.5497 | -1.1823e-10 | 64.7017
-9.7771e-12 | 64.9547 | -1.5689e-10 | 65.1067
-8.4128e-12 | 65.3597 | -1.5962e-10 | 65.5118
-6.3665e-12 | 65.7648 | -1.6917e-10 | 65.9158
-2.5011e-12 | 66.1688 | -1.4097e-10 | 66.3198
-2.9559e-12 | 66.5738 | -1.5621e-10 | 66.7248
-5.6843e-12 | 66.9788 | -1.4825e-10 | 67.1308
-7.7307e-12 | 67.3829 | -1.4211e-10 | 67.5359
-5.2296e-12 | 67.7879 | -1.4552e-10 | 67.9409
-4.0927e-12 | 68.1929 | -1.4847e-10 | 68.3449
-7.0486e-12 | 68.5979 | -1.4006e-10 | 68.7489
-7.7307e-12 | 69.0029 | -1.5348e-10 | 69.152
-5.6843e-12 | 69.407 | -1.5416e-10 | 69.556
-5.2296e-12 | 69.812 | -1.4802e-10 | 69.961
-4.3201e-12 | 70.217 | -1.5461e-10 | 70.366
-7.5033e-12 | 70.621 | -1.5052e-10 | 70.77
-4.5475e-12 | 71.0261 | -1.5348e-10 | 71.1751
-4.3201e-12 | 71.4311 | -1.3279e-10 | 71.5801
-6.3665e-12 | 71.8351 | -1.5893e-10 | 71.9841
-7.0486e-12 | 72.2411 | -1.3438e-10 | 72.3891
-6.8212e-12 | 72.6452 | -1.4461e-10 | 72.7942
-3.1832e-12 | 73.0502 | -1.6644e-10 | 73.1992
-3.638e-12 | 73.4562 | -1.4438e-10 | 73.6042
-5.0022e-12 | 73.8592 | -1.4302e-10 | 74.0082
-1.0232e-11 | 74.2643 | -1.4825e-10 | 74.4133
-6.5938e-12 | 74.6703 | -1.7621e-10 | 74.8183
-2.7285e-12 | 75.0743 | -1.4734e-10 | 75.2223
-5.457e-12 | 75.4783 | -1.6507e-10 | 75.6393
-5.9117e-12 | 75.8833 | -1.4484e-10 | 76.0443
-4.3201e-12 | 76.2874 | -1.4688e-10 | 76.4494
-4.3201e-12 | 76.6934 | -1.4984e-10 | 76.8544
-4.5475e-12 | 77.0994 | -1.4643e-10 | 77.2604
-7.5033e-12 | 77.5034 | -1.6234e-10 | 77.6644
-6.3665e-12 | 77.9075 | -1.4711e-10 | 78.0675
-4.5475e-12 | 78.3115 | -1.4597e-10 | 78.4725
-4.3201e-12 | 78.7165 | -1.4438e-10 | 78.8775
-4.3201e-12 | 79.1215 | -1.6576e-10 | 79.2825
-6.3665e-12 | 79.5266 | -1.4393e-10 | 79.6866
-4.3201e-12 | 79.9316 | -1.4529e-10 | 80.0916
-2.5011e-12 | 80.3356 | -1.5052e-10 | 80.4976
-3.8654e-12 | 80.7406 | -1.4415e-10 | 80.9026
-6.5938e-12 | 81.1466 | -1.769e-10 | 81.3067
-4.5475e-12 | 81.5517 | -1.4211e-10 | 81.7117
-3.1832e-12 | 81.9557 | -1.5507e-10 | 82.1167
-5.0022e-12 | 82.3597 | -1.4143e-10 | 82.5207
-5.9117e-12 | 82.7647 | -1.7212e-10 | 82.9257
-7.0486e-12 | 83.1698 | -1.6234e-10 | 83.3308
-3.638e-12 | 83.5728 | -1.5734e-10 | 83.7358
-2.5011e-12 | 83.9778
-3.8654e-12 | 84.3818
-5.9117e-12 | 84.7869
-2.7285e-12 | 85.1919
-3.638e-12 | 85.5969
-5.2296e-12 | 86.0019
-5.9117e-12 | 86.4079
-5.6843e-12 | 86.812
-3.8654e-12 | 87.217
-5.9117e-12 | 87.622
-6.8212e-12 | 88.028
-5.457e-12 | 88.4321
-4.5475e-12 | 88.8381
-4.0927e-12 | 89.2421
-5.2296e-12 | 89.6481
-6.3665e-12 | 90.0521
-5.457e-12 | 90.4572
-6.3665e-12 | 90.8612
-5.6843e-12 | 91.2662

## 640uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-6.1328e-12 | 1.481e-12 | -2.1321e-10 | 1.638e-11
GEMs On |  | GEMs On
-4.7748e-12 | 0.31302 | -2.1714e-10 | 0.31102
-7.276e-12 | 0.99706 | -2.2419e-10 | 0.99406
-7.5033e-12 | 1.4031 | -1.9872e-10 | 1.4001
-4.3201e-12 | 1.8071 | -2.2442e-10 | 1.8051
-3.638e-12 | 2.2121 | -2.3397e-10 | 2.2081
-7.0486e-12 | 2.6161 | -1.7553e-10 | 2.6121
-1.1141e-11 | 3.0202 | -2.1805e-10 | 3.0172
-4.5475e-12 | 3.4252 | -2.1078e-10 | 3.4212
-2.9559e-12 | 3.8322 | -2.135e-10 | 3.8252
-7.0486e-12 | 4.2382 | -2.2874e-10 | 4.2312
-8.4128e-12 | 4.6433 | -2.4738e-10 | 4.6363
-4.0927e-12 | 5.0493 | -1.96e-10 | 5.0413
-5.2296e-12 | 5.4533 | -1.9645e-10 | 5.4453
-7.5033e-12 | 5.8583 | -2.2214e-10 | 5.8493
-6.8212e-12 | 6.2624 | -1.8144e-10 | 6.2544
-4.3201e-12 | 6.6674 | -1.96e-10 | 6.6584
-6.5938e-12 | 7.0724 | -2.3101e-10 | 7.0634
-7.0486e-12 | 7.4774 | -2.3988e-10 | 7.4674
-8.4128e-12 | 7.8825 | -2.1669e-10 | 7.8724
-4.3201e-12 | 8.2875 | -2.151e-10 | 8.2775
-6.1391e-12 | 8.6915 | -2.1691e-10 | 8.6825
-1.0687e-11 | 9.0965 | -1.9554e-10 | 9.0865
-5.9117e-12 | 9.5015 | -2.2533e-10 | 9.4895
-5.9117e-12 | 9.9066 | -2.1896e-10 | 9.8946
-4.3201e-12 | 10.3106 | -2.2101e-10 | 10.2986
-5.457e-12 | 10.7136 | -1.7189e-10 | 10.7046
-9.0949e-12 | 11.1186 | -2.3056e-10 | 11.1106
-5.457e-12 | 11.5237 | -2.0896e-10 | 11.5157
-4.3201e-12 | 11.9277 | -2.1851e-10 | 11.9207
-7.5033e-12 | 12.3337 | -2.2328e-10 | 12.3257
-8.1855e-12 | 12.7387 | -1.9622e-10 | 12.7317
-3.638e-12 | 13.1427 | -2.1305e-10 | 13.1367
-5.9117e-12 | 13.5478 | -2.0191e-10 | 13.5418
-4.3201e-12 | 13.9518 | -1.9668e-10 | 13.9468
-7.276e-12 | 14.3558 | -2.2624e-10 | 14.3508
-6.5938e-12 | 14.7608 | -1.9941e-10 | 14.7568
-5.457e-12 | 15.1659 | -2.3215e-10 | 15.1609
-6.8212e-12 | 15.5719 | -2.2874e-10 | 15.5659
-9.5497e-12 | 15.9759 | -2.0827e-10 | 15.9699
-3.4106e-12 | 16.3829 | -2.1214e-10 | 16.3749
-5.9117e-12 | 16.788 | -2.3715e-10 | 16.779
-5.6843e-12 | 17.194 | -2.2442e-10 | 17.185
-5.457e-12 | 17.599 | -1.9554e-10 | 17.59
-4.5475e-12 | 18.004 | -2.1168e-10 | 17.994
-5.6843e-12 | 18.408 | -2.0282e-10 | 18.3981
-4.0927e-12 | 18.8141 | -2.4238e-10 | 18.8031
-8.4128e-12 | 19.2181 | -2.0191e-10 | 19.2071
-3.1832e-12 | 19.6221 | -1.9099e-10 | 19.6121
-4.0927e-12 | 20.0282 | -1.985e-10 | 20.0161
-5.6843e-12 | 20.4342 | -2.3533e-10 | 20.4232
-5.2296e-12 | 20.8382 | -2.1805e-10 | 20.8262
-4.3201e-12 | 21.2432 | -2.1942e-10 | 21.2312
-5.2296e-12 | 21.6492 | -2.417e-10 | 21.6362
-6.8212e-12 | 22.0533 | -2.3738e-10 | 22.0413
-5.2296e-12 | 22.4593 | -2.0759e-10 | 22.4463
-4.0927e-12 | 22.8653 | -2.0486e-10 | 22.8513
-6.5938e-12 | 23.2703 | -2.2442e-10 | 23.2553
-7.9581e-12 | 23.6754 | -1.9531e-10 | 23.6604
-7.0486e-12 | 24.0794 | -1.9418e-10 | 24.0654
-3.638e-12 | 24.4844 | -2.1805e-10 | 24.4694
-3.1832e-12 | 24.8884 | -2.1191e-10 | 24.8744
-6.5938e-12 | 25.2924 | -1.9395e-10 | 25.2784
-7.276e-12 | 25.6965 | -2.0623e-10 | 25.6835
-5.6843e-12 | 26.1025 | -2.5057e-10 | 26.0895
-4.0927e-12 | 26.5065 | -2.0918e-10 | 26.4935
-6.1391e-12 | 26.9125 | -1.9259e-10 | 26.8985
-7.276e-12 | 27.3176 | -2.4102e-10 | 27.3036
-5.6843e-12 | 27.7216 | -2.2033e-10 | 27.7076
-4.0927e-12 | 28.1256 | -1.935e-10 | 28.1126
-4.7748e-12 | 28.5296 | -2.1964e-10 | 28.5176
-8.6402e-12 | 28.9347 | -1.9941e-10 | 28.9227
-3.638e-12 | 29.3397 | -2.1782e-10 | 29.3267
-6.1391e-12 | 29.7447 | -2.0168e-10 | 29.7307
-8.4128e-12 | 30.1497 | -2.2919e-10 | 30.1357
-7.276e-12 | 30.5538 | -2.5648e-10 | 30.5407
-7.0486e-12 | 30.9578 | -1.9759e-10 | 30.9468
-6.1391e-12 | 31.3618 | -2.0805e-10 | 31.3528
-5.6843e-12 | 31.7648 | -2.1328e-10 | 31.7568
-6.1391e-12 | 32.1708 | -2.1191e-10 | 32.1618
-6.8212e-12 | 32.5739 | -2.1259e-10 | 32.5669
-7.276e-12 | 32.9789 | -1.9918e-10 | 32.9719
-5.9117e-12 | 33.3829 | -2.0441e-10 | 33.3759
-5.2296e-12 | 33.7869 | -2.2305e-10 | 33.7809
-8.4128e-12 | 34.192 | -2.3306e-10 | 34.185
-7.5033e-12 | 34.597 | -2.1464e-10 | 34.591
-6.8212e-12 | 35.002 | -2.5443e-10 | 34.996
-2.7285e-12 | 35.408 | -2.0509e-10 | 35.401
-6.8212e-12 | 35.812 | -2.2942e-10 | 35.8051
-9.3223e-12 | 36.2161 | -2.1691e-10 | 36.2091
-5.457e-12 | 36.6211 | -1.9941e-10 | 36.6121
-7.0486e-12 | 37.0261 | -2.2806e-10 | 37.0171
-8.4128e-12 | 37.4311 | -2.1305e-10 | 37.4211
-7.276e-12 | 37.8362 | -2.1419e-10 | 37.8252
-5.457e-12 | 38.2422 | -2.0509e-10 | 38.2302
-7.0486e-12 | 38.6462 | -2.0782e-10 | 38.6342
-7.0486e-12 | 39.0502 | -1.9645e-10 | 39.0382
-5.457e-12 | 39.4563 | -1.9668e-10 | 39.4433
-4.7748e-12 | 39.8603 | -2.1782e-10 | 39.8483
-4.5475e-12 | 40.2653 | -2.2328e-10 | 40.2533
-5.9117e-12 | 40.6713 | -2.1532e-10 | 40.6583
-6.5938e-12 | 41.0763 | -1.9281e-10 | 41.0633
-6.8212e-12 | 41.4814 | -2.1805e-10 | 41.4694
-5.457e-12 | 41.8854 | -2.0918e-10 | 41.8744
-7.276e-12 | 42.2904 | -2.4852e-10 | 42.2794
-6.1391e-12 | 42.6944 | -1.9872e-10 | 42.6834
-5.9117e-12 | 43.0995 | -1.9418e-10 | 43.0885
-7.276e-12 | 43.5045 | -2.3579e-10 | 43.4925
-6.5938e-12 | 43.9085 | -2.1669e-10 | 43.8965
-6.1391e-12 | 44.3145 | -1.985e-10 | 44.3035
-5.2296e-12 | 44.7196 | -2.3988e-10 | 44.7086
-7.276e-12 | 45.1236 | -2.2237e-10 | 45.1136
-6.5938e-12 | 45.5276 | -2.4102e-10 | 45.5186
-7.0486e-12 | 45.9316 | -2.0668e-10 | 45.9236
-7.276e-12 | 46.3376 | -2.1646e-10 | 46.3276
-5.9117e-12 | 46.7417 | -2.0691e-10 | 46.7327
-6.5938e-12 | 47.1467 | -2.1078e-10 | 47.1377
-4.0927e-12 | 47.5517 | -1.8576e-10 | 47.5417
-6.8212e-12 | 47.9577 | -2.0009e-10 | 47.9467
-5.9117e-12 | 48.3628 | -2.3965e-10 | 48.3528
-7.0486e-12 | 48.7678 | -2.2624e-10 | 48.7568
-5.9117e-12 | 49.1728 | -1.9554e-10 | 49.1618
-8.4128e-12 | 49.5768 | -2.1896e-10 | 49.5658
-6.8212e-12 | 49.9819 | -1.9486e-10 | 49.9709
-5.2296e-12 | 50.3869 | -1.9486e-10 | 50.3759
-5.9117e-12 | 50.7919 | -2.1305e-10 | 50.7809
-4.5475e-12 | 51.1969 | -2.4033e-10 | 51.1849
-6.1391e-12 | 51.602 | -2.1805e-10 | 51.59
-6.1391e-12 | 52.007 | -1.9509e-10 | 51.994
-6.5938e-12 | 52.412 | -2.2283e-10 | 52.399
-6.1391e-12 | 52.816 | -1.6462e-10 | 52.802
-7.9581e-12 | 53.221 | -2.1441e-10 | 53.207
-6.8212e-12 | 53.6261 | -2.2533e-10 | 53.6141
-7.0486e-12 | 54.0321 | -1.8986e-10 | 54.0191
-2.9559e-12 | 54.4371 | -2.1987e-10 | 54.4241
-7.276e-12 | 54.8411 | -2.3488e-10 | 54.8271
-5.457e-12 | 55.2452 | -2.1714e-10 | 55.2312
-4.0927e-12 | 55.6492 | -2.1078e-10 | 55.6362
-5.2296e-12 | 56.0552 | -2.0918e-10 | 56.0392
-8.6402e-12 | 56.4592 | -1.8372e-10 | 56.4452
-7.5033e-12 | 56.8642 | -2.0191e-10 | 56.8503
-5.457e-12 | 57.2703 | -1.985e-10 | 57.2543
-6.1391e-12 | 57.6763 | -1.9099e-10 | 57.6593
-5.2296e-12 | 58.0803 | -1.8645e-10 | 58.0643
-7.276e-12 | 58.4853 | -1.9872e-10 | 58.4693
-5.457e-12 | 58.8894 | -2.2169e-10 | 58.8744
-6.1391e-12 | 59.2954 | -2.1646e-10 | 59.2784
-5.457e-12 | 59.6984 | -2.4397e-10 | 59.6824
-5.9117e-12 | 60.1024 | -2.0577e-10 | 60.0874
-5.457e-12 | 60.5075 | -2.2533e-10 | 60.4915
-3.638e-12 | 60.9125 | -2.4284e-10 | 60.8975
-3.638e-12 | 61.3175 | -2.3397e-10 | 61.3025
-8.1855e-12 | 61.7215 | -1.9486e-10 | 61.7075
-5.6843e-12 | 62.1255 | -2.0873e-10 | 62.1125
-7.0486e-12 | 62.5316 | -2.1964e-10 | 62.5156
-8.1855e-12 | 62.9356 | -1.9622e-10 | 62.9216
-8.1855e-12 | 63.3406 | -2.0782e-10 | 63.3246
-7.5033e-12 | 63.7467 | -2.2624e-10 | 63.7286
-4.7748e-12 | 64.1517 | -1.935e-10 | 64.1327
-3.638e-12 | 64.5567 | -2.2123e-10 | 64.5377
-6.1391e-12 | 64.9607 | -2.1191e-10 | 64.9427
-8.1855e-12 | 65.3657 | -2.2123e-10 | 65.3467
-5.457e-12 | 65.7708 | -2.2442e-10 | 65.7508
-2.9559e-12 | 66.1758 | -2.0441e-10 | 66.1568
-8.1855e-12 | 66.5808 | -2.1078e-10 | 66.5608
-7.5033e-12 | 66.9868 | -2.1396e-10 | 66.9658
-4.7748e-12 | 67.3909 | -2.2192e-10 | 67.3709
-5.457e-12 | 67.7939 | -2.2828e-10 | 67.7759
-6.8212e-12 | 68.1999 | -2.1419e-10 | 68.1819
-8.1855e-12 | 68.6049 | -2.2237e-10 | 68.5869
-4.7748e-12 | 69.0109 | -2.2055e-10 | 68.992
-3.638e-12 | 69.415 | -1.96e-10 | 69.398
-5.2296e-12 | 69.819 | -2.1009e-10 | 69.803
-6.8212e-12 | 70.224 | -2.2033e-10 | 70.208
-6.5938e-12 | 70.629 | -1.8304e-10 | 70.614
-5.2296e-12 | 71.0331 | -2.2624e-10 | 71.0191
-5.9117e-12 | 71.4371 | -2.3465e-10 | 71.4241
-6.8212e-12 | 71.8431 | -2.0327e-10 | 71.8281
-5.9117e-12 | 72.2471 | -2.3442e-10 | 72.2331
-5.2296e-12 | 72.6522 | -2.0532e-10 | 72.6372
-5.9117e-12 | 73.0562 | -2.2737e-10 | 73.0422
-6.5938e-12 | 73.4612 | -1.8508e-10 | 73.4462
-7.0486e-12 | 73.8662 | -2.1578e-10 | 73.8522
-5.2296e-12 | 74.2703 | -1.9713e-10 | 74.2572
-5.2296e-12 | 74.6743 | -1.8531e-10 | 74.6603
-9.7771e-12 | 75.0793 | -2.3215e-10 | 75.0643
-6.5938e-12 | 75.4833 | -1.9622e-10 | 75.4693
-7.0486e-12 | 75.8873 | -2.3374e-10 | 75.8743
-6.5938e-12 | 76.2924 | -1.9645e-10 | 76.2784
-5.457e-12 | 76.6964 | -2.1919e-10 | 76.6834
-5.9117e-12 | 77.1014 | -2.3942e-10 | 77.0894
-7.5033e-12 | 77.5064 | -2.0259e-10 | 77.4944
-6.5938e-12 | 77.9125 | -2.0896e-10 | 77.8995
-5.9117e-12 | 78.3155 | -2.1419e-10 | 78.3055
-6.8212e-12 | 78.7205 | -2.16e-10 | 78.7085
-7.0486e-12 | 79.1255 | -2.1828e-10 | 79.1135
-6.5938e-12 | 79.5315 | -2.1646e-10 | 79.5175
-5.9117e-12 | 79.9346 | -2.0304e-10 | 79.9206
-6.8212e-12 | 80.3406 | -1.935e-10 | 80.3256
-5.6843e-12 | 80.7456 | -2.2555e-10 | 80.7316
-6.1391e-12 | 81.1516 | -2.0759e-10 | 81.1356
-6.5938e-12 | 81.5557 | -1.8736e-10 | 81.5407
-5.2296e-12 | 81.9597 | -2.0736e-10 | 81.9467
-2.2737e-12 | 82.3647 | -2.0032e-10 | 82.3517
-6.1391e-12 | 82.7687 | -2.1942e-10 | 82.7567
-5.9117e-12 | 83.1738 | -2.1964e-10 | 83.1608
-5.9117e-12 | 83.5788 | -2.0691e-10 | 83.5648
-5.6843e-12 | 83.9838 | -2.2737e-10 | 83.9698
-5.457e-12 | 84.3878
-8.4128e-12 | 84.7929
-5.9117e-12 | 85.1979
-6.1391e-12 | 85.6019
-5.6843e-12 | 86.0069
-6.8212e-12 | 86.4119
-6.1391e-12 | 86.816
-7.9581e-12 | 87.221
-5.2296e-12 | 87.626

## 650uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-7.5416e-12 | 2.9761e-12 | -3.0469e-10 | 2.3527e-11
GEMs On |  | GEMs On
-5.457e-12 | 0.31402 | -3.1559e-10 | 0.31002
-1.1141e-11 | 0.99906 | -3.1059e-10 | 0.99306
-1.2278e-11 | 1.4041 | -3.2333e-10 | 1.3971
-4.7748e-12 | 1.8091 | -3.0036e-10 | 1.8021
-3.638e-12 | 2.2141 | -3.3219e-10 | 2.2081
-8.6402e-12 | 2.6181 | -3.0195e-10 | 2.6132
-1.1596e-11 | 3.0232 | -3.2901e-10 | 3.0202
-5.2296e-12 | 3.4282 | -2.5852e-10 | 3.4252
-2.9559e-12 | 3.8342 | -2.8717e-10 | 3.8302
-5.6843e-12 | 4.2382 | -3.1287e-10 | 4.2342
-9.5497e-12 | 4.6423 | -2.8899e-10 | 4.6383
-7.9581e-12 | 5.0483 | -3.0445e-10 | 5.0433
-4.5475e-12 | 5.4523 | -3.2264e-10 | 5.4493
-3.4106e-12 | 5.8563 | -2.8604e-10 | 5.8533
-9.0949e-12 | 6.2614 | -3.4697e-10 | 6.2594
-9.7771e-12 | 6.6664 | -3.14e-10 | 6.6644
-6.1391e-12 | 7.0704 | -3.1582e-10 | 7.0684
-3.638e-12 | 7.4744 | -2.7831e-10 | 7.4744
-6.1391e-12 | 7.8795 | -2.6807e-10 | 7.8785
-1.1823e-11 | 8.2855 | -3.0195e-10 | 8.2825
-3.4106e-12 | 8.6915 | -3.1673e-10 | 8.6865
-3.1832e-12 | 9.0975 | -3.1173e-10 | 9.0925
-8.1855e-12 | 9.5015 | -2.5648e-10 | 9.4965
-9.0949e-12 | 9.9076 | -2.9172e-10 | 9.9006
-5.457e-12 | 10.3136 | -3.256e-10 | 10.3046
-2.0464e-12 | 10.7176 | -3.2765e-10 | 10.7096
-8.6402e-12 | 11.1226 | -3.7471e-10 | 11.1156
-1.296e-11 | 11.5277 | -2.974e-10 | 11.5207
-1.0687e-11 | 11.9317 | -3.1628e-10 | 11.9267
-5.457e-12 | 12.3367 | -2.7194e-10 | 12.3317
-5.9117e-12 | 12.7407 | -3.1605e-10 | 12.7367
-1.0004e-11 | 13.1457 | -2.7308e-10 | 13.1418
-1.0459e-11 | 13.5508 | -2.7308e-10 | 13.5458
-5.2296e-12 | 13.9548 | -2.9786e-10 | 13.9508
-2.7285e-12 | 14.3588 | -3.0536e-10 | 14.3558
-1.0687e-11 | 14.7648 | -3.0332e-10 | 14.7608
-1.2278e-11 | 15.1709 | -3.2219e-10 | 15.1659
-5.6843e-12 | 15.5759 | -3.0013e-10 | 15.5709
-3.1832e-12 | 15.9799 | -2.8217e-10 | 15.9749
-8.4128e-12 | 16.3859 | -2.9172e-10 | 16.3799
-1.2506e-11 | 16.791 | -3.0923e-10 | 16.785
-6.1391e-12 | 17.197 | -2.8581e-10 | 17.19
-4.5475e-12 | 17.602 | -2.9013e-10 | 17.595
-8.4128e-12 | 18.006 | -3.1878e-10 | 17.999
-1.2506e-11 | 18.411 | -3.1855e-10 | 18.4041
-7.276e-12 | 18.8161 | -2.949e-10 | 18.8101
-3.4106e-12 | 19.2211 | -3.2333e-10 | 19.2141
-4.7748e-12 | 19.6251 | -2.9149e-10 | 19.6191
-1.0687e-11 | 20.0301 | -3.3492e-10 | 20.0241
-8.4128e-12 | 20.4342 | -3.1332e-10 | 20.4292
-4.7748e-12 | 20.8402 | -3.0195e-10 | 20.8332
-2.0464e-12 | 21.2442 | -2.8353e-10 | 21.2372
-7.5033e-12 | 21.6492 | -3.0286e-10 | 21.6422
-1.0914e-11 | 22.0553 | -3.2333e-10 | 22.0473
-4.0927e-12 | 22.4613 | -2.7012e-10 | 22.4513
-2.9559e-12 | 22.8643 | -2.7626e-10 | 22.8553
-7.0486e-12 | 23.2693 | -2.5921e-10 | 23.2593
-1.2278e-11 | 23.6744 | -2.8513e-10 | 23.6643
-1.0459e-11 | 24.0794 | -3.2469e-10 | 24.0694
-2.0464e-12 | 24.4844 | -2.9672e-10 | 24.4734
-6.8212e-12 | 24.8904 | -2.8217e-10 | 24.8774
-1.0687e-11 | 25.2945 | -2.9627e-10 | 25.2814
-9.3223e-12 | 25.6995 | -2.9354e-10 | 25.6865
-3.4106e-12 | 26.1035 | -2.9831e-10 | 26.0915
-5.2296e-12 | 26.5085 | -2.8422e-10 | 26.4955
-8.1855e-12 | 26.9135 | -2.9263e-10 | 26.9005
-1.1141e-11 | 27.3176 | -2.9377e-10 | 27.3056
-4.3201e-12 | 27.7226 | -2.7808e-10 | 27.7096
-2.9559e-12 | 28.1286 | -3.1787e-10 | 28.1136
-1.4779e-11 | 28.5336 | -3.2082e-10 | 28.5186
-9.0949e-12 | 28.9387 | -3.2856e-10 | 28.9227
-4.0927e-12 | 29.3447 | -3.2742e-10 | 29.3287
-5.6843e-12 | 29.7497 | -2.9513e-10 | 29.7337
-1.1823e-11 | 30.1557 | -3.0127e-10 | 30.1377
-1.6144e-11 | 30.5607 | -3.1241e-10 | 30.5418
-6.5938e-12 | 30.9648 | -3.0309e-10 | 30.9468
-3.1832e-12 | 31.3708 | -3.4265e-10 | 31.3518
-1.3188e-11 | 31.7758 | -2.8763e-10 | 31.7558
-1.0914e-11 | 32.1818 | -2.9809e-10 | 32.1598
-6.5938e-12 | 32.5869 | -2.8263e-10 | 32.5639
-5.6843e-12 | 32.9919 | -3.2514e-10 | 32.9689
-9.5497e-12 | 33.3959 | -2.9627e-10 | 33.3729
-1.0914e-11 | 33.8029 | -3.3219e-10 | 33.7779
-5.9117e-12 | 34.208 | -3.0082e-10 | 34.181
-3.638e-12 | 34.612 | -3.5379e-10 | 34.587
-8.6402e-12 | 35.017 | -3.1059e-10 | 34.991
-1.1823e-11 | 35.422 | -3.0286e-10 | 35.395
-6.5938e-12 | 35.827 | -2.9536e-10 | 35.7991
-2.0464e-12 | 36.2321 | -3.181e-10 | 36.2031
-9.3223e-12 | 36.6371 | -3.3333e-10 | 36.6081
-1.0914e-11 | 37.0411 | -3.14e-10 | 37.0121
-8.1855e-12 | 37.4471 | -3.0127e-10 | 37.4161
-2.7285e-12 | 37.8522 | -2.6489e-10 | 37.8222
-5.457e-12 | 38.2572 | -3.0582e-10 | 38.2272
-5.9117e-12 | 38.6612 | -2.9945e-10 | 38.6322
-1.0914e-11 | 39.0652 | -3.0536e-10 | 39.0362
-4.5475e-12 | 39.4703 | -3.0673e-10 | 39.4403
-2.7285e-12 | 39.8753 | -3.1105e-10 | 39.8453
-1.2506e-11 | 40.2813 | -3.5789e-10 | 40.2503
-1.1823e-11 | 40.6843 | -2.774e-10 | 40.6543
-1.2506e-11 | 41.0894 | -3.1628e-10 | 41.0594
-2.7285e-12 | 41.4954 | -2.6466e-10 | 41.4644
-3.4106e-12 | 41.8994 | -3.3651e-10 | 41.8684
-8.4128e-12 | 42.3044 | -3.0013e-10 | 42.2734
-1.2051e-11 | 42.7084 | -3.2219e-10 | 42.6794
-3.4106e-12 | 43.1135 | -2.9104e-10 | 43.0835
-5.9117e-12 | 43.5185 | -3.0536e-10 | 43.4875
-6.1391e-12 | 43.9215 | -3.0309e-10 | 43.8935
-1.387e-11 | 44.3265 | -2.9172e-10 | 44.2985
-1.0914e-11 | 44.7306 | -3.0764e-10 | 44.7026
-4.7748e-12 | 45.1366 | -2.9627e-10 | 45.1076
-3.4106e-12 | 45.5406 | -3.1196e-10 | 45.5116
-1.1141e-11 | 45.9456 | -2.7853e-10 | 45.9156
-6.8212e-12 | 46.3526 | -3.6016e-10 | 46.3197
2.2737e-13 | 46.7577 | -3.1423e-10 | 46.7257
-6.5938e-12 | 47.1627 | -3.2378e-10 | 47.1307
-9.7771e-12 | 47.5677 | -2.9945e-10 | 47.5357
-9.5497e-12 | 47.9727 | -2.7512e-10 | 47.9397
-5.2296e-12 | 48.3768 | -3.4242e-10 | 48.3438
-5.457e-12 | 48.7818 | -3.1832e-10 | 48.7488
-1.1141e-11 | 49.1868 | -3.3015e-10 | 49.1548
-1.0914e-11 | 49.5918 | -3.2992e-10 | 49.5588
-9.7771e-12 | 49.9969 | -3.2423e-10 | 49.9629
-2.0464e-12 | 50.4009 | -3.1582e-10 | 50.3689
-1.0687e-11 | 50.8069 | -2.8967e-10 | 50.7739
-1.2506e-11 | 51.2129 | -3.4129e-10 | 51.1779
-4.7748e-12 | 51.618 | -2.9263e-10 | 51.584
-4.0927e-12 | 52.023 | -3.1537e-10 | 51.988
-7.276e-12 | 52.428 | -3.065e-10 | 52.393
-9.3223e-12 | 52.833 | -3.4879e-10 | 52.797
-8.1855e-12 | 53.237 | -2.6921e-10 | 53.202
-3.4106e-12 | 53.6401 | -3.0059e-10 | 53.6081
-5.9117e-12 | 54.0461 | -2.8604e-10 | 54.0121
-9.3223e-12 | 54.4511 | -3.4834e-10 | 54.4161
-1.1141e-11 | 54.8561 | -2.9013e-10 | 54.8211
-6.5938e-12 | 55.2602 | -2.9786e-10 | 55.2262
-5.6843e-12 | 55.6652 | -3.0309e-10 | 55.6322
-8.4128e-12 | 56.0682 | -2.6853e-10 | 56.0362
-9.3223e-12 | 56.4742 | -2.9013e-10 | 56.4402
-7.9581e-12 | 56.8783 | -2.9831e-10 | 56.8453
-4.3201e-12 | 57.2843 | -3.3219e-10 | 57.2503
-8.6402e-12 | 57.6893 | -3.4652e-10 | 57.6553
-7.9581e-12 | 58.0953 | -3.09e-10 | 58.0593
-4.3201e-12 | 58.5003 | -2.799e-10 | 58.4653
-4.7748e-12 | 58.9044 | -3.5197e-10 | 58.8704
-7.5033e-12 | 59.3094 | -2.6716e-10 | 59.2754
-1.0004e-11 | 59.7134 | -3.2128e-10 | 59.6804
-9.0949e-12 | 60.1194 | -3.4424e-10 | 60.0844
-6.8212e-12 | 60.5235 | -2.9127e-10 | 60.4895
-4.7748e-12 | 60.9285 | -2.7603e-10 | 60.8935
-9.0949e-12 | 61.3335 | -3.0263e-10 | 61.2975
-1.0459e-11 | 61.7375 | -3.2514e-10 | 61.7015
-6.1391e-12 | 62.1435 | -3.09e-10 | 62.1065
-6.5938e-12 | 62.5496 | -2.7444e-10 | 62.5116
-9.0949e-12 | 62.9536 | -3.2742e-10 | 62.9166
-9.7771e-12 | 63.3566 | -3.1014e-10 | 63.3206
-7.0486e-12 | 63.7617 | -2.7944e-10 | 63.7257
-5.9117e-12 | 64.1677 | -3.3606e-10 | 64.1317
-1.1141e-11 | 64.5727 | -3.3651e-10 | 64.5367
-9.7771e-12 | 64.9777 | -2.7694e-10 | 64.9417
-7.0486e-12 | 65.3817 | -2.9513e-10 | 65.3477
-6.1391e-12 | 65.7868 | -2.9308e-10 | 65.7528
-5.6843e-12 | 66.1918 | -3.0036e-10 | 66.1568
-8.6402e-12 | 66.5968 | -3.3856e-10 | 66.5608
-9.5497e-12 | 67.0008 | -2.5761e-10 | 66.9658
-1.0687e-11 | 67.4059 | -2.974e-10 | 67.3718
-6.8212e-12 | 67.8129 | -3.2105e-10 | 67.7759
-1.1141e-11 | 68.2169 | -3.1787e-10 | 68.1809
-7.5033e-12 | 68.6229 | -3.0218e-10 | 68.5839
-6.5938e-12 | 69.0269 | -2.5193e-10 | 68.989
-7.5033e-12 | 69.433 | -2.9672e-10 | 69.392
-8.4128e-12 | 69.839 | -3.6107e-10 | 69.798
-6.8212e-12 | 70.243 | -3.231e-10 | 70.203
-7.0486e-12 | 70.648 | -3.3492e-10 | 70.607
-5.457e-12 | 71.0531 | -3.1832e-10 | 71.0111
-8.1855e-12 | 71.4581 | -2.9354e-10 | 71.4141
-6.1391e-12 | 71.8641 | -2.6489e-10 | 71.8201
-5.9117e-12 | 72.2681 | -2.633e-10 | 72.2251
-7.9581e-12 | 72.6732 | -3.065e-10 | 72.6302
-1.0004e-11 | 73.0772 | -3.1537e-10 | 73.0362
-1.0004e-11 | 73.4822 | -2.8808e-10 | 73.4402
-6.5938e-12 | 73.8872 | -2.9013e-10 | 73.8452
-8.1855e-12 | 74.2922 | -2.9399e-10 | 74.2502
-1.5007e-11 | 74.6973 | -2.8126e-10 | 74.6553
-9.3223e-12 | 75.1023 | -2.9377e-10 | 75.0613
-6.1391e-12 | 75.5073 | -2.8035e-10 | 75.4663
-6.1391e-12 | 75.9133 | -3.1559e-10 | 75.8713
-6.1391e-12 | 76.3184 | -3.3265e-10 | 76.2764
-8.1855e-12 | 76.7234 | -2.8012e-10 | 76.6814
-6.8212e-12 | 77.1284 | -2.7308e-10 | 77.0844
-6.1391e-12 | 77.5324 | -3.1991e-10 | 77.4894
-7.5033e-12 | 77.9375 | -3.1332e-10 | 77.8945
-7.0486e-12 | 78.3415 | -3.1059e-10 | 78.2995
-5.9117e-12 | 78.7455 | -2.7899e-10 | 78.7045
-8.1855e-12 | 79.1525 | -3.0627e-10 | 79.1095
-9.0949e-12 | 79.5566 | -3.2128e-10 | 79.5145
-1.1596e-11 | 79.9616 | -3.5584e-10 | 79.9196
-6.8212e-12 | 80.3676 | -3.1991e-10 | 80.3246
-5.6843e-12 | 80.7706 | -2.4738e-10 | 80.7286
-5.9117e-12 | 81.1746 | -2.7057e-10 | 81.1336
-9.3223e-12 | 81.5787 | -3.4015e-10 | 81.5387
-7.9581e-12 | 81.9837 | -2.8331e-10 | 81.9427
-5.9117e-12 | 82.3877 | -2.9831e-10 | 82.3477
-7.276e-12 | 82.7917 | -2.949e-10 | 82.7517
-8.4128e-12 | 83.1968 | -3.0946e-10 | 83.1568
-5.9117e-12 | 83.6008
-9.0949e-12 | 84.0058
-5.9117e-12 | 84.4108
-5.6843e-12 | 84.8149
-9.7771e-12 | 85.2199
-8.6402e-12 | 85.6249
-5.9117e-12 | 86.0299
-7.0486e-12 | 86.4339
-9.7771e-12 | 86.838

## 660uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-7.9804e-12 | 2.5856e-12 | -4.3111e-10 | 3.3603e-11
GEMs On |  | GEMs On
-8.4128e-12 | 0.31402 | -4.2587e-10 | 0.31602
-5.6843e-12 | 0.99806 | -3.9199e-10 | 1.0011
-9.3223e-12 | 1.4031 | -5.0295e-10 | 1.4051
-7.276e-12 | 1.8081 | -4.5225e-10 | 1.8101
-5.9117e-12 | 2.2141 | -4.4315e-10 | 2.2151
-5.9117e-12 | 2.6191 | -4.4588e-10 | 2.6181
-8.4128e-12 | 3.0232 | -4.386e-10 | 3.0232
-7.5033e-12 | 3.4272 | -4.5316e-10 | 3.4362
-9.3223e-12 | 3.8322 | -4.9567e-10 | 3.8402
-7.276e-12 | 4.2382 | -4.1746e-10 | 4.2442
-9.3223e-12 | 4.6443 | -4.5793e-10 | 4.6503
-8.4128e-12 | 5.0483 | -4.4929e-10 | 5.0553
-6.8212e-12 | 5.4543 | -4.4565e-10 | 5.4943
-5.9117e-12 | 5.8593 | -3.6471e-10 | 5.9003
-1.2051e-11 | 6.2644 | -4.195e-10 | 6.3054
-8.1855e-12 | 6.6704 | -3.6812e-10 | 6.7104
-8.1855e-12 | 7.0754 | -4.4429e-10 | 7.1144
-7.5033e-12 | 7.4804 | -4.6566e-10 | 7.5184
-8.6402e-12 | 7.8845 | -3.7721e-10 | 7.9345
-7.5033e-12 | 8.2895 | -3.4129e-10 | 8.3385
-7.5033e-12 | 8.6935 | -4.5657e-10 | 8.7815
-6.3665e-12 | 9.0975 | -5.2546e-10 | 9.2065
-9.7771e-12 | 9.5015 | -3.9722e-10 | 9.6196
-6.8212e-12 | 9.9066 | -4.7589e-10 | 10.0246
-8.1855e-12 | 10.3116 | -4.0336e-10 | 10.4436
-1.0004e-11 | 10.7156 | -4.4793e-10 | 10.8496
-8.4128e-12 | 11.1216 | -4.236e-10 | 11.2586
-8.6402e-12 | 11.5267 | -4.3428e-10 | 11.6697
-7.0486e-12 | 11.9337 | -4.3519e-10 | 12.0797
-7.276e-12 | 12.3377 | -4.0313e-10 | 12.4847
-6.8212e-12 | 12.7427 | -4.5588e-10 | 12.8897
-7.276e-12 | 13.1468 | -4.7476e-10 | 13.2958
-1.6598e-11 | 13.5518 | -4.261e-10 | 13.7008
-8.1855e-12 | 13.9568 | -4.3156e-10 | 14.1068
-7.0486e-12 | 14.3608 | -4.2633e-10 | 14.5148
-8.6402e-12 | 14.7658 | -4.286e-10 | 14.9208
-1.0232e-11 | 15.1699 | -4.3133e-10 | 15.3259
-5.457e-12 | 15.5769 | -4.2974e-10 | 15.7309
-6.1391e-12 | 15.9819 | -4.145e-10 | 16.1359
-9.3223e-12 | 16.3869 | -4.3951e-10 | 16.5409
-9.3223e-12 | 16.791 | -4.643e-10 | 16.948
-5.9117e-12 | 17.197 | -4.1814e-10 | 17.354
-4.7748e-12 | 17.602 | -4.6339e-10 | 17.759
-1.0232e-11 | 18.008 | -4.286e-10 | 18.163
-1.3188e-11 | 18.413 | -4.5225e-10 | 18.5681
-2.5011e-12 | 18.8181 | -4.7203e-10 | 18.9741
-5.9117e-12 | 19.2221 | -4.5793e-10 | 19.3781
-1.1141e-11 | 19.6281 | -4.1132e-10 | 19.7821
-8.8676e-12 | 20.0342 | -4.0268e-10 | 20.1871
-4.7748e-12 | 20.4382 | -4.5452e-10 | 20.5902
-7.0486e-12 | 20.8442 | -4.7589e-10 | 20.9972
-1.0004e-11 | 21.2492 | -4.2132e-10 | 21.4002
-1.0232e-11 | 21.6532 | -4.1723e-10 | 21.8062
-6.3665e-12 | 22.0583 | -4.2382e-10 | 22.2103
-8.1855e-12 | 22.4643 | -4.0222e-10 | 22.6173
-8.8676e-12 | 22.8683 | -4.6612e-10 | 23.0213
-1.1823e-11 | 23.2733 | -4.2928e-10 | 23.4263
-8.6402e-12 | 23.6783 | -4.1859e-10 | 23.8304
-5.0022e-12 | 24.0834 | -4.3701e-10 | 24.2354
-3.638e-12 | 24.4874 | -4.2451e-10 | 24.6404
-1.2278e-11 | 24.8914 | -4.1041e-10 | 25.0454
-8.8676e-12 | 25.2975 | -4.6953e-10 | 25.4505
-4.3201e-12 | 25.7025 | -4.6998e-10 | 25.8535
-7.5033e-12 | 26.1055 | -3.854e-10 | 26.2575
-1.0687e-11 | 26.5095 | -4.4065e-10 | 26.6625
-1.1823e-11 | 26.9145 | -3.5584e-10 | 27.0676
-5.0022e-12 | 27.3206 | -4.2519e-10 | 27.4726
-5.0022e-12 | 27.7246 | -3.9722e-10 | 27.8776
-7.5033e-12 | 28.1306 | -4.3929e-10 | 28.2816
-1.2506e-11 | 28.5346 | -4.3678e-10 | 28.6866
-9.7771e-12 | 28.9407 | -4.1882e-10 | 29.0917
-5.0022e-12 | 29.3447 | -4.3178e-10 | 29.4957
-8.1855e-12 | 29.7497 | -4.4497e-10 | 29.9007
-1.2278e-11 | 30.1547 | -4.2769e-10 | 30.3067
-1.0914e-11 | 30.5598 | -4.1587e-10 | 30.7118
-4.7748e-12 | 30.9648 | -4.1859e-10 | 31.1158
-7.0486e-12 | 31.3698 | -4.1109e-10 | 31.5198
-1.0914e-11 | 31.7738 | -3.9017e-10 | 31.9238
-8.6402e-12 | 32.1788 | -4.5748e-10 | 32.3289
-8.4128e-12 | 32.5839 | -4.4497e-10 | 32.7339
-1.0232e-11 | 32.9879 | -3.904e-10 | 33.1389
-7.5033e-12 | 33.3919 | -4.3815e-10 | 33.5449
-1.0232e-11 | 33.7969 | -4.3246e-10 | 33.9499
-1.2051e-11 | 34.202 | -4.3269e-10 | 34.356
-3.8654e-12 | 34.608 | -4.7339e-10 | 34.761
-6.3665e-12 | 35.012 | -4.0654e-10 | 35.166
-1.1823e-11 | 35.417 | -4.6771e-10 | 35.571
-8.1855e-12 | 35.822 | -5.0818e-10 | 35.9751
-3.8654e-12 | 36.2281 | -4.8863e-10 | 36.3801
-5.0022e-12 | 36.6321 | -4.4793e-10 | 36.7841
-5.9117e-12 | 37.0381 | -4.1996e-10 | 37.1901
-6.1391e-12 | 37.4421 | -4.4383e-10 | 37.5951
-5.457e-12 | 37.8472 | -4.8612e-10 | 37.9992
-4.5475e-12 | 38.2522 | -4.0336e-10 | 38.4032
-9.7771e-12 | 38.6562 | -4.3519e-10 | 38.8072
-1.0914e-11 | 39.0612 | -4.045e-10 | 39.2122
-3.4106e-12 | 39.4663 | -4.7589e-10 | 39.6183
-2.2737e-12 | 39.8723 | -4.1177e-10 | 40.0243
-9.7771e-12 | 40.2763 | -3.688e-10 | 40.4283
-1.3642e-11 | 40.6813 | -4.4088e-10 | 40.8333
-6.3665e-12 | 41.0864 | -4.145e-10 | 41.2384
-2.5011e-12 | 41.4914 | -4.2542e-10 | 41.6414
-8.4128e-12 | 41.8964 | -4.0541e-10 | 42.0474
-1.2506e-11 | 42.3004 | -4.6862e-10 | 42.4514
-5.6843e-12 | 42.7064 | -4.0086e-10 | 42.8565
-3.4106e-12 | 43.1115 | -3.8517e-10 | 43.2615
-6.3665e-12 | 43.5165 | -4.7362e-10 | 43.6655
-1.1823e-11 | 43.9215 | -4.8044e-10 | 44.0705
-6.3665e-12 | 44.3275 | -4.4383e-10 | 44.4745
-8.4128e-12 | 44.7326 | -4.4383e-10 | 44.8796
-8.1855e-12 | 45.1386 | -4.0541e-10 | 45.2836
-1.0232e-11 | 45.5436 | -3.9518e-10 | 45.6886
-5.457e-12 | 45.9496 | -4.6293e-10 | 46.0916
-4.5475e-12 | 46.3537 | -5.1455e-10 | 46.4967
-5.6843e-12 | 46.7577 | -4.4247e-10 | 46.9007
-1.1369e-11 | 47.1627 | -4.0586e-10 | 47.3067
-9.3223e-12 | 47.5677 | -4.7521e-10 | 47.7117
-4.7748e-12 | 47.9727 | -3.929e-10 | 48.1168
-7.276e-12 | 48.3768 | -4.7294e-10 | 48.5208
-1.1369e-11 | 48.7828 | -3.7085e-10 | 48.9258
-7.9581e-12 | 49.1878 | -4.4793e-10 | 49.3308
-8.1855e-12 | 49.5918 | -4.618e-10 | 49.7338
-4.5475e-12 | 49.9959 | -4.643e-10 | 50.1379
-9.3223e-12 | 50.4009 | -4.6316e-10 | 50.5419
-1.0914e-11 | 50.8069 | -4.2337e-10 | 50.9459
-7.276e-12 | 51.2119 | -4.3383e-10 | 51.3519
-5.0022e-12 | 51.617 | -3.4515e-10 | 51.757
-8.1855e-12 | 52.022 | -4.5998e-10 | 52.163
-9.5497e-12 | 52.427 | -3.5652e-10 | 52.568
-3.638e-12 | 52.832 | -4.7294e-10 | 52.972
-3.8654e-12 | 53.2381 | -4.1064e-10 | 53.3781
-9.3223e-12 | 53.6431 | -4.1109e-10 | 53.7821
-1.1141e-11 | 54.0481 | -4.7999e-10 | 54.1871
-6.8212e-12 | 54.4541 | -4.3883e-10 | 54.5931
-7.0486e-12 | 54.8581 | -3.9836e-10 | 54.9971
-9.3223e-12 | 55.2642 | -4.5111e-10 | 55.4032
-1.0232e-11 | 55.6692 | -4.0723e-10 | 55.8072
-5.9117e-12 | 56.0752 | -4.2132e-10 | 56.2102
-7.9581e-12 | 56.4792 | -4.9954e-10 | 56.6152
-4.3201e-12 | 56.8843 | -3.7835e-10 | 57.0203
-1.0687e-11 | 57.2883 | -4.5611e-10 | 57.4243
-7.5033e-12 | 57.6933 | -4.4179e-10 | 57.8293
-5.457e-12 | 58.0973 | -3.979e-10 | 58.2343
-7.0486e-12 | 58.5023 | -4.7635e-10 | 58.6384
-8.6402e-12 | 58.9074 | -4.2905e-10 | 59.0424
-8.6402e-12 | 59.3124 | -4.2724e-10 | 59.4464
-6.1391e-12 | 59.7164 | -4.4588e-10 | 59.8504
-6.1391e-12 | 60.1214 | -4.7658e-10 | 60.2555
-9.7771e-12 | 60.5255 | -4.4815e-10 | 60.6605
-8.4128e-12 | 60.9305 | -3.5493e-10 | 61.0645
-8.1855e-12 | 61.3355 | -3.7539e-10 | 61.4695
-6.8212e-12 | 61.7405 | -3.9063e-10 | 61.8745
-5.457e-12 | 62.1446 | -4.2587e-10 | 62.2796
-1.0687e-11 | 62.5476 | -3.9381e-10 | 62.6846
-9.5497e-12 | 62.9526 | -4.5407e-10 | 63.0886
-9.5497e-12 | 63.3566 | -4.2519e-10 | 63.4936
-7.276e-12 | 63.7617 | -4.3701e-10 | 63.8987
-9.3223e-12 | 64.1667 | -4.2132e-10 | 64.3037
-1.0914e-11 | 64.5727 | -3.9427e-10 | 64.7087
-5.457e-12 | 64.9767 | -4.0882e-10 | 65.1137
-5.457e-12 | 65.3817 | -4.2542e-10 | 65.5198
-5.0022e-12 | 65.7858 | -4.145e-10 | 65.9228
-1.1823e-11 | 66.1898 | -4.4406e-10 | 66.3268
-9.5497e-12 | 66.5948 | -4.3633e-10 | 66.7328
-6.3665e-12 | 66.9988 | -4.3929e-10 | 67.1378
-4.7748e-12 | 67.4039 | -3.954e-10 | 67.5419
-1.0687e-11 | 67.8079 | -4.2064e-10 | 67.9469
-1.1823e-11 | 68.2139 | -4.7476e-10 | 68.3529
-7.5033e-12 | 68.6199 | -3.7153e-10 | 68.7589
-6.8212e-12 | 69.0259 | -4.1041e-10 | 69.163
-1.387e-11 | 69.431 | -4.286e-10 | 69.567
-8.6402e-12 | 69.837 | -4.5702e-10 | 69.971
-3.8654e-12 | 70.241 | -4.361e-10 | 70.375
-3.638e-12 | 70.645 | -4.7794e-10 | 70.779
-8.4128e-12 | 71.0501 | -5.0932e-10 | 71.1841
-9.5497e-12 | 71.4541 | -4.0109e-10 | 71.5891
-3.4106e-12 | 71.8591 | -3.904e-10 | 71.9931
-4.3201e-12 | 72.2651 | -4.3906e-10 | 72.3961
-9.5497e-12 | 72.6692 | -4.4201e-10 | 72.8022
-9.5497e-12 | 73.0742 | -4.3906e-10 | 73.2072
-5.457e-12 | 73.4802 | -4.6452e-10 | 73.6122
-4.7748e-12 | 73.8852 | -4.1291e-10 | 74.0162
-7.0486e-12 | 74.2903 | -4.7703e-10 | 74.4203
-1.1369e-11 | 74.6943 | -4.1496e-10 | 74.8243
-7.5033e-12 | 75.0993 | -4.3224e-10 | 75.2293
-7.276e-12 | 75.5033 | -4.07e-10 | 75.6333
-5.457e-12 | 75.9093 | -4.2974e-10 | 76.0383
-9.7771e-12 | 76.3144 | -4.0291e-10 | 76.4424
-8.8676e-12 | 76.7194 | -4.1746e-10 | 76.8474
-6.1391e-12 | 77.1244 | -4.5748e-10 | 77.2534
-1.3415e-11 | 77.5294 | -3.7289e-10 | 77.6584
-1.2278e-11 | 77.9345 | -4.1496e-10 | 78.0625
-8.8676e-12 | 78.3395 | -4.2587e-10 | 78.4665
-7.5033e-12 | 78.7465 | -3.8563e-10 | 78.8715
-5.457e-12 | 79.1505 | -4.2269e-10 | 79.2765
-1.3188e-11 | 79.5536 | -4.3042e-10 | 79.6816
-8.6402e-12 | 79.9596 | -4.4542e-10 | 80.0866
-6.8212e-12 | 80.3656 | -3.9586e-10 | 80.4916
-1.0232e-11 | 80.7706 | -4.1382e-10 | 80.8976
-1.1823e-11 | 81.1766 | -4.5088e-10 | 81.3026
-9.7771e-12 | 81.5817 | -3.588e-10 | 81.7077
-8.6402e-12 | 81.9867 | -4.7089e-10 | 82.1127
-7.5033e-12 | 82.3917 | -4.1814e-10 | 82.5177
-1.1141e-11 | 82.7967 | -4.1496e-10 | 82.9227

## 670uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.0238e-11 | 2.545e-12 | -6.1518e-10 | 4.7551e-11
GEMs On |  | GEMs On
-9.3223e-12 | 0.31502 | -6.0686e-10 | 0.30902
-6.8212e-12 | 0.99806 | -5.8753e-10 | 0.99306
-8.1855e-12 | 1.4041 | -6.7325e-10 | 1.3991
-8.6402e-12 | 1.8101 | -6.2664e-10 | 1.8031
-6.3665e-12 | 2.2151 | -5.7184e-10 | 2.2081
-4.7748e-12 | 2.6212 | -5.4297e-10 | 2.6132
-1.0004e-11 | 3.0252 | -5.8662e-10 | 3.0182
-1.0687e-11 | 3.4312 | -6.5643e-10 | 3.4222
-1.0459e-11 | 3.8362 | -5.4683e-10 | 3.8282
-5.9117e-12 | 4.2412 | -6.7826e-10 | 4.2332
-1.1369e-11 | 4.6463 | -5.5184e-10 | 4.6383
-1.296e-11 | 5.0523 | -6.3073e-10 | 5.0423
-7.9581e-12 | 5.4583 | -6.7303e-10 | 5.4473
-1.2733e-11 | 5.8693 | -5.4297e-10 | 5.8523
-9.7771e-12 | 6.2744 | -6.7325e-10 | 6.2564
-8.6402e-12 | 6.6774 | -6.1095e-10 | 6.6604
-8.6402e-12 | 7.0814 | -6.5143e-10 | 7.0654
-9.0949e-12 | 7.4864 | -7.0372e-10 | 7.4704
-9.3223e-12 | 7.8915 | -6.2437e-10 | 7.8755
-1.0687e-11 | 8.2965 | -5.8458e-10 | 8.2805
-6.1391e-12 | 8.7015 | -6.2505e-10 | 8.6845
-9.3223e-12 | 9.1055 | -6.2209e-10 | 9.0885
-1.0459e-11 | 9.5115 | -6.3756e-10 | 9.4935
-7.7307e-12 | 9.9166 | -6.7621e-10 | 9.8986
-1.0232e-11 | 10.3216 | -6.5779e-10 | 10.3036
-1.2506e-11 | 10.7276 | -6.1391e-10 | 10.7076
-1.0004e-11 | 11.1326 | -6.2892e-10 | 11.1126
-7.276e-12 | 11.5387 | -6.1527e-10 | 11.5167
-9.3223e-12 | 11.9427 | -6.4256e-10 | 11.9217
-9.0949e-12 | 12.3477 | -6.3756e-10 | 12.3257
-1.0232e-11 | 12.7547 | -6.039e-10 | 12.7297
-1.1141e-11 | 13.1598 | -6.1164e-10 | 13.1357
-8.1855e-12 | 13.5638 | -5.9936e-10 | 13.5388
-9.0949e-12 | 13.9688 | -5.9845e-10 | 13.9438
-1.0004e-11 | 14.3728 | -6.3505e-10 | 14.3488
-9.0949e-12 | 14.7779 | -6.8098e-10 | 14.7538
-1.0687e-11 | 15.1829 | -5.4933e-10 | 15.1589
-7.5033e-12 | 15.5889 | -6.7166e-10 | 15.5649
-7.7307e-12 | 15.9939 | -6.1027e-10 | 15.9689
-2.0464e-11 | 16.3989 | -5.8139e-10 | 16.3729
-1.0232e-11 | 16.805 | -6.23e-10 | 16.778
-1.0004e-11 | 17.21 | -6.0459e-10 | 17.184
-1.0687e-11 | 17.622 | -6.014e-10 | 17.59
-8.6402e-12 | 18.028 | -5.8276e-10 | 17.995
-1.0004e-11 | 18.4331 | -5.7025e-10 | 18.4001
-1.0004e-11 | 18.8381 | -6.9144e-10 | 18.8051
-9.3223e-12 | 19.2431 | -5.3751e-10 | 19.2101
-1.0232e-11 | 19.6481 | -6.2232e-10 | 19.6141
-1.0687e-11 | 20.0541 | -5.6775e-10 | 20.0191
-1.1369e-11 | 20.4602 | -6.446e-10 | 20.4232
-1.0687e-11 | 20.8652 | -6.4551e-10 | 20.8272
-7.5033e-12 | 21.2712 | -6.3483e-10 | 21.2342
-2.3192e-11 | 21.6742 | -6.5234e-10 | 21.6382
-1.1369e-11 | 22.0803 | -5.9504e-10 | 22.0423
-9.3223e-12 | 22.4863 | -5.6684e-10 | 22.4473
-1.1596e-11 | 22.8903 | -6.1937e-10 | 22.8503
-1.0232e-11 | 23.2953 | -5.3728e-10 | 23.2553
-1.1141e-11 | 23.7004 | -5.2887e-10 | 23.6593
-9.3223e-12 | 24.1054 | -5.9049e-10 | 24.0644
-8.6402e-12 | 24.5114 | -5.8048e-10 | 24.4704
-1.1823e-11 | 24.9164 | -6.6143e-10 | 24.8744
-9.3223e-12 | 25.3224 | -5.4979e-10 | 25.2795
-7.5033e-12 | 25.7275 | -6.3437e-10 | 25.6845
-8.6402e-12 | 26.1325 | -6.155e-10 | 26.0885
-1.1141e-11 | 26.5385 | -6.9463e-10 | 26.4925
-9.0949e-12 | 26.9435 | -5.5502e-10 | 26.8965
-7.9581e-12 | 27.3496 | -6.0868e-10 | 27.3026
-1.0004e-11 | 27.7556 | -6.3233e-10 | 27.7066
-1.5461e-11 | 28.1606 | -5.6548e-10 | 28.1106
-1.0232e-11 | 28.5666 | -5.7071e-10 | 28.5156
-7.7307e-12 | 28.9717 | -6.2255e-10 | 28.9206
-9.7771e-12 | 29.3757 | -6.6075e-10 | 29.3247
-9.7771e-12 | 29.7807 | -6.1596e-10 | 29.7297
-9.3223e-12 | 30.1847 | -5.3251e-10 | 30.1347
-9.3223e-12 | 30.5897 | -5.5661e-10 | 30.5407
-1.0232e-11 | 30.9938 | -6.2414e-10 | 30.9448
-1.1369e-11 | 31.3998 | -6.8781e-10 | 31.3498
-1.0459e-11 | 31.8048 | -5.6934e-10 | 31.7538
-1.0232e-11 | 32.2088 | -5.8276e-10 | 32.1598
-1.0232e-11 | 32.6139 | -7.7353e-10 | 32.5629
-9.0949e-12 | 33.0189 | -5.9208e-10 | 32.9679
-1.1596e-11 | 33.4239 | -6.1209e-10 | 33.3719
-8.8676e-12 | 33.8289 | -6.4892e-10 | 33.7769
-8.8676e-12 | 34.235 | -5.9117e-10 | 34.181
-1.2051e-11 | 34.641 | -5.914e-10 | 34.585
-1.0004e-11 | 35.046 | -6.8985e-10 | 34.991
-1.0004e-11 | 35.451 | -6.1709e-10 | 35.394
-1.0232e-11 | 35.8571 | -5.7025e-10 | 35.7991
-1.0687e-11 | 36.2621 | -6.7348e-10 | 36.2041
-1.0459e-11 | 36.6681 | -5.9345e-10 | 36.6091
-9.3223e-12 | 37.0731 | -5.6957e-10 | 37.0131
-1.0459e-11 | 37.4791 | -6.0709e-10 | 37.4181
-1.0232e-11 | 37.8842 | -6.7348e-10 | 37.8232
-1.0687e-11 | 38.2892 | -5.648e-10 | 38.2282
-1.2733e-11 | 38.6932 | -6.6052e-10 | 38.6332
-8.8676e-12 | 39.0982 | -5.4933e-10 | 39.0372
-7.7307e-12 | 39.5023 | -6.4415e-10 | 39.4423
-9.7771e-12 | 39.9083 | -6.1209e-10 | 39.8473
-7.9581e-12 | 40.3123 | -5.839e-10 | 40.2533
-8.6402e-12 | 40.7163 | -6.4847e-10 | 40.6573
-9.7771e-12 | 41.1213 | -6.5529e-10 | 41.0613
-9.0949e-12 | 41.5264 | -6.9144e-10 | 41.4674
-1.1141e-11 | 41.9304 | -5.7639e-10 | 41.8714
-9.7771e-12 | 42.3354 | -5.7344e-10 | 42.2764
-1.1141e-11 | 42.7404 | -5.914e-10 | 42.6804
-8.8676e-12 | 43.1455 | -6.7303e-10 | 43.0855
-9.0949e-12 | 43.5485 | -5.8003e-10 | 43.4905
-1.0004e-11 | 43.9535 | -5.773e-10 | 43.8955
-1.5234e-11 | 44.3595 | -6.7189e-10 | 44.2995
-1.819e-11 | 44.7676 | -6.3619e-10 | 44.7036
-1.0459e-11 | 45.1726 | -5.6821e-10 | 45.1096
-1.1141e-11 | 45.5776 | -7.6534e-10 | 45.5136
-1.0004e-11 | 45.9826 | -6.0777e-10 | 45.9186
-9.7771e-12 | 46.3877 | -6.3665e-10 | 46.3227
-1.0459e-11 | 46.7937 | -6.9508e-10 | 46.7267
-7.9581e-12 | 47.1997 | -6.246e-10 | 47.1317
-1.296e-11 | 47.6047 | -6.9394e-10 | 47.5357
-1.1369e-11 | 48.0097 | -5.748e-10 | 47.9407
-1.0459e-11 | 48.4158 | -6.1436e-10 | 48.3448
-1.0004e-11 | 48.8208 | -6.6325e-10 | 48.7498
-1.1369e-11 | 49.2258 | -6.6052e-10 | 49.1548
-1.1823e-11 | 49.6318 | -6.321e-10 | 49.5588
-9.7771e-12 | 50.0369 | -6.2983e-10 | 49.9639
-9.0949e-12 | 50.4429 | -5.6184e-10 | 50.3679
-1.0004e-11 | 50.8469 | -6.0527e-10 | 50.7719
-9.3223e-12 | 51.2539 | -6.08e-10 | 51.1769
-1.0687e-11 | 51.659 | -5.6275e-10 | 51.581
-7.9581e-12 | 52.064 | -5.8481e-10 | 51.985
-8.8676e-12 | 52.469 | -5.7366e-10 | 52.39
-1.0687e-11 | 52.873 | -5.5502e-10 | 52.794
-1.6598e-11 | 53.2771 | -6.0777e-10 | 53.199
-7.276e-12 | 53.6821 | -5.7798e-10 | 53.6041
-1.0004e-11 | 54.0861 | -5.8503e-10 | 54.0091
-1.1823e-11 | 54.4911 | -6.587e-10 | 54.4151
-1.0232e-11 | 54.8961 | -7.0872e-10 | 54.8191
-7.276e-12 | 55.3022 | -5.1409e-10 | 55.2242
-1.2506e-11 | 55.7072 | -5.6957e-10 | 55.6282
-8.1855e-12 | 56.1132 | -7.6398e-10 | 56.0332
-1.6598e-11 | 56.5182 | -6.7985e-10 | 56.4382
-1.0232e-11 | 56.9233 | -6.6257e-10 | 56.8413
-1.1369e-11 | 57.3293 | -6.1141e-10 | 57.2453
-7.5033e-12 | 57.7363 | -6.1573e-10 | 57.6513
-8.8676e-12 | 58.1413 | -6.2164e-10 | 58.0543
-1.7508e-11 | 58.5463 | -5.4138e-10 | 58.4603
-1.2051e-11 | 58.9524 | -6.4506e-10 | 58.8644
-7.5033e-12 | 59.3584 | -5.5707e-10 | 59.2704
-8.6402e-12 | 59.7634 | -5.3819e-10 | 59.6744
-1.1596e-11 | 60.1684 | -6.4529e-10 | 60.0794
-1.0687e-11 | 60.5735 | -5.5843e-10 | 60.4845
-1.1369e-11 | 60.9795 | -5.9913e-10 | 60.8875
-1.1596e-11 | 61.3865 | -6.0436e-10 | 61.2935
-1.1141e-11 | 61.7915 | -6.0027e-10 | 61.6975
-1.1141e-11 | 62.1976 | -5.5911e-10 | 62.1035
-1.3188e-11 | 62.6026 | -5.6775e-10 | 62.5096
-1.2051e-11 | 63.0076 | -7.2555e-10 | 62.9146
-1.0459e-11 | 63.4126 | -6.6666e-10 | 63.3196
-6.8212e-12 | 63.8167 | -6.2346e-10 | 63.7246
-7.7307e-12 | 64.2227 | -6.0413e-10 | 64.1277
-1.6371e-11 | 64.6267 | -5.9617e-10 | 64.5327
-1.0004e-11 | 65.0327 | -6.0163e-10 | 64.9367
-5.2296e-12 | 65.4377 | -6.0118e-10 | 65.3407
-8.6402e-12 | 65.8418 | -6.1846e-10 | 65.7458
-1.1596e-11 | 66.2458 | -6.8326e-10 | 66.1518
-9.7771e-12 | 66.6508 | -6.4392e-10 | 66.5558
-1.0004e-11 | 67.0568 | -6.2482e-10 | 66.9588
-1.1823e-11 | 67.4629 | -6.3528e-10 | 67.3638
-1.2506e-11 | 67.8679 | -6.346e-10 | 67.7679
-7.5033e-12 | 68.2729 | -5.7889e-10 | 68.1729
-6.5938e-12 | 68.6779 | -5.9845e-10 | 68.5769
-1.2506e-11 | 69.082 | -6.08e-10 | 68.9819
-1.1141e-11 | 69.488 | -6.0777e-10 | 69.386
-7.9581e-12 | 69.894 | -6.0641e-10 | 69.79
-8.8676e-12 | 70.299 | -5.8208e-10 | 70.194
-1.296e-11 | 70.703 | -5.9254e-10 | 70.598
-1.0687e-11 | 71.1081 | -6.4915e-10 | 71.0041
-7.5033e-12 | 71.5131 | -5.5638e-10 | 71.4081
-7.7307e-12 | 71.9171 | -5.8048e-10 | 71.8131
-1.1823e-11 | 72.3201 | -5.7548e-10 | 72.2171
-1.3188e-11 | 72.7252 | -6.0572e-10 | 72.6222
-8.6402e-12 | 73.1302 | -5.9504e-10 | 73.0272
-5.2296e-12 | 73.5352 | -6.0959e-10 | 73.4322
-1.1369e-11 | 73.9402 | -5.8367e-10 | 73.8362
-1.5916e-11 | 74.3462 | -6.1686e-10 | 74.2403
-4.3201e-12 | 74.7523 | -6.1914e-10 | 74.6453
-7.9581e-12 | 75.1563 | -6.1596e-10 | 75.0503
-1.4325e-11 | 75.5603 | -6.9213e-10 | 75.4563
-1.2733e-11 | 75.9664 | -7.0304e-10 | 75.8613
-6.8212e-12 | 76.3724 | -6.8576e-10 | 76.2644
-8.8676e-12 | 76.7764 | -6.1277e-10 | 76.6684
-1.4552e-11 | 77.1804 | -6.3255e-10 | 77.0744
-1.2051e-11 | 77.5854 | -7.0031e-10 | 77.4794
-9.0949e-12 | 77.9905 | -6.8189e-10 | 77.8834
-8.8676e-12 | 78.3955 | -6.2323e-10 | 78.2875
-1.3642e-11 | 78.8005 | -5.8185e-10 | 78.6925
-7.5033e-12 | 79.2065 | -5.3501e-10 | 79.0985
-7.276e-12 | 79.6115 | -6.4642e-10 | 79.5025
-1.5007e-11 | 80.0156 | -5.7821e-10 | 79.9066
-1.3642e-11 | 80.4206 | -5.7958e-10 | 80.3116
-7.9581e-12 | 80.8266 | -5.2341e-10 | 80.7156
-9.7771e-12 | 81.2327 | -6.1095e-10 | 81.1206
-1.0004e-11 | 81.6367 | -5.8662e-10 | 81.5247
-1.5689e-11 | 82.0427 | -5.4024e-10 | 81.9307
-9.0949e-12 | 82.4467 | -6.562e-10 | 82.3357
-8.8676e-12 | 82.8517 | -6.2209e-10 | 82.7417
-1.3188e-11 | 83.2558
-1.3188e-11 | 83.6618

## 680uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.2225e-11 | 3.4025e-12 | -8.775e-10 | 6.9002e-11
GEMs On |  | GEMs On
-9.0949e-12 | 0.30902 | -8.2468e-10 | 0.31302
-1.5689e-11 | 0.99406 | -8.9926e-10 | 0.99806
-1.296e-11 | 1.3991 | -8.3128e-10 | 1.4041
-1.4779e-11 | 1.8051 | -7.899e-10 | 1.8081
-9.5497e-12 | 2.2111 | -8.1945e-10 | 2.2131
-1.387e-11 | 2.6151 | -8.0877e-10 | 2.6172
-1.4325e-11 | 3.0212 | -8.2036e-10 | 3.0232
-9.0949e-12 | 3.4262 | -8.6106e-10 | 3.4282
-8.4128e-12 | 3.8302 | -8.2855e-10 | 3.8322
-1.3642e-11 | 4.2352 | -9.2541e-10 | 4.2362
-1.1596e-11 | 4.6403 | -8.1559e-10 | 4.6413
-1.1141e-11 | 5.0453 | -1.0195e-09 | 5.0463
-1.1141e-11 | 5.4493 | -7.8307e-10 | 5.4503
-1.4779e-11 | 5.8543 | -9.1586e-10 | 5.8543
-8.4128e-12 | 6.2614 | -9.2655e-10 | 6.2594
-9.5497e-12 | 6.6664 | -8.7152e-10 | 6.6644
-1.2506e-11 | 7.0714 | -8.2855e-10 | 7.0684
-1.3188e-11 | 7.4764 | -9.386e-10 | 7.4734
-9.3223e-12 | 7.8825 | -8.4856e-10 | 7.8774
-9.7771e-12 | 8.2865 | -8.5356e-10 | 8.2815
-1.8645e-11 | 8.6915 | -7.5693e-10 | 8.6865
-1.3188e-11 | 9.0955 | -8.3287e-10 | 9.0895
-7.0486e-12 | 9.5015 | -8.6698e-10 | 9.4955
-9.3223e-12 | 9.9066 | -9.3041e-10 | 9.8996
-1.2051e-11 | 10.3106 | -8.3514e-10 | 10.3046
-1.6371e-11 | 10.7156 | -8.1195e-10 | 10.7086
-8.1855e-12 | 11.1206 | -9.0404e-10 | 11.1136
-7.0486e-12 | 11.5257 | -8.1263e-10 | 11.5187
-1.296e-11 | 11.9317 | -1.0002e-09 | 11.9237
-8.4128e-12 | 12.3367 | -9.7657e-10 | 12.3297
-5.6843e-12 | 12.7427 | -9.0063e-10 | 12.7357
-1.7735e-11 | 13.1468 | -9.4519e-10 | 13.1408
-1.296e-11 | 13.5528 | -7.7762e-10 | 13.5458
-9.7771e-12 | 13.9588 | -9.511e-10 | 13.9508
-1.4552e-11 | 14.3638 | -8.1377e-10 | 14.3548
-1.728e-11 | 14.7678 | -9.3655e-10 | 14.7598
-1.2506e-11 | 15.1739 | -9.4792e-10 | 15.1649
-9.0949e-12 | 15.5799 | -9.1063e-10 | 15.5699
-9.5497e-12 | 15.9829 | -8.356e-10 | 15.9749
-1.4552e-11 | 16.3879 | -7.8876e-10 | 16.3789
-1.6826e-11 | 16.793 | -8.8266e-10 | 16.784
-7.9581e-12 | 17.198 | -8.3446e-10 | 17.188
-1.4325e-11 | 17.603 | -8.7766e-10 | 17.592
-1.1596e-11 | 18.007 | -8.1195e-10 | 17.996
-1.2051e-11 | 18.413 | -8.9153e-10 | 18.4011
-6.5938e-12 | 18.8191 | -7.8285e-10 | 18.8051
-1.8872e-11 | 19.2241 | -8.8153e-10 | 19.2101
-1.4552e-11 | 19.6291 | -8.7448e-10 | 19.6161
-8.1855e-12 | 20.0352 | -9.1268e-10 | 20.0211
-7.9581e-12 | 20.4402 | -9.1109e-10 | 20.4262
-1.2278e-11 | 20.8462 | -7.6943e-10 | 20.8302
-1.3188e-11 | 21.2512 | -8.9631e-10 | 21.2352
-7.9581e-12 | 21.6572 | -8.074e-10 | 21.6402
-9.0949e-12 | 22.0623 | -9.0722e-10 | 22.0453
-1.296e-11 | 22.4663 | -7.9785e-10 | 22.4493
-9.5497e-12 | 22.8723 | -9.0995e-10 | 22.8533
-6.1391e-12 | 23.2773 | -9.5156e-10 | 23.2593
-1.4552e-11 | 23.6834 | -8.4538e-10 | 23.6643
-1.2506e-11 | 24.0884 | -8.5174e-10 | 24.0684
-1.5007e-11 | 24.4944 | -9.2155e-10 | 24.4734
-8.1855e-12 | 24.8984 | -8.7448e-10 | 24.8794
-1.0687e-11 | 25.3035 | -8.954e-10 | 25.2834
-1.296e-11 | 25.7085 | -9.0586e-10 | 25.6875
-6.5938e-12 | 26.1135 | -8.4265e-10 | 26.0915
-7.0486e-12 | 26.5185 | -9.5315e-10 | 26.4955
-1.4325e-11 | 26.9235 | -9.1518e-10 | 26.9015
-1.1823e-11 | 27.3286 | -9.6884e-10 | 27.3056
-6.8212e-12 | 27.7336 | -8.5174e-10 | 27.7106
-9.3223e-12 | 28.1376 | -8.6993e-10 | 28.1146
-1.3188e-11 | 28.5426 | -8.1741e-10 | 28.5196
-1.3188e-11 | 28.9477 | -9.0313e-10 | 28.9246
-8.6402e-12 | 29.3527 | -8.2787e-10 | 29.3297
-7.7307e-12 | 29.7567 | -9.7293e-10 | 29.7347
-1.5007e-11 | 30.1617 | -7.5352e-10 | 30.1397
-1.3415e-11 | 30.5678 | -1.0839e-09 | 30.5448
-1.0459e-11 | 30.9718 | -8.1354e-10 | 30.9488
-1.8645e-11 | 31.3778 | -8.0695e-10 | 31.3538
-1.5689e-11 | 31.7838 | -8.3264e-10 | 31.7578
-7.276e-12 | 32.1888 | -9.8294e-10 | 32.1628
-6.8212e-12 | 32.5929 | -9.2973e-10 | 32.5699
-1.8872e-11 | 32.9989 | -9.7089e-10 | 32.9749
-1.2278e-11 | 33.4039 | -8.8266e-10 | 33.3789
-1.0914e-11 | 33.8099 | -7.9149e-10 | 33.7829
-8.6402e-12 | 34.215 | -8.7402e-10 | 34.188
-1.2051e-11 | 34.619 | -8.5311e-10 | 34.594
-1.3415e-11 | 35.024 | -8.5288e-10 | 34.998
-1.0914e-11 | 35.429 | -9.1813e-10 | 35.402
-1.0914e-11 | 35.8351 | -8.9153e-10 | 35.8081
-1.9782e-11 | 36.2401 | -8.788e-10 | 36.2121
-1.2506e-11 | 36.6451 | -9.195e-10 | 36.6171
-1.5916e-11 | 37.0501 | -8.5493e-10 | 37.0221
-9.0949e-12 | 37.4551 | -8.4765e-10 | 37.4271
-1.3188e-11 | 37.8602 | -8.6425e-10 | 37.8332
-9.7771e-12 | 38.2652 | -9.3905e-10 | 38.2372
-1.0004e-11 | 38.6702 | -7.9899e-10 | 38.6412
-1.5916e-11 | 39.0752 | -9.843e-10 | 39.0472
-1.1596e-11 | 39.4813 | -8.4128e-10 | 39.4503
-9.0949e-12 | 39.8863 | -9.3314e-10 | 39.8553
-8.1855e-12 | 40.2923 | -8.8517e-10 | 40.2593
-1.4779e-11 | 40.6983 | -8.4538e-10 | 40.6643
-1.2278e-11 | 41.1033 | -9.17e-10 | 41.0703
-1.0459e-11 | 41.5084 | -9.4838e-10 | 41.4754
-1.0004e-11 | 41.9134 | -9.4724e-10 | 41.8804
-1.2051e-11 | 42.3194 | -8.2514e-10 | 42.2844
-2.0691e-11 | 42.7244 | -7.6375e-10 | 42.6904
-9.0949e-12 | 43.1295 | -8.6857e-10 | 43.0955
-2.1828e-11 | 43.5345 | -8.6266e-10 | 43.5005
-1.5916e-11 | 43.9395 | -8.24e-10 | 43.9055
-9.5497e-12 | 44.3455 | -1.033e-09 | 44.3095
-8.4128e-12 | 44.7496 | -9.2496e-10 | 44.7136
-1.3415e-11 | 45.1546 | -8.763e-10 | 45.1186
-1.3188e-11 | 45.5586 | -8.1945e-10 | 45.5236
-1.4552e-11 | 45.9636 | -9.0836e-10 | 45.9266
-7.276e-12 | 46.3697 | -9.4406e-10 | 46.3306
-1.1596e-11 | 46.7747 | -8.3742e-10 | 46.7347
-1.1596e-11 | 47.1807 | -7.9331e-10 | 47.1397
-9.3223e-12 | 47.5867 | -1.0073e-09 | 47.5447
-1.2051e-11 | 47.9907 | -7.8285e-10 | 47.9487
-2.0691e-11 | 48.3958 | -8.4833e-10 | 48.3528
-1.1596e-11 | 48.8008 | -8.9585e-10 | 48.7588
-2.4556e-11 | 49.2058 | -8.3742e-10 | 49.1648
-8.4128e-12 | 49.6108 | -9.7634e-10 | 49.5698
-1.2051e-11 | 50.0159 | -8.8039e-10 | 49.9749
-1.3188e-11 | 50.4199 | -8.9472e-10 | 50.3789
-8.1855e-12 | 50.8239 | -8.24e-10 | 50.7839
-9.0949e-12 | 51.2299 | -9.4428e-10 | 51.1889
-1.5007e-11 | 51.636 | -1.0364e-09 | 51.593
-1.1823e-11 | 52.04 | -8.8244e-10 | 51.996
-8.1855e-12 | 52.445 | -9.3814e-10 | 52.401
-1.0914e-11 | 52.85 | -8.6834e-10 | 52.804
-1.2278e-11 | 53.2561 | -8.2491e-10 | 53.208
-1.1596e-11 | 53.6611 | -8.3924e-10 | 53.6121
-1.7053e-11 | 54.0661 | -8.8903e-10 | 54.0161
-1.3415e-11 | 54.4731 | -9.6611e-10 | 54.4211
-1.387e-11 | 54.8781 | -8.7198e-10 | 54.8251
-8.6402e-12 | 55.2842 | -9.6179e-10 | 55.2302
-1.0459e-11 | 55.6892 | -8.7448e-10 | 55.6342
-1.5916e-11 | 56.0932 | -1.0129e-09 | 56.0382
-1.1823e-11 | 56.4992 | -8.3264e-10 | 56.4432
-8.1855e-12 | 56.9053 | -8.0581e-10 | 56.8473
-1.4552e-11 | 57.3103 | -1.0309e-09 | 57.2523
-1.5007e-11 | 57.7153 | -8.5083e-10 | 57.6573
-9.3223e-12 | 58.1203 | -8.1741e-10 | 58.0623
-1.3415e-11 | 58.5264 | -7.6398e-10 | 58.4663
-1.3642e-11 | 58.9324 | -8.7743e-10 | 58.8724
-1.5461e-11 | 59.3384 | -8.7152e-10 | 59.2764
-9.3223e-12 | 59.7444 | -9.8908e-10 | 59.6814
-1.2051e-11 | 60.1504 | -9.3473e-10 | 60.0854
-1.2506e-11 | 60.5575 | -7.9785e-10 | 60.4915
-8.4128e-12 | 60.9625 | -8.7402e-10 | 60.8965
-1.6371e-11 | 61.3675 | -1.0368e-09 | 61.2995
-1.4779e-11 | 61.7725 | -1.0473e-09 | 61.7035
-1.0687e-11 | 62.1776 | -8.6425e-10 | 62.1086
-2.0236e-11 | 62.5816 | -1.0332e-09 | 62.5136
-1.0687e-11 | 62.9876 | -7.8876e-10 | 62.9166
-1.3642e-11 | 63.3926 | -1.0332e-09 | 63.3206
-1.1141e-11 | 63.7976 | -8.7539e-10 | 63.7257
-1.2051e-11 | 64.2027 | -8.6129e-10 | 64.1297
-1.5689e-11 | 64.6077 | -7.6693e-10 | 64.5347
-2.2055e-11 | 65.0137 | -7.567e-10 | 64.9387
-9.5497e-12 | 65.4177 | -9.2746e-10 | 65.3437
-8.4128e-12 | 65.8228 | -7.9149e-10 | 65.7488
-1.3642e-11 | 66.2278 | -9.0722e-10 | 66.1528
-2.2737e-11 | 66.6328 | -8.9244e-10 | 66.5578
-1.1596e-11 | 67.0388 | -8.4651e-10 | 66.9628
-1.6371e-11 | 67.4449 | -8.813e-10 | 67.3679
-1.5916e-11 | 67.8509 | -8.1832e-10 | 67.7729
-1.1596e-11 | 68.2559 | -8.3446e-10 | 68.1769
-1.0459e-11 | 68.6619 | -9.8817e-10 | 68.5819
-1.1141e-11 | 69.0679 | -9.3019e-10 | 68.9869
-1.3642e-11 | 69.472 | -1.0059e-09 | 69.392
-9.0949e-12 | 69.877 | -8.4424e-10 | 69.797
-1.1141e-11 | 70.282 | -8.2605e-10 | 70.202
-1.2278e-11 | 70.687 | -9.2837e-10 | 70.606
-1.296e-11 | 71.0931 | -7.6557e-10 | 71.0091
-1.3642e-11 | 71.4971 | -9.1427e-10 | 71.4141
-1.0004e-11 | 71.9011 | -9.4042e-10 | 71.8191
-1.4552e-11 | 72.3061 | -9.2859e-10 | 72.2241
-1.296e-11 | 72.7132 | -8.522e-10 | 72.6282
-1.387e-11 | 73.1182 | -8.738e-10 | 73.0332
-9.7771e-12 | 73.5242 | -8.3219e-10 | 73.4382
-1.6371e-11 | 73.9282 | -8.0695e-10 | 73.8422
-2.0236e-11 | 74.3342 | -8.215e-10 | 74.2472
-9.5497e-12 | 74.7393 | -8.9676e-10 | 74.6513
-1.1141e-11 | 75.1443 | -7.9399e-10 | 75.0553
-1.2506e-11 | 75.5493 | -7.5556e-10 | 75.4613
-1.3642e-11 | 75.9553 | -8.863e-10 | 75.8663
-1.0004e-11 | 76.3594 | -8.5947e-10 | 76.2714
-1.2051e-11 | 76.7644 | -8.1559e-10 | 76.6764
-1.4325e-11 | 77.1694 | -9.0108e-10 | 77.0794
-1.1823e-11 | 77.5744 | -1.0036e-09 | 77.4834
-1.0004e-11 | 77.9795 | -8.672e-10 | 77.8885
-1.2278e-11 | 78.3855 | -8.4788e-10 | 78.2925
-1.2506e-11 | 78.7905 | -8.5925e-10 | 78.6995
-9.5497e-12 | 79.1975 | -8.2969e-10 | 79.1035
-9.5497e-12 | 79.6016 | -8.0991e-10 | 79.5075
-1.3642e-11 | 80.0076 | -8.6584e-10 | 79.9116
-1.5689e-11 | 80.4116 | -8.1695e-10 | 80.3156
-1.1823e-11 | 80.8186 | -8.8266e-10 | 80.7216
-1.1823e-11 | 81.2237 | -7.7284e-10 | 81.1246
-1.4779e-11 | 81.6287 | -9.3883e-10 | 81.5297
-1.1823e-11 | 82.0327 | -8.2582e-10 | 81.9347
-1.1141e-11 | 82.4387 | -7.9262e-10 | 82.3397
-1.3188e-11 | 82.8447 | -7.7898e-10 | 82.7437
-1.2506e-11 | 83.2498 | -9.7407e-10 | 83.1488

## 690uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.5997e-11 | 4.231e-12 | -1.2262e-09 | 9.5664e-11
GEMs On |  | GEMs On
-1.4325e-11 | 0.31002 | -1.3047e-09 | 0.31502
-1.1823e-11 | 0.99506 | -1.1498e-09 | 0.99806
-1.6144e-11 | 1.4031 | -1.2112e-09 | 1.4031
-1.1141e-11 | 1.8091 | -1.1951e-09 | 1.8071
-1.0004e-11 | 2.2141 | -1.2626e-09 | 2.2111
-1.2278e-11 | 2.6202 | -1.3595e-09 | 2.6161
-1.2506e-11 | 3.0252 | -1.1685e-09 | 3.0202
-1.0687e-11 | 3.4312 | -1.3067e-09 | 3.4252
-1.3642e-11 | 3.8362 | -1.3895e-09 | 3.8292
-1.0687e-11 | 4.2422 | -1.0687e-09 | 4.2332
-1.728e-11 | 4.6483 | -1.2913e-09 | 4.6383
-2.0691e-11 | 5.0533 | -1.1548e-09 | 5.0413
-1.7053e-11 | 5.4593 | -1.1985e-09 | 5.4463
-1.2506e-11 | 5.8643 | -1.0536e-09 | 5.8513
-1.2506e-11 | 6.2704 | -1.1582e-09 | 6.2554
-1.2051e-11 | 6.6774 | -1.1441e-09 | 6.6604
-1.5916e-11 | 7.0824 | -1.1935e-09 | 7.0664
-9.3223e-12 | 7.4894 | -1.2076e-09 | 7.4704
-9.3223e-12 | 7.8945 | -1.3174e-09 | 7.8745
-1.7053e-11 | 8.2995 | -1.3197e-09 | 8.2795
-1.2278e-11 | 8.7035 | -1.1362e-09 | 8.6845
-1.3415e-11 | 9.1085 | -1.3003e-09 | 9.0895
-9.7771e-12 | 9.5135 | -1.1444e-09 | 9.4935
-1.296e-11 | 9.9196 | -1.1621e-09 | 9.8986
-1.296e-11 | 10.3256 | -1.2838e-09 | 10.3016
-2.3874e-11 | 10.7306 | -1.261e-09 | 10.7056
-1.4325e-11 | 11.1366 | -1.4263e-09 | 11.1106
-8.1855e-12 | 11.5397 | -1.2478e-09 | 11.5157
-1.387e-11 | 11.9447 | -1.1075e-09 | 11.9197
-1.4779e-11 | 12.3497 | -1.299e-09 | 12.3237
-1.0687e-11 | 12.7547 | -1.1337e-09 | 12.7277
-1.0687e-11 | 13.1587 | -1.2749e-09 | 13.1308
-1.2051e-11 | 13.5648 | -1.1178e-09 | 13.5348
-1.296e-11 | 13.9698 | -1.0727e-09 | 13.9398
-1.2051e-11 | 14.3758 | -1.0778e-09 | 14.3448
-9.0949e-12 | 14.7808 | -1.3422e-09 | 14.7488
-1.387e-11 | 15.1859 | -1.1928e-09 | 15.1519
-1.2051e-11 | 15.5919 | -1.2053e-09 | 15.5569
-1.296e-11 | 16.0039 | -1.0684e-09 | 15.9609
-2.6375e-11 | 16.4099 | -1.2976e-09 | 16.3649
-1.4552e-11 | 16.816 | -1.103e-09 | 16.769
-1.728e-11 | 17.222 | -1.2237e-09 | 17.173
-1.3188e-11 | 17.626 | -1.1864e-09 | 17.578
-1.1596e-11 | 18.031 | -1.0443e-09 | 17.981
-1.296e-11 | 18.4361 | -1.0632e-09 | 18.385
-1.3188e-11 | 18.8411 | -1.2797e-09 | 18.7901
-1.387e-11 | 19.2471 | -1.2644e-09 | 19.1941
-1.5007e-11 | 19.6531 | -1.2251e-09 | 19.5981
-1.8645e-11 | 20.0582 | -1.3943e-09 | 20.0021
-1.4325e-11 | 20.4652 | -1.2624e-09 | 20.4062
-1.2506e-11 | 20.8702 | -1.1364e-09 | 20.8102
-1.3415e-11 | 21.2762 | -1.1378e-09 | 21.2142
-1.5007e-11 | 21.6812 | -1.3856e-09 | 21.6182
-1.1596e-11 | 22.0863 | -1.3024e-09 | 22.0233
-1.3188e-11 | 22.4913 | -1.2903e-09 | 22.4283
-1.5916e-11 | 22.8973 | -1.1157e-09 | 22.8333
-2.7285e-11 | 23.3023 | -1.1566e-09 | 23.2383
-1.2051e-11 | 23.7084 | -1.1996e-09 | 23.6413
-1.3415e-11 | 24.1134 | -1.2576e-09 | 24.0454
-2.5238e-11 | 24.5184 | -1.2781e-09 | 24.4504
-1.4552e-11 | 24.9244 | -1.2246e-09 | 24.8554
-1.4779e-11 | 25.3295 | -1.2353e-09 | 25.2604
-2.4102e-11 | 25.7345 | -1.1294e-09 | 25.6635
-1.296e-11 | 26.1395 | -1.1951e-09 | 26.0685
-1.4325e-11 | 26.5445 | -1.3338e-09 | 26.4735
-1.3188e-11 | 26.9495 | -1.3545e-09 | 26.8775
-1.6144e-11 | 27.3556 | -1.2367e-09 | 27.2806
-1.4325e-11 | 27.7606 | -1.1105e-09 | 27.6856
-1.296e-11 | 28.1656 | -1.2149e-09 | 28.0906
-1.387e-11 | 28.5706 | -1.1594e-09 | 28.4946
-1.4325e-11 | 28.9747 | -1.1528e-09 | 28.8997
-1.2278e-11 | 29.3817 | -1.1689e-09 | 29.3037
-2.5693e-11 | 29.7867 | -1.2492e-09 | 29.7077
-1.3188e-11 | 30.1927 | -1.1466e-09 | 30.1107
-1.296e-11 | 30.5987 | -1.1626e-09 | 30.5158
-1.387e-11 | 31.0038 | -1.1798e-09 | 30.9198
-1.5461e-11 | 31.4088 | -1.1632e-09 | 31.3238
-1.4325e-11 | 31.8148 | -1.3292e-09 | 31.7288
-1.3642e-11 | 32.2198 | -1.357e-09 | 32.1328
-1.1823e-11 | 32.6249 | -1.2503e-09 | 32.5369
-2.1373e-11 | 33.0299 | -1.2601e-09 | 32.9409
-1.5689e-11 | 33.4349 | -9.9135e-10 | 33.3449
-1.4779e-11 | 33.8399 | -1.3429e-09 | 33.7499
-1.2278e-11 | 34.246 | -1.3461e-09 | 34.154
-1.4325e-11 | 34.651 | -1.1592e-09 | 34.558
-1.3642e-11 | 35.058 | -1.2651e-09 | 34.963
-1.2506e-11 | 35.463 | -1.2035e-09 | 35.367
-1.4325e-11 | 35.867 | -1.2208e-09 | 35.7721
-1.4779e-11 | 36.2721 | -1.1378e-09 | 36.1751
-1.6371e-11 | 36.6771 | -1.3395e-09 | 36.5791
-1.4552e-11 | 37.0811 | -1.4104e-09 | 36.9841
-1.3415e-11 | 37.4871 | -1.0866e-09 | 37.3881
-1.3415e-11 | 37.8922 | -1.151e-09 | 37.7922
-1.3642e-11 | 38.2972 | -1.4393e-09 | 38.1952
-1.5461e-11 | 38.7012 | -1.3395e-09 | 38.6002
-1.9327e-11 | 39.1062 | -1.269e-09 | 39.0042
-1.5916e-11 | 39.5113 | -1.1262e-09 | 39.4083
-1.5007e-11 | 39.9163 | -1.2158e-09 | 39.8123
-1.3415e-11 | 40.3213 | -1.153e-09 | 40.2163
-1.387e-11 | 40.7263 | -1.2403e-09 | 40.6203
-1.5007e-11 | 41.1324 | -1.3222e-09 | 41.0254
-1.6144e-11 | 41.5374 | -1.1903e-09 | 41.4294
-1.4779e-11 | 41.9414 | -1.1403e-09 | 41.8324
-1.4779e-11 | 42.3474 | -1.261e-09 | 42.2374
-1.8417e-11 | 42.7525 | -1.1037e-09 | 42.6414
-1.3188e-11 | 43.1585 | -1.3702e-09 | 43.0465
-1.387e-11 | 43.5635 | -1.3702e-09 | 43.4495
-1.4325e-11 | 43.9685 | -1.3579e-09 | 43.8535
-1.7508e-11 | 44.3735 | -1.2874e-09 | 44.2565
-2.251e-11 | 44.7806 | -1.1412e-09 | 44.6606
-1.387e-11 | 45.1856 | -1.2765e-09 | 45.0646
-1.3642e-11 | 45.5906 | -1.2785e-09 | 45.4686
-1.5689e-11 | 45.9956 | -1.2997e-09 | 45.8736
-1.6144e-11 | 46.4006 | -1.078e-09 | 46.2777
-1.5007e-11 | 46.8057 | -1.2046e-09 | 46.6817
-2.3647e-11 | 47.2107 | -1.2344e-09 | 47.0867
-2.6375e-11 | 47.6157 | -1.3852e-09 | 47.4917
-1.8645e-11 | 48.0207 | -1.1426e-09 | 47.8957
-1.8417e-11 | 48.4258 | -1.1817e-09 | 48.3008
-1.5689e-11 | 48.8298 | -1.1819e-09 | 48.7048
-1.3642e-11 | 49.2358 | -1.3238e-09 | 49.1088
-1.6371e-11 | 49.6408 | -1.2567e-09 | 49.5128
-1.819e-11 | 50.0459 | -1.2915e-09 | 49.9179
-1.4325e-11 | 50.4509 | -1.2576e-09 | 50.3229
-1.4779e-11 | 50.8559 | -1.1996e-09 | 50.7269
-1.6826e-11 | 51.2599 | -1.1819e-09 | 51.1309
-2.4102e-11 | 51.6669 | -1.1255e-09 | 51.5359
-1.5689e-11 | 52.073 | -1.1262e-09 | 51.941
-1.7508e-11 | 52.478 | -1.0969e-09 | 52.346
-2.2055e-11 | 52.882 | -1.3479e-09 | 52.75
-1.8417e-11 | 53.2871 | -1.2535e-09 | 53.154
-1.3188e-11 | 53.6921 | -1.3904e-09 | 53.5571
-1.296e-11 | 54.0971 | -1.098e-09 | 53.9611
-1.5916e-11 | 54.5021 | -1.1632e-09 | 54.3661
-2.4784e-11 | 54.9081 | -1.093e-09 | 54.7711
-1.3642e-11 | 55.3132 | -1.2576e-09 | 55.1752
-1.4325e-11 | 55.7182 | -1.1714e-09 | 55.5802
-1.9327e-11 | 56.1242 | -1.2224e-09 | 55.9842
-1.5689e-11 | 56.5292 | -1.2074e-09 | 56.3882
-1.3642e-11 | 56.9353 | -1.4224e-09 | 56.7923
-1.6371e-11 | 57.3393 | -1.2415e-09 | 57.1983
-1.8417e-11 | 57.7443 | -1.2276e-09 | 57.6023
-1.7735e-11 | 58.1493 | -1.1887e-09 | 58.0063
-1.3642e-11 | 58.5534 | -1.0996e-09 | 58.4113
-1.4779e-11 | 58.9594 | -1.2956e-09 | 58.8154
-2.4102e-11 | 59.3644 | -1.2033e-09 | 59.2194
-1.7735e-11 | 59.7694 | -1.1953e-09 | 59.6234
-1.387e-11 | 60.1754 | -1.3799e-09 | 60.0294
-1.3642e-11 | 60.5805 | -1.1298e-09 | 60.4315
-1.728e-11 | 60.9855 | -1.1996e-09 | 60.8345
-1.9327e-11 | 61.3905 | -1.2478e-09 | 61.2385
-1.3642e-11 | 61.7955 | -1.1482e-09 | 61.6425
-1.6144e-11 | 62.2016 | -1.0298e-09 | 62.0476
-1.8645e-11 | 62.6076 | -1.5111e-09 | 62.4526
-1.5007e-11 | 63.0136 | -1.0573e-09 | 62.8566
-1.2506e-11 | 63.4186 | -1.1682e-09 | 63.2616
-1.8417e-11 | 63.8237 | -1.221e-09 | 63.6646
-1.819e-11 | 64.2297 | -1.3133e-09 | 64.0677
-1.296e-11 | 64.6357 | -1.1428e-09 | 64.4717
-1.5461e-11 | 65.0417 | -1.2853e-09 | 64.8757
-2.3192e-11 | 65.4477 | -1.1876e-09 | 65.2797
-2.1146e-11 | 65.8528 | -1.1689e-09 | 65.6848
-1.4552e-11 | 66.2578 | -1.2726e-09 | 66.0888
-1.7508e-11 | 66.6638 | -1.3954e-09 | 66.4938
-2.2283e-11 | 67.0688 | -9.8635e-10 | 66.8988
-2.5693e-11 | 67.4749 | -1.3822e-09 | 67.3029
-1.2051e-11 | 67.8799 | -1.1767e-09 | 67.7069
-1.9782e-11 | 68.2849 | -1.1814e-09 | 68.1119
-1.7735e-11 | 68.6909 | -1.2658e-09 | 68.5169
-1.5461e-11 | 69.097 | -1.2003e-09 | 68.9209
-1.5689e-11 | 69.503 | -1.3777e-09 | 69.325
-1.7735e-11 | 69.908 | -1.3263e-09 | 69.73
-1.7053e-11 | 70.312 | -1.2685e-09 | 70.134
-1.7735e-11 | 70.718 | -1.2942e-09 | 70.537
-1.3415e-11 | 71.1241 | -1.2446e-09 | 70.9411
-1.728e-11 | 71.5291 | -1.3404e-09 | 71.3451
-1.728e-11 | 71.9341 | -1.3831e-09 | 71.7491
-1.4779e-11 | 72.3401 | -1.1794e-09 | 72.1531
-1.728e-11 | 72.7452 | -1.251e-09 | 72.5592
-1.9554e-11 | 73.1502 | -1.3356e-09 | 72.9622
-1.5461e-11 | 73.5572 | -1.2112e-09 | 73.3672
-1.6826e-11 | 73.9642 | -1.2412e-09 | 73.7712
-1.7508e-11 | 74.3692 | -1.2533e-09 | 74.1762
-1.728e-11 | 74.7743 | -1.306e-09 | 74.5803
-1.5689e-11 | 75.1793 | -1.2433e-09 | 74.9833
-1.4779e-11 | 75.5843 | -1.3301e-09 | 75.3883
-1.7735e-11 | 75.9904 | -1.2435e-09 | 75.7923
-3.6152e-11 | 76.3964 | -1.2412e-09 | 76.1964
-1.4552e-11 | 76.8014 | -1.2969e-09 | 76.6014
-1.5689e-11 | 77.2064 | -1.2392e-09 | 77.0044
-1.8417e-11 | 77.6124 | -1.2044e-09 | 77.4094
-1.6826e-11 | 78.0175 | -1.0762e-09 | 77.8135
-1.7508e-11 | 78.4235 | -1.3545e-09 | 78.2175
-1.6144e-11 | 78.8305 | -1.2019e-09 | 78.6225
-2.251e-11 | 79.2335 | -1.2312e-09 | 79.0275
-1.7508e-11 | 79.6406 | -1.1942e-09 | 79.4315
-1.5689e-11 | 80.0446 | -1.166e-09 | 79.8356
-4.0245e-11 | 80.4516 | -1.3204e-09 | 80.2406
-2.16e-11 | 80.8556 | -1.2631e-09 | 80.6456
-1.6144e-11 | 81.2617 | -1.0273e-09 | 81.0496
-2.3192e-11 | 81.6677 | -1.2301e-09 | 81.4537
-2.0918e-11 | 82.0737 | -1.0657e-09 | 81.8587
-1.6826e-11 | 82.4787 | -1.113e-09 | 82.2627
-1.9554e-11 | 82.8837 | -1.3274e-09 | 82.6667
-1.5916e-11 | 83.2888 | -1.2537e-09 | 83.0717
-1.9327e-11 | 83.6948 | -1.2062e-09 | 83.4778

## 700uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-2.4088e-11 | 7.6611e-12 | -1.7561e-09 | 1.3439e-10
GEMs On |  | GEMs On
-9.5497e-12 | 0.30902 | -1.6812e-09 | 0.53003
-1.6598e-11 | 0.99406 | -1.882e-09 | 1.3021
-1.5461e-11 | 1.4001 | -1.6466e-09 | 1.7071
-1.4097e-11 | 1.8051 | -1.7469e-09 | 2.1111
-1.4552e-11 | 2.2101 | -1.8433e-09 | 2.5161
-1.1369e-11 | 2.6161 | -1.7942e-09 | 2.9202
-1.5007e-11 | 3.0212 | -1.8249e-09 | 3.3242
-1.3415e-11 | 3.4272 | -1.8269e-09 | 3.7282
-1.387e-11 | 3.8322 | -1.6269e-09 | 4.1332
-1.2733e-11 | 4.2382 | -1.686e-09 | 4.5363
-1.3415e-11 | 4.6433 | -1.7164e-09 | 4.9403
-1.4779e-11 | 5.0473 | -1.502e-09 | 5.3443
-1.3642e-11 | 5.4533 | -1.56e-09 | 5.7493
-1.387e-11 | 5.8573 | -1.8742e-09 | 6.1554
-1.4779e-11 | 6.2634 | -1.8529e-09 | 6.5594
-1.2506e-11 | 6.6674 | -1.9195e-09 | 6.9644
-1.3642e-11 | 7.0734 | -1.7965e-09 | 7.3694
-1.4097e-11 | 7.4774 | -1.4729e-09 | 7.7734
6.3665e-12 | 7.8835 | -1.9238e-09 | 8.1785
-1.2733e-11 | 8.2875 | -1.7112e-09 | 8.5825
-1.2506e-11 | 8.6935 | -1.7278e-09 | 8.9875
-1.4097e-11 | 9.0995 | -1.6885e-09 | 9.3915
-1.3188e-11 | 9.5055 | -1.7008e-09 | 9.7956
-4.4793e-11 | 9.9116 | -1.6764e-09 | 10.1996
-1.9099e-11 | 10.3166 | -1.6523e-09 | 10.6046
-1.387e-11 | 10.7216 | -1.8088e-09 | 11.0086
-1.5234e-11 | 11.1276 | -1.6089e-09 | 11.4126
-1.8645e-11 | 11.5327 | -1.7064e-09 | 11.8177
-3.6607e-11 | 11.9377 | -1.5432e-09 | 12.2217
-2.3647e-11 | 12.3417 | -1.962e-09 | 12.6257
-1.8645e-11 | 12.7477 | -1.5882e-09 | 13.0297
-1.4097e-11 | 13.1517 | -1.6689e-09 | 13.4348
-1.4097e-11 | 13.5578 | -1.6375e-09 | 13.8398
-1.4097e-11 | 13.9628 | -1.7674e-09 | 14.2428
-1.3642e-11 | 14.3678 | -1.8815e-09 | 14.6458
-1.5007e-11 | 14.7728 | -1.9484e-09 | 15.0509
-2.0009e-11 | 15.1789 | -1.8392e-09 | 15.4549
-1.3642e-11 | 15.5839 | -2.0182e-09 | 15.8599
-1.5007e-11 | 15.9899 | -1.6478e-09 | 16.2649
-1.6598e-11 | 16.3949 | -1.7199e-09 | 16.67
-1.2278e-11 | 16.801 | -1.7408e-09 | 17.074
-1.4097e-11 | 17.205 | -1.7594e-09 | 17.478
-1.387e-11 | 17.61 | -1.774e-09 | 17.884
-1.5234e-11 | 18.015 | -1.726e-09 | 18.288
-1.4779e-11 | 18.419 | -1.8188e-09 | 18.6931
-1.0914e-11 | 18.8241 | -1.7001e-09 | 19.0961
-1.4097e-11 | 19.2291 | -1.834e-09 | 19.5001
-1.5234e-11 | 19.6351 | -1.6369e-09 | 19.9051
-1.5461e-11 | 20.0402 | -1.6444e-09 | 20.3092
-1.5234e-11 | 20.4462 | -1.897e-09 | 20.7132
-2.7512e-11 | 20.8512 | -1.8749e-09 | 21.1182
-1.6598e-11 | 21.2552 | -1.5232e-09 | 21.5222
-1.3642e-11 | 21.6602 | -1.834e-09 | 21.9262
-1.6144e-11 | 22.0663 | -1.9165e-09 | 22.3303
-1.5234e-11 | 22.4723 | -1.9406e-09 | 22.7353
-1.5234e-11 | 22.8763 | -1.8201e-09 | 23.1393
-2.7285e-11 | 23.2813 | -1.5889e-09 | 23.5444
-1.6371e-11 | 23.6864 | -1.6223e-09 | 23.9494
-2.0464e-11 | 24.0914 | -1.5827e-09 | 24.3524
-1.5234e-11 | 24.4964 | -1.7599e-09 | 24.7574
-2.0236e-11 | 24.9014 | -1.6857e-09 | 25.1624
-1.6371e-11 | 25.3074 | -1.7544e-09 | 25.5685
-1.7053e-11 | 25.7135 | -1.7712e-09 | 25.9725
-1.8417e-11 | 26.1175 | -1.8388e-09 | 26.3765
-1.728e-11 | 26.5225 | -1.864e-09 | 26.7815
-1.7053e-11 | 26.9285 | -1.8729e-09 | 27.1866
-1.5916e-11 | 27.3336 | -1.5993e-09 | 27.5916
-1.4552e-11 | 27.7396 | -1.733e-09 | 27.9956
-1.6371e-11 | 28.1456 | -1.5755e-09 | 28.4006
-1.5916e-11 | 28.5506 | -1.997e-09 | 28.8046
-1.4097e-11 | 28.9547 | -1.8508e-09 | 29.2087
-1.387e-11 | 29.3597 | -1.7953e-09 | 29.6127
-1.387e-11 | 29.7647 | -1.733e-09 | 30.0187
-1.9782e-11 | 30.1697 | -1.5405e-09 | 30.4227
-1.7053e-11 | 30.5757 | -2.0843e-09 | 30.8268
-1.6371e-11 | 30.9798 | -1.8542e-09 | 31.2308
-1.6144e-11 | 31.3848 | -1.7371e-09 | 31.6348
-1.6144e-11 | 31.7898 | -1.6087e-09 | 32.0388
-1.4097e-11 | 32.1948 | -1.625e-09 | 32.4429
-1.6598e-11 | 32.5989 | -1.7908e-09 | 32.8469
-5.8662e-11 | 33.0039 | -1.7212e-09 | 33.2509
-1.5916e-11 | 33.4089 | -1.9518e-09 | 33.6549
-1.9099e-11 | 33.8139 | -1.6985e-09 | 34.059
-1.7508e-11 | 34.221 | -1.7524e-09 | 34.463
-2.2283e-11 | 34.627 | -1.9086e-09 | 34.868
-1.5461e-11 | 35.031 | -1.8586e-09 | 35.273
-1.5234e-11 | 35.437 | -1.8479e-09 | 35.678
-1.9327e-11 | 35.843 | -1.6769e-09 | 36.0821
-1.4097e-11 | 36.2491 | -1.5398e-09 | 36.4851
-1.6598e-11 | 36.6551 | -1.4852e-09 | 36.8901
-3.979e-11 | 37.0611 | -1.5327e-09 | 37.2941
-1.7963e-11 | 37.4671 | -1.5493e-09 | 37.6972
-1.6598e-11 | 37.8732 | -1.7433e-09 | 38.1012
-2.0009e-11 | 38.2782 | -1.819e-09 | 38.5062
-1.9327e-11 | 38.6832 | -1.7283e-09 | 38.9112
-1.7053e-11 | 39.0882 | -1.7778e-09 | 39.3162
-1.5916e-11 | 39.4943 | -1.8158e-09 | 39.7223
-1.728e-11 | 39.9013 | -2.0407e-09 | 40.1263
-2.774e-11 | 40.3073 | -1.8088e-09 | 40.5293
-1.6598e-11 | 40.7123 | -1.7833e-09 | 40.9323
-1.7963e-11 | 41.1174 | -2.0314e-09 | 41.3354
-1.7508e-11 | 41.5224 | -1.7135e-09 | 41.7404
-2.4784e-11 | 41.9274 | -1.965e-09 | 42.1464
-1.387e-11 | 42.3324 | -1.7706e-09 | 42.5504
-1.9327e-11 | 42.7384 | -1.8451e-09 | 42.9555
-2.1828e-11 | 43.1435 | -1.6817e-09 | 43.3595
-1.9782e-11 | 43.5495 | -1.6039e-09 | 43.7635
-1.6144e-11 | 43.9545 | -1.9586e-09 | 44.1675
-1.8417e-11 | 44.3605 | -1.6109e-09 | 44.5716
-1.6371e-11 | 44.7656 | -1.9547e-09 | 44.9756
-1.8645e-11 | 45.1706 | -1.6587e-09 | 45.3796
-1.4097e-11 | 45.5756 | -1.8485e-09 | 45.7846
-1.9327e-11 | 45.9816 | -1.6639e-09 | 46.1896
-2.0464e-11 | 46.3877 | -1.8292e-09 | 46.5927
-1.6598e-11 | 46.7937 | -1.6635e-09 | 46.9967
-2.16e-11 | 47.1987 | -1.849e-09 | 47.4007
-1.9099e-11 | 47.6037 | -1.6698e-09 | 47.8047
-2.0918e-11 | 48.0087 | -1.6548e-09 | 48.2078
-1.7508e-11 | 48.4148 | -1.8135e-09 | 48.6118
-1.7735e-11 | 48.8208 | -1.3088e-09 | 49.0158
-1.7053e-11 | 49.2258 | -1.8608e-09 | 49.4198
-1.8872e-11 | 49.6318 | -1.7621e-09 | 49.8248
-1.4779e-11 | 50.0369 | -1.6253e-09 | 50.2279
-2.0236e-11 | 50.4429 | -1.6123e-09 | 50.6319
-1.8645e-11 | 50.8489 | -1.4663e-09 | 51.0379
-2.2283e-11 | 51.2539 | -1.6228e-09 | 51.4429
-1.5916e-11 | 51.6589 | -1.8574e-09 | 51.848
-3.0013e-11 | 52.066 | -1.497e-09 | 52.252
-1.6144e-11 | 52.47 | -1.6073e-09 | 52.657
-2.1146e-11 | 52.875 | -1.8028e-09 | 53.061
-3.0695e-11 | 53.2801 | -1.723e-09 | 53.4651
-3.0695e-11 | 53.6861 | -1.8406e-09 | 53.8681
-2.0918e-11 | 54.0921 | -1.8474e-09 | 54.2721
-1.8417e-11 | 54.4961 | -1.7844e-09 | 54.6761
-3.7062e-11 | 54.9021 | -1.9086e-09 | 55.0821
-2.0009e-11 | 55.3092 | -1.8417e-09 | 55.4872
-1.8872e-11 | 55.7142 | -1.7926e-09 | 55.8912
-1.5461e-11 | 56.1192 | -1.7551e-09 | 56.2952
-1.7735e-11 | 56.5242 | -1.7162e-09 | 56.6992
-1.8645e-11 | 56.9293 | -1.61e-09 | 57.1033
-3.7744e-11 | 57.3333 | -1.5013e-09 | 57.5063
-2.16e-11 | 57.7383 | -1.7792e-09 | 57.9103
-2.251e-11 | 58.1433 | -1.8936e-09 | 58.3143
-3.3424e-11 | 58.5493 | -1.9497e-09 | 58.7174
-2.1146e-11 | 58.9544 | -1.9165e-09 | 59.1234
-1.8645e-11 | 59.3614 | -1.8417e-09 | 59.5284
-1.8872e-11 | 59.7674 | -1.8147e-09 | 59.9334
-2.2283e-11 | 60.1724 | -1.9422e-09 | 60.3374
-1.8417e-11 | 60.5785 | -1.7503e-09 | 60.7415
-1.7508e-11 | 60.9845 | -1.6955e-09 | 61.1465
-2.4102e-11 | 61.3895 | -1.9102e-09 | 61.5505
-2.0464e-11 | 61.7965 | -1.9131e-09 | 61.9555
-1.8872e-11 | 62.2026 | -1.794e-09 | 62.3606
-2.2737e-11 | 62.6076 | -1.6846e-09 | 62.7646
-3.0013e-11 | 63.0136 | -1.7255e-09 | 63.1686
-1.9327e-11 | 63.4196 | -1.9495e-09 | 63.5736
-1.7963e-11 | 63.8257 | -1.7333e-09 | 63.9777
-2.1146e-11 | 64.2307 | -1.8433e-09 | 64.3827
-2.3874e-11 | 64.6367 | -1.5741e-09 | 64.7857
-2.9104e-11 | 65.0427 | -1.8636e-09 | 65.1917
-2.0918e-11 | 65.4467 | -1.6807e-09 | 65.5957
-2.0464e-11 | 65.8518 | -1.8865e-09 | 66.0008
-3.979e-11 | 66.2578 | -1.8313e-09 | 66.4048
-1.9099e-11 | 66.6628 | -1.4907e-09 | 66.8088
-2.3874e-11 | 67.0688 | -1.7151e-09 | 67.2128
-2.251e-11 | 67.4739 | -1.5662e-09 | 67.6169
-2.1828e-11 | 67.8799 | -1.6837e-09 | 68.0209
-1.9099e-11 | 68.2849 | -1.4973e-09 | 68.4239
-2.1146e-11 | 68.6899 | -2.0684e-09 | 68.8289
-2.0009e-11 | 69.0949 | -1.8069e-09 | 69.234
-1.7735e-11 | 69.5 | -1.832e-09 | 69.64
-4.2291e-11 | 69.905 | -1.7824e-09 | 70.044
-2.1828e-11 | 70.311 | -1.7683e-09 | 70.449
-2.0236e-11 | 70.716 | -1.8238e-09 | 70.8541
-1.8417e-11 | 71.1221 | -1.8856e-09 | 71.2581
-2.0009e-11 | 71.5271 | -1.8974e-09 | 71.6621
-2.2737e-11 | 71.9331 | -1.7076e-09 | 72.0661
-1.8872e-11 | 72.3401 | -1.8472e-09 | 72.4711
-2.0236e-11 | 72.7452 | -1.6205e-09 | 72.8752
-3.206e-11 | 73.1492 | -1.6346e-09 | 73.2812
-2.0009e-11 | 73.5552 | -1.8085e-09 | 73.6852
-1.728e-11 | 73.9612 | -1.8906e-09 | 74.0892
-1.9782e-11 | 74.3672 | -1.9365e-09 | 74.4933
-2.4784e-11 | 74.7723 | -1.8995e-09 | 74.8973
-1.9327e-11 | 75.1783 | -1.9938e-09 | 75.3013
-2.0236e-11 | 75.5833 | -1.4886e-09 | 75.7053
-2.0918e-11 | 75.9883 | -1.6209e-09 | 76.1094
-2.0236e-11 | 76.3934 | -1.638e-09 | 76.5144
-1.8417e-11 | 76.7994 | -1.8254e-09 | 76.9174
-1.8645e-11 | 77.2054 | -1.8633e-09 | 77.3214
-2.5238e-11 | 77.6104 | -1.8435e-09 | 77.7254
-2.3647e-11 | 78.0145 | -1.6582e-09 | 78.1295
-2.16e-11 | 78.4205 | -1.8306e-09 | 78.5355
-2.0464e-11 | 78.8255 | -1.7872e-09 | 78.9395
-6.1391e-11 | 79.2315 | -1.8101e-09 | 79.3435
-2.251e-11 | 79.6376 | -1.8313e-09 | 79.7486
-1.7963e-11 | 80.0426 | -1.6651e-09 | 80.1526
-2.1828e-11 | 80.4476 | -1.7544e-09 | 80.5576
-3.2742e-11 | 80.8526 | -1.7421e-09 | 80.9626
-1.9327e-11 | 81.2587 | -1.6717e-09 | 81.3657
-1.9327e-11 | 81.6627 | -1.7355e-09 | 81.7687
-2.1373e-11 | 82.0687 | -1.9218e-09 | 82.1727
-2.7285e-11 | 82.4747 | -1.6739e-09 | 82.5757
-2.0236e-11 | 82.8807 | -1.6064e-09 | 82.9808
-1.8645e-11 | 83.2868
-2.2965e-11 | 83.6918
-2.3192e-11 | 84.0958
-1.9782e-11 | 84.5018
-2.2737e-11 | 84.9049
-2.251e-11 | 85.3109
-2.4784e-11 | 85.7149
-3.5016e-11 | 86.1209
-1.8417e-11 | 86.5259
-2.1828e-11 | 86.932
-2.8876e-11 | 87.336
-2.3647e-11 | 87.741
-1.9782e-11 | 88.146
-2.0918e-11 | 88.5521
-3.0013e-11 | 88.9571
-2.0918e-11 | 89.3621
-1.8645e-11 | 89.7671
-2.2965e-11 | 90.1732
-2.4102e-11 | 90.5782
-2.4329e-11 | 90.9852
-2.251e-11 | 91.3902
-2.3647e-11 | 91.7952
-2.3647e-11 | 92.2013
-2.16e-11 | 92.6063
-2.3874e-11 | 93.0123
-2.4784e-11 | 93.4163
-2.4784e-11 | 93.8224
-1.9327e-11 | 94.2284
-2.0918e-11 | 94.6334
-3.5698e-11 | 95.0394
-2.8876e-11 | 95.4445
-2.1373e-11 | 95.8495
-2.16e-11 | 96.2545
-2.5238e-11 | 96.6585
-2.2737e-11 | 97.0645
-2.4329e-11 | 97.4696
-2.3192e-11 | 97.8756
-4.3656e-11 | 98.2816
-3.0695e-11 | 98.6866
-2.5011e-11 | 99.0927
-2.3874e-11 | 99.4987
-2.5238e-11 | 99.9037
-2.5693e-11 | 100.3087
-2.7285e-11 | 100.7148
-2.4102e-11 | 101.1198
-3.115e-11 | 101.5248
-2.251e-11 | 101.9288
-2.6375e-11 | 102.3349
-3.3424e-11 | 102.7409
-2.6148e-11 | 103.1469
-2.3874e-11 | 103.5519
-2.5693e-11 | 103.9579
-2.6375e-11 | 104.363
-2.251e-11 | 104.769
-2.2737e-11 | 105.175
-2.4102e-11 | 105.58
-2.2737e-11 | 105.9861
-2.16e-11 | 106.3901
-2.683e-11 | 106.7951
-3.6607e-11 | 107.2001
-2.16e-11 | 107.6062
-2.16e-11 | 108.0122
-2.6375e-11 | 108.4182
-2.4784e-11 | 108.8222
-2.4784e-11 | 109.2282
-2.1373e-11 | 109.6323
-2.6148e-11 | 110.0383
-3.2742e-11 | 110.4423
-2.5011e-11 | 110.8473
-2.5238e-11 | 111.2544
-2.8649e-11 | 111.6584
-2.8876e-11 | 112.0634
-2.7285e-11 | 112.4694
-2.4102e-11 | 112.8745
-3.1378e-11 | 113.2795
-2.5693e-11 | 113.6845
-2.1828e-11 | 114.0905
-3.0241e-11 | 114.4955
-2.7967e-11 | 114.9016
-2.7512e-11 | 115.3066
-3.4333e-11 | 115.7126
-2.4102e-11 | 116.1176
-3.115e-11 | 116.5227
-2.683e-11 | 116.9277
-2.0009e-11 | 117.3337
-3.0468e-11 | 117.7387
-2.7967e-11 | 118.1438
-2.683e-11 | 118.5478
-3.0241e-11 | 118.9518
-2.5693e-11 | 119.3588
-2.774e-11 | 119.7628
-2.683e-11 | 120.1679
-2.5466e-11 | 120.5729
-3.5016e-11 | 120.9789
-2.9559e-11 | 121.3849
-2.4784e-11 | 121.793
-2.2283e-11 | 122.197
-2.774e-11 | 122.605
-2.9104e-11 | 123.011
-3.0468e-11 | 123.4151
-2.3192e-11 | 123.8211
-2.9559e-11 | 124.2271
-2.683e-11 | 124.6331
-2.1828e-11 | 125.0382
-2.6148e-11 | 125.4432
-2.9104e-11 | 125.8502
-2.3192e-11 | 126.2552
-2.3647e-11 | 126.6612
-2.683e-11 | 127.0663
-2.8649e-11 | 127.4703
-3.0468e-11 | 127.8743
-2.5466e-11 | 128.2793
-2.683e-11 | 128.6854
-3.0695e-11 | 129.0894
-2.6148e-11 | 129.4954
-2.2283e-11 | 129.9004
-2.7285e-11 | 130.3065
-2.7285e-11 | 130.7125
-3.0468e-11 | 131.1175
-2.8876e-11 | 131.5215
-2.8876e-11 | 131.9275
-2.9331e-11 | 132.3326
-2.4784e-11 | 132.7386
-2.5466e-11 | 133.1446
-4.3656e-11 | 133.5496
-2.9331e-11 | 133.9567
-2.6148e-11 | 134.3607
-3.0013e-11 | 134.7657
-3.115e-11 | 135.1707
-2.9331e-11 | 135.5758
-2.5011e-11 | 135.9818
-2.7512e-11 | 136.3878
-2.8876e-11 | 136.7918
-2.8876e-11 | 137.1978
-2.4784e-11 | 137.6019
-2.7285e-11 | 138.0059
-2.9331e-11 | 138.4119
-2.7512e-11 | 138.8179
-2.6603e-11 | 139.224
-5.2296e-11 | 139.629
-2.8194e-11 | 140.034
-2.5011e-11 | 140.439
-3.0468e-11 | 140.8451
-2.9104e-11 | 141.2491
-3.0695e-11 | 141.6541
-2.7512e-11 | 142.0601
-2.7512e-11 | 142.4641
-3.5016e-11 | 142.8702
-2.9331e-11 | 143.2752
-2.683e-11 | 143.6792
-2.6603e-11 | 144.0842
-2.5238e-11 | 144.4893
-3.9563e-11 | 144.8953
-2.7285e-11 | 145.3013
-2.8876e-11 | 145.7073
-3.1832e-11 | 146.1114
-3.7289e-11 | 146.5174
-2.6375e-11 | 146.9234
-2.774e-11 | 147.3274
-3.206e-11 | 147.7324
-2.8876e-11 | 148.1385
-2.774e-11 | 148.5435
-2.8876e-11 | 148.9475
-3.7971e-11 | 149.3535
-3.0468e-11 | 149.7596
-2.7512e-11 | 150.1656
-2.9331e-11 | 150.5706
-2.9331e-11 | 150.9756
-2.8194e-11 | 151.3807
-2.774e-11 | 151.7867
-3.206e-11 | 152.1917
-3.0468e-11 | 152.5957
-4.9113e-11 | 153.0008
-2.6603e-11 | 153.4058
-3.0013e-11 | 153.8108
-3.6607e-11 | 154.2148
-2.683e-11 | 154.6208
-3.0241e-11 | 155.0289
-3.6607e-11 | 155.4339
-3.5016e-11 | 155.8389
-2.774e-11 | 156.2439
-2.8876e-11 | 156.65
-2.8194e-11 | 157.055
-3.1832e-11 | 157.46
-3.0241e-11 | 157.865
-3.2969e-11 | 158.2721
-3.0695e-11 | 158.6771
-2.7967e-11 | 159.0831
-2.9331e-11 | 159.4881
-3.2514e-11 | 159.8941
-3.1832e-11 | 160.2992
-2.9559e-11 | 160.7042
-3.0695e-11 | 161.1102
-3.2742e-11 | 161.5152
-3.1605e-11 | 161.9213
-2.9331e-11 | 162.3263
-2.7967e-11 | 162.7313
-4.7294e-11 | 163.1373
-3.2742e-11 | 163.5414
-2.9331e-11 | 163.9464
-3.0695e-11 | 164.3514
-3.4333e-11 | 164.7584
-3.2969e-11 | 165.1634
-2.9331e-11 | 165.5685
-2.9559e-11 | 165.9735
-3.115e-11 | 166.3785
-3.3879e-11 | 166.7835
-3.0695e-11 | 167.1886
-2.7967e-11 | 167.5936
-3.5016e-11 | 167.9986
-3.4333e-11 | 168.4036
-1.5461e-11 | 168.8097
-2.6603e-11 | 169.2147
-3.1605e-11 | 169.6177
-3.0241e-11 | 170.0227
-2.9559e-11 | 170.4277
-2.8876e-11 | 170.8318
-5.3433e-11 | 171.2388
-2.8194e-11 | 171.6438
-3.206e-11 | 172.0488
-3.2969e-11 | 172.4539
-3.547e-11 | 172.8599
-2.8194e-11 | 173.2649
-3.1605e-11 | 173.6709
-3.3424e-11 | 174.077
-3.7289e-11 | 174.481
-2.9559e-11 | 174.886
-3.0695e-11 | 175.292
-3.0241e-11 | 175.698

## 710uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 720uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 730uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 740uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 750uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 760uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 770uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 780uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 790uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 800uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 810uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On