## Data Summary
QC5 Effective gain
Username: | Christos Roskas | QC5 Template | V2 |  | Environment |  |  |  |  | For Measuring Rate |  |  |  | For Measuring Current
Amplifier Settings |  | Date | 2018-04-05 |  | Vmon | Imon | Time | Pressure | Temp | Source Off |  | Source On |  | Source Off |  | Source On
[10 empty] | Counts | Err | Counts | Err | Current | Err | Current | Err
Make & Model: | ORTEC 474 |  |  |  | (V) | (uA) | (HH:MM) | (mbar) | (Deg C) | (N) | (N) | (N) | (N) | (A) | (A) | (A) | (A)
Coarse Gain: | 4x |  |  | Ar/CO2:70/30 Range | 3499.6 | 700 | 13:00 | 1006 | 22 | 368 | 19.1833 | 111811 | 334.3815 | -1.7362e-11 | 5.6878e-12 | -9.1142e-10 | 6.4454e-11
Fine Gain: | 4.5 |  |  |  | 3449.6 | 690 | 13:10 | 1006 | 22 | 293 | 17.1172 | 107751 | 328.2545 | -1.0173e-11 | 3.0022e-12 | -6.48e-10 | 5.0319e-11
Diff Time (ns): | 100 |  |  |  | 3399.8 | 680 | 13:20 | 1006 | 21.9 | 279 | 16.7033 | 100975 | 317.7656 | -8.3928e-12 | 2.1697e-12 | -4.5082e-10 | 3.4871e-11
Int Time (ns): | 100 |  |  |  | 3349.8 | 670 | 13:30 | 1006 | 21.8 | 202 | 14.2127 | 87947 | 296.5586 | -6.4082e-12 | 1.6449e-12 | -3.1701e-10 | 2.274e-11
Detector Settings |  |  |  |  | 3300 | 660 | 13:40 | 1006 | 22 | 132 | 11.4891 | 51365 | 226.6385 | -5.61e-12 | 1.4536e-12 | -2.276e-10 | 1.555e-11
[5 empty] | 3250 | 650 | 13:50 | 1006 | 22 | 76 | 8.7178 | 38666 | 196.6367 | -4.2993e-12 | 5.4629e-12 | -1.612e-10 | 1.0926e-11
Serial Number: | GE1/1-X-L-GHENT-0010 |  |  |  | 3200 | 640 | 14:00 | 1006 | 21.9 | 54 | 7.3485 | 21512 | 146.6697 | -3.7816e-12 | 1.2229e-12 | -1.1421e-10 | 7.8612e-12
Position (ieta,iphi): | 4.2 |  |  |  | 3150.2 | 630 | 14:10 | 1006 | 22 | 29 | 5.3852 | 5184 | 72 | -2.5416e-12 | 1.1449e-12 | -8.23e-11 | 5.8498e-12
Gas (X/Y/Z): | Ar/CO2 |  |  |  | 3100.4 | 620 | 14:20 | 1006 | 22.1 | 6 | 2.4495 | 165 | 12.8452 | -1.991e-12 | 2.5218e-12 | -5.8737e-11 | 4.2557e-12
Gas Frac (%X/%Y/%Z): | 70/30 |  |  |  | 3050.4 | 610 | 14:30 | 1006 | 22 | 2 | 1.4142 | 9 | 3 | -1.4536e-12 | 1.2518e-12 | -4.1791e-11 | 3.0499e-12
Flow Rate (L/hr): | 5 |  |  |  | 3001 | 600 | 14:40 | 1006 | 22 | 5 | 2.2361 | 8 | 2.8284 | -1.2445e-12 | 1.1404e-12 | -3.0576e-11 | 2.4799e-12
HV Circuit R_Equiv (MOhm): | 5 |  |  |  | 2951.2 | 590 | 14:50 | 1006 | 21.9 | 2 | 1.4142 | 7 | 2.6458 | -6.2351e-13 | 1.0642e-12 | -2.2181e-11 | 2.0045e-12
Total HV Divider R (MOhm): | 4.7 |  |  |  | 2901.2 | 580 | 15:00 | 1006 | 21.9 | 3 | 1.7321 | 8 | 2.8284 | -3.8283e-13 | 1.0057e-12 | -1.5735e-11 | 1.5844e-12
R_Drift (MOhm): | 1.126 |  |  |  | 2851.4 | 570 | 15:10 | 1006 | 21.8 | 1 | 1 | 6 | 2.4495 | 2.3625e-13 | 1.2587e-12 | -1.1134e-11 | 1.3843e-12
R_GEM1 (MOhm): | 0.563 |  |  |  | 2801.4 | 560 | 15:20 | 1006 | 21.9 | 1 | 1 | 5 | 2.2361 | 2.5431e-13 | 1.0736e-12 | -7.7679e-12 | 1.243e-12
R_Transfer1 (MOhm): | 0.438 |  |  |  | 2751.6 | 550 | 15:30 | 1006 | 22 | 1 | 1 | 5 | 2.2361 | 7.7832e-13 | 1.0708e-12 | -5.6739e-12 | 1.1752e-12
R_GEM2 (MOhm): | 0.551
R_Transfer2 (MOhm): | 0.875 |  |  |  |  | Fe55 | Cd109 | Cu X-Ray
R_GEM3 (MOhm): | 0.525 |  |  | N_primary | 346 | 196 | 322 | 346 | Electron Charge |  | 1.602e-19
R_Induction (MOhm): | 0.626 |  |  | Sigma N_primary | 2.9 | 1.8 | 2.8 | 2.9
Discriminator Settings |  |  |  | After T/P Corrections
Make & Model: | ORTEC 935 |  |  | Imon Equiv | V_mon | V_Drift | Rate | Rate Err | Current | Current Err | Gain | Gain Err
Threshold (mV): | -65 |  |  | (uA) | (V) | (V) | (Hz) | (Hz) | (A) | (A) | (A.U.) | (A.U.)
Walk Adjust (mV): | -0.001 |  |  | 666.5363 | 3332.3007 | #VALUE! | 1857.3833 | 5.5822 | -8.9406e-10 | 6.4705e-11 | 8684.1279 | 633.2232
Width (ns): | 180 |  |  | 657.0144 | 3284.691 | #VALUE! | 1790.9667 | 5.4783 | -6.3782e-10 | 5.0409e-11 | 6195.2539 | 492.7268
Picoammeter Settings |  |  |  | 647.4924 | 3237.2717 | #VALUE! | 1678.2667 | 5.3034 | -4.4243e-10 | 3.4939e-11 | 4297.3722 | 341.5167
|  |  |  | 637.9705 | 3189.6619 | #VALUE! | 1462.4167 | 4.9483 | -3.106e-10 | 2.28e-11 | 3016.9039 | 223.0795
Make & Model: | Keithley 6487 |  |  | 628.4485 | 3142.2426 | #VALUE! | 853.8833 | 3.7822 | -2.2199e-10 | 1.5618e-11 | 2156.253 | 152.912
Triax Red: | Signal |  |  | 618.9266 | 3094.6329 | #VALUE! | 643.1667 | 3.2805 | -1.569e-10 | 1.2215e-11 | 1523.9898 | 119.421
Triax Black: | ground |  |  | 609.4046 | 3047.0232 | #VALUE! | 357.6333 | 2.4476 | -1.1042e-10 | 7.9558e-12 | 1072.5614 | 77.8633
Triax Green: | ground |  |  | 599.8827 | 2999.6039 | #VALUE! | 85.9167 | 1.2034 | -7.9758e-11 | 5.9608e-12 | 774.7011 | 58.3076
Scalar Settings |  |  |  | 590.3607 | 2952.1846 | #VALUE! | 2.65 | 0.21794 | -5.6746e-11 | 4.9468e-12 | 551.1833 | 48.2989
|  |  |  | 580.8388 | 2904.5748 | #VALUE! | 0.11667 | 0.055277 | -4.0337e-11 | 3.2967e-12 | 391.7989 | 32.2112
|  |  |  | 571.3168 | 2857.5364 | #VALUE! | 0.05 | 0.060093 | -2.9332e-11 | 2.7295e-12 | 284.9043 | 26.6336
Make & Model: | ORTEC 974 |  |  | 561.7949 | 2810.1171 | #VALUE! | 0.083333 | 0.05 | -2.1557e-11 | 2.2695e-12 | 209.3869 | 22.1225
Acquisition Time (s): | 60 |  |  | 552.2729 | 2762.5074 | #VALUE! | 0.083333 | 0.055277 | -1.5352e-11 | 1.8766e-12 | 149.1148 | 18.2764
Source Settings |  |  |  | 542.751 | 2715.0881 | #VALUE! | 0.083333 | 0.044096 | -1.137e-11 | 1.871e-12 | 110.4382 | 18.1998
|  |  |  | 533.2291 | 2667.4783 | #VALUE! | 0.066667 | 0.040825 | -8.0222e-12 | 1.6425e-12 | 77.9206 | 15.9687
Source: | Ag (Silver) |  |  | 523.7071 | 2620.0591 | #VALUE! | 0.066667 | 0.040825 | -6.4522e-12 | 1.5899e-12 | 62.6712 | 15.4527
HV (kV): | 40
Current (uA): | 5 |  |  | Summary |  |  | Correction parameters (P5) |  |  |  | For Ana. Framwork
Activity (MBq): | Enter value |  |  | Avrg temperature (C) | 295.1 |  | P0 | 964.4 |  |  | 666,536318942621,8684,12787096408,633,223154085955
Filter Status: | 3Cu tape layer + 5mm Al filter |  |  | Err temperature (C) | 0.079057 |  | T0 | 297.1 |  |  | 657,014371529155,6195,25390199678,492,726752468736
Collimator Status: | Removed |  |  | Avrg pressure (mbar) | 1006 | [5 empty] | 647,492424115689,4297,37219901696,341,516737621952
|  |  |  | Err pressure (mbar) | 0 | [5 empty] | 637,970476702223,3016,90390708619,223,079530375398
|  |  |  | Expo_fit_1_param_1 | 6.2971e-07 | [5 empty] | 628,448529288757,2156,25302501424,152,911968028814
|  |  |  | Expo_fit_1_param_2 | 0.034945 | [5 empty] | 618,926581875291,1523,9897849875,119,421039803819
609,404634461825,1072,56143817742,77,8632725606408
599,882687048359,774,701121252752,58,3075692595878
590,360739634893,551,183299257074,48,2988998788842
580,838792221427,391,798875319985,32,2112353180728
571,316844807961,284,904322059001,26,633571660944
561,794897394495,209,386932160293,22,122458786892
552,272949981029,149,114847023711,18,2763720269177
542,751002567563,110,438237939855,18,1997645189751
533,229055154097,77,9206464168817,15,9686827738859
523,707107740631,62,6712035645046,15,452667547258

## 550uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
7.7832e-13 | 1.0708e-12 | -5.6739e-12 | 1.1752e-12
GEMs On |  | GEMs On
9.0949e-13 | 0.31102 | -6.8212e-12 | 0.31202
-4.5475e-13 | 0.99506 | -5.0022e-12 | 0.99706
1.819e-12 | 1.3991 | -6.3665e-12 | 1.4031
-6.8212e-13 | 1.8031 | -6.1391e-12 | 1.8081
1.5916e-12 | 2.2071 | -4.3201e-12 | 2.2131
9.0949e-13 | 2.6122 | -5.0022e-12 | 2.6191
6.8212e-13 | 3.0182 | -5.0022e-12 | 3.0232
1.3642e-12 | 3.4212 | -4.7748e-12 | 3.4272
1.819e-12 | 3.8272 | -6.1391e-12 | 3.8352
1.819e-12 | 4.2332 | -5.2296e-12 | 4.2402
6.8212e-13 | 4.6433 | -6.3665e-12 | 4.6453
2.0464e-12 | 5.0783 | -5.6843e-12 | 5.0503
2.0464e-12 | 5.4833 | -6.1391e-12 | 5.4553
-1.1369e-12 | 5.8883 | -4.0927e-12 | 5.8593
-9.0949e-13 | 6.2934 | -5.2296e-12 | 6.2644
-6.8212e-13 | 6.6974 | -5.6843e-12 | 6.6694
2.9559e-12 | 7.1034 | -4.3201e-12 | 7.0744
2.2737e-13 | 7.5084 | -5.457e-12 | 7.4804
4.5475e-13 | 7.9115 | -7.0486e-12 | 7.8865
3.1832e-12 | 8.3165 | -5.2296e-12 | 8.2925
-2.2737e-13 | 8.7215 | -4.7748e-12 | 8.6985
2.2737e-13 | 9.1265 | -3.8654e-12 | 9.1035
1.5916e-12 | 9.5305 | -5.457e-12 | 9.5085
2.0464e-12 | 9.9356 | -6.3665e-12 | 9.9146
1.3642e-12 | 10.3406 | -5.457e-12 | 10.3196
2.2737e-13 | 10.7446 | -6.8212e-12 | 10.7496
6.8212e-13 | 11.1496 | -6.8212e-12 | 11.1546
6.8212e-13 | 11.5527 | -3.638e-12 | 11.5757
9.0949e-13 | 11.9577 | -7.7307e-12 | 11.9867
2.7285e-12 | 12.4167 | -8.1855e-12 | 12.3917
1.3642e-12 | 12.8227 | -4.0927e-12 | 12.7967
-6.8212e-13 | 13.2288 | -4.0927e-12 | 13.2018
4.5475e-13 | 13.6448 | -7.0486e-12 | 13.6078
-2.2737e-13 | 14.0498 | -7.5033e-12 | 14.0138
1.3642e-12 | 14.4548 | -5.457e-12 | 14.4188
-2.2737e-13 | 14.8619 | -5.6843e-12 | 14.8239
-2.2737e-13 | 15.2659 | -6.8212e-12 | 15.2299
4.5475e-13 | 15.6889 | -6.8212e-12 | 15.6349
3.1832e-12 | 16.0939 | -5.457e-12 | 16.0399
2.0464e-12 | 16.583 | -5.0022e-12 | 16.4459
6.8212e-13 | 16.999 | -6.5938e-12 | 16.851
2.2737e-13 | 17.409 | -6.5938e-12 | 17.256
1.5916e-12 | 17.814 | -5.6843e-12 | 17.661
3.1832e-12 | 18.22 | -7.9581e-12 | 18.09
9.0949e-13 | 18.7441 | -3.8654e-12 | 18.5221
-2.2737e-13 | 19.1831 | -5.6843e-12 | 18.9331
1.3642e-12 | 19.5891 | -7.0486e-12 | 19.3451
2.2737e-13 | 19.9971 | -4.3201e-12 | 19.7701
-9.0949e-13 | 20.4012 | -7.7307e-12 | 20.1761
6.8212e-13 | 20.8042 | -5.0022e-12 | 20.5812
4.5475e-13 | 21.2082 | -4.7748e-12 | 20.9882
1.819e-12 | 21.6152 | -6.5938e-12 | 21.3932
1.3642e-12 | 22.0203 | -6.5938e-12 | 21.8012
-6.8212e-13 | 22.4273 | -5.2296e-12 | 22.2063
3.1832e-12 | 22.8333 | -5.6843e-12 | 22.6143
3.638e-12 | 23.2403 | -7.5033e-12 | 23.0573
6.8212e-13 | 23.6454 | -6.3665e-12 | 23.4623
2.2737e-13 | 24.0494 | -6.1391e-12 | 23.8674
4.5475e-13 | 24.4534 | -5.6843e-12 | 24.2864
1.3642e-12 | 24.8574 | -6.1391e-12 | 24.7554
2.0464e-12 | 25.2624 | -4.7748e-12 | 25.1764
-9.0949e-13 | 25.6675 | -7.0486e-12 | 25.5925
-6.8212e-13 | 26.0715 | -6.1391e-12 | 25.9985
3.1832e-12 | 26.4765 | -6.1391e-12 | 26.4025
1.819e-12 | 26.8795 | -4.7748e-12 | 26.8075
2.0464e-12 | 27.2836 | -5.457e-12 | 27.2126
2.2737e-13 | 27.6896 | -4.0927e-12 | 27.6176
9.0949e-13 | 28.0936 | -5.0022e-12 | 28.0236
1.819e-12 | 28.4986 | -4.7748e-12 | 28.4336
2.2737e-13 | 28.9027 | -5.6843e-12 | 28.8387
2.2737e-13 | 29.3067 | -6.8212e-12 | 29.2437
-6.8212e-13 | 29.7117 | -4.0927e-12 | 29.6497
4.5475e-13 | 30.1157 | -6.3665e-12 | 30.0547
3.638e-12 | 30.5207 | -7.9581e-12 | 30.4587
-6.8212e-13 | 30.9248 | -3.8654e-12 | 30.8668
6.8212e-13 | 31.3288 | -6.1391e-12 | 31.2738
2.2737e-13 | 31.7328 | -7.9581e-12 | 31.6798
-2.2737e-13 | 32.1358 | -7.5033e-12 | 32.0838
6.8212e-13 | 32.5409 | -3.8654e-12 | 32.4889
2.0464e-12 | 32.9429 | -6.1391e-12 | 32.8939
-2.2737e-13 | 33.3479 | -7.9581e-12 | 33.2999
6.8212e-13 | 33.7509 | -6.8212e-12 | 33.7059
1.3642e-12 | 34.1549 | -3.8654e-12 | 34.112
1.5916e-12 | 34.56 | -7.7307e-12 | 34.517
1.3642e-12 | 34.965 | -5.0022e-12 | 34.924
2.0464e-12 | 35.369 | -4.0927e-12 | 35.329
6.8212e-13 | 35.773 | -5.6843e-12 | 35.734
6.8212e-13 | 36.1781 | -5.457e-12 | 36.1391
4.5475e-13 | 36.5811 | -6.8212e-12 | 36.5451
4.5475e-13 | 36.9851 | -5.457e-12 | 36.9491
4.5475e-13 | 37.3891 | -5.457e-12 | 37.3551
-6.8212e-13 | 37.7932 | -7.0486e-12 | 37.7602
1.819e-12 | 38.1972 | -5.6843e-12 | 38.1652
1.819e-12 | 38.6012 | -7.0486e-12 | 38.5702
6.8212e-13 | 39.0062 | -5.457e-12 | 38.9752
1.5916e-12 | 39.4093 | -6.3665e-12 | 39.3813
6.8212e-13 | 39.8143 | -5.0022e-12 | 39.7853
-4.5475e-13 | 40.2193 | -5.2296e-12 | 40.1903
4.5475e-13 | 40.6223 | -6.3665e-12 | 40.5963
9.0949e-13 | 41.0273 | -5.0022e-12 | 41.0003
2.5011e-12 | 41.4304 | -4.7748e-12 | 41.4054
6.8212e-13 | 41.8364 | -5.0022e-12 | 41.8104
4.5475e-13 | 42.2394 | -4.0927e-12 | 42.2164
1.819e-12 | 42.6434 | -5.6843e-12 | 42.6214
6.8212e-13 | 43.0485 | -6.1391e-12 | 43.0285
-9.0949e-13 | 43.4525 | -5.6843e-12 | 43.4335
2.5011e-12 | 43.8585 | -4.3201e-12 | 43.8385
1.3642e-12 | 44.2635 | -6.5938e-12 | 44.2425
6.8212e-13 | 44.6686 | -5.0022e-12 | 44.6465
2.2737e-13 | 45.0726 | -7.0486e-12 | 45.0516
1.819e-12 | 45.4766 | -5.2296e-12 | 45.4566
1.5916e-12 | 45.8806 | -5.0022e-12 | 45.8616
9.0949e-13 | 46.2856 | -5.457e-12 | 46.2657
1.5916e-12 | 46.6897 | -5.457e-12 | 46.6707
-2.2737e-13 | 47.0947 | -2.9559e-12 | 47.0757
6.8212e-13 | 47.4987 | -7.7307e-12 | 47.4817
3.1832e-12 | 47.9027 | -6.8212e-12 | 47.8867
6.8212e-13 | 48.3078 | -5.6843e-12 | 48.2928
-4.5475e-13 | 48.7118 | -4.0927e-12 | 48.6978
-9.0949e-13 | 49.1168 | -5.457e-12 | 49.1028
-9.0949e-13 | 49.5198 | -4.7748e-12 | 49.5088
9.0949e-13 | 49.9229 | -7.0486e-12 | 49.9129
-2.2737e-13 | 50.3269 | -2.9559e-12 | 50.3189
6.8212e-13 | 50.7309 | -5.0022e-12 | 50.7239
-2.2737e-13 | 51.1359 | -6.8212e-12 | 51.1289
-9.0949e-13 | 51.5389 | -5.2296e-12 | 51.5339
3.1832e-12 | 51.944 | -5.0022e-12 | 51.939
1.819e-12 | 52.349 | -4.7748e-12 | 52.344
-4.5475e-13 | 52.754 | -6.1391e-12 | 52.75
1.3642e-12 | 53.158 | -6.3665e-12 | 53.154
-1.1369e-12 | 53.5621 | -4.0927e-12 | 53.5591
2.0464e-12 | 53.9661 | -5.457e-12 | 53.9631
6.8212e-13 | 54.3711 | -8.8676e-12 | 54.3681
6.8212e-13 | 54.7751 | -5.6843e-12 | 54.7721
9.0949e-13 | 55.1792 | -5.0022e-12 | 55.1772
2.0464e-12 | 55.5832 | -4.7748e-12 | 55.5822
2.0464e-12 | 55.9882 | -5.6843e-12 | 55.9862
1.3642e-12 | 56.3932 | -7.0486e-12 | 56.3912
-1.1369e-12 | 56.7963 | -5.6843e-12 | 56.7953
2.2737e-13 | 57.2013 | -3.638e-12 | 57.1983
-2.2737e-13 | 57.6063 | -6.1391e-12 | 57.6043
1.5916e-12 | 58.0113 | -7.7307e-12 | 58.0083
-6.8212e-13 | 58.4163 | -5.6843e-12 | 58.4143
2.2737e-13 | 58.8204 | -6.1391e-12 | 58.8184
1.5916e-12 | 59.2244 | -5.2296e-12 | 59.2234
-9.0949e-13 | 59.6294 | -9.3223e-12 | 59.6284
6.8212e-13 | 60.0334 | -5.0022e-12 | 60.0344
9.0949e-13 | 60.4395 | -5.0022e-12 | 60.4405
6.8212e-13 | 60.8445 | -5.6843e-12 | 60.8455
9.0949e-13 | 61.2475 | -5.6843e-12 | 61.2505
-4.5475e-13 | 61.6515 | -6.1391e-12 | 61.6565
1.3642e-12 | 62.0555 | -4.7748e-12 | 62.0615
9.0949e-13 | 62.4616 | -7.5033e-12 | 62.4666
1.3642e-12 | 62.8646 | -5.6843e-12 | 62.8706
-2.2737e-13 | 63.2686 | -5.2296e-12 | 63.2756
-4.5475e-13 | 63.6746 | -5.2296e-12 | 63.6796
9.0949e-13 | 64.0787 | -6.1391e-12 | 64.0857
1.3642e-12 | 64.4837 | -6.1391e-12 | 64.4907
1.3642e-12 | 64.8887 | -5.2296e-12 | 64.8947
2.2737e-13 | 65.2937 | -2.7285e-12 | 65.3007
2.2737e-13 | 65.6978 | -5.6843e-12 | 65.7048
2.2737e-13 | 66.1028 | -4.7748e-12 | 66.1098
2.0464e-12 | 66.5068 | -5.2296e-12 | 66.5148
6.8212e-13 | 66.9118 | -4.3201e-12 | 66.9198
1.5916e-12 | 67.3149 | -4.0927e-12 | 67.3248
2.0464e-12 | 67.7199 | -6.5938e-12 | 67.7299
-2.2737e-13 | 68.1239 | -5.6843e-12 | 68.1359
1.5916e-12 | 68.5289 | -2.9559e-12 | 68.5399
9.0949e-13 | 68.9339 | -5.0022e-12 | 68.9449
2.2737e-13 | 69.338 | -5.2296e-12 | 69.35
2.0464e-12 | 69.742 | -5.457e-12 | 69.755
-4.5475e-13 | 70.147 | -5.2296e-12 | 70.16
1.5916e-12 | 70.551 | -6.1391e-12 | 70.567
-4.5475e-13 | 70.9551 | -5.6843e-12 | 70.9721
2.2737e-13 | 71.3591
2.2737e-13 | 71.7641
1.5916e-12 | 72.1681
2.2737e-13 | 72.5731
-1.5916e-12 | 72.9762
6.8212e-13 | 73.3812
1.5916e-12 | 73.7852
-2.2737e-13 | 74.1892
-9.0949e-13 | 74.5933
-4.5475e-13 | 74.9993
2.2737e-13 | 75.4023
4.5475e-13 | 75.8053
6.8212e-13 | 76.2114
4.5475e-13 | 76.6154
4.5475e-13 | 77.0204
9.0949e-13 | 77.4434
4.5475e-13 | 77.8474
6.8212e-13 | 78.2525
4.5475e-13 | 78.6565
2.5011e-12 | 79.0605
4.5475e-13 | 79.5055
6.8212e-13 | 79.9146
-4.5475e-13 | 80.3186
-1.1369e-12 | 80.7226
1.3642e-12 | 81.1276
1.3642e-12 | 81.5317
6.8212e-13 | 81.9357
-9.0949e-13 | 82.3407
-4.5475e-13 | 82.7437
2.7285e-12 | 83.1478
1.819e-12 | 83.5528
1.3642e-12 | 83.9578
2.2737e-13 | 84.3628
4.5475e-13 | 84.7678

## 560uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
2.5431e-13 | 1.0736e-12 | -7.7679e-12 | 1.243e-12
GEMs On |  | GEMs On
1.3642e-12 | 0.31302 | -7.9581e-12 | 0.31302
-9.0949e-13 | 0.99706 | -8.6402e-12 | 1.0001
-2.2737e-13 | 1.4021 | -1.0459e-11 | 1.4041
-2.2737e-13 | 1.8061 | -6.1391e-12 | 1.8101
2.2737e-13 | 2.2091 | -9.0949e-12 | 2.2161
1.5916e-12 | 2.6131 | -6.8212e-12 | 2.6202
0 | 3.0182 | -7.7307e-12 | 3.0262
2.2737e-13 | 3.4212 | -8.1855e-12 | 3.4312
4.5475e-13 | 3.8262 | -8.1855e-12 | 3.8432
-1.1369e-12 | 4.2312 | -6.5938e-12 | 4.2492
2.2737e-13 | 4.6373 | -7.7307e-12 | 4.6543
-1.1369e-12 | 5.0423 | -7.0486e-12 | 5.0583
0 | 5.4463 | -7.7307e-12 | 5.4653
-4.5475e-13 | 5.8483 | -8.6402e-12 | 5.8693
1.5916e-12 | 6.2524 | -6.1391e-12 | 6.2744
1.5916e-12 | 6.6574 | -6.5938e-12 | 6.6784
4.5475e-13 | 7.0624 | -6.5938e-12 | 7.0834
1.3642e-12 | 7.4664 | -8.1855e-12 | 7.4904
0 | 7.8715 | -7.7307e-12 | 7.8975
2.2737e-13 | 8.2725 | -7.9581e-12 | 8.3015
2.2737e-13 | 8.6765 | -6.8212e-12 | 8.7065
1.1369e-12 | 9.0805 | -8.1855e-12 | 9.1125
-2.2737e-13 | 9.4855 | -6.3665e-12 | 9.5165
-9.0949e-13 | 9.8896 | -8.1855e-12 | 9.9216
1.3642e-12 | 10.2926 | -8.6402e-12 | 10.3726
4.5475e-13 | 10.6976 | -6.1391e-12 | 10.7786
9.0949e-13 | 11.1016 | -7.7307e-12 | 11.1906
0 | 11.5067 | -8.6402e-12 | 11.6057
-2.2737e-13 | 11.9097 | -6.3665e-12 | 12.0237
1.3642e-12 | 12.3127 | -7.0486e-12 | 12.4317
1.3642e-12 | 12.7167 | -5.2296e-12 | 12.8367
-2.2737e-13 | 13.1218 | -5.457e-12 | 13.2438
0 | 13.5248 | -8.1855e-12 | 13.6498
0 | 13.9268 | -7.9581e-12 | 14.0568
-9.0949e-13 | 14.3308 | -6.8212e-12 | 14.4628
9.0949e-13 | 14.7348 | -8.1855e-12 | 14.8679
2.9559e-12 | 15.1389 | -9.0949e-12 | 15.2749
4.5475e-13 | 15.5429 | -7.5033e-12 | 15.6799
1.3642e-12 | 15.9479 | -6.5938e-12 | 16.0869
-9.0949e-13 | 16.3519 | -7.0486e-12 | 16.4949
2.2737e-12 | 16.756 | -6.8212e-12 | 16.913
2.2737e-13 | 17.16 | -8.6402e-12 | 17.327
-9.0949e-13 | 17.563 | -6.5938e-12 | 17.738
4.5475e-13 | 17.967 | -9.5497e-12 | 18.143
1.1369e-12 | 18.3711 | -8.8676e-12 | 18.5481
0 | 18.7771 | -7.0486e-12 | 18.9531
0 | 19.1791 | -6.5938e-12 | 19.3581
2.2737e-13 | 19.5811 | -6.8212e-12 | 19.7651
0 | 19.9861 | -7.7307e-12 | 20.1701
0 | 20.3892 | -1.0459e-11 | 20.5762
-2.2737e-13 | 20.7912 | -6.8212e-12 | 20.9812
1.1369e-12 | 21.1952 | -9.5497e-12 | 21.3852
-1.1369e-12 | 21.6002 | -8.6402e-12 | 21.7933
4.5475e-13 | 22.0053 | -8.8676e-12 | 22.1993
-2.2737e-13 | 22.4083 | -7.9581e-12 | 22.6043
0 | 22.8113 | -1.0687e-11 | 23.0133
1.1369e-12 | 23.2143 | -7.0486e-12 | 23.4203
0 | 23.6183 | -8.8676e-12 | 23.8254
-2.2737e-13 | 24.0214 | -7.5033e-12 | 24.2294
0 | 24.4244 | -7.9581e-12 | 24.6364
0 | 24.8274 | -6.5938e-12 | 25.0494
1.3642e-12 | 25.2294 | -6.1391e-12 | 25.4625
-9.0949e-13 | 25.6345 | -9.0949e-12 | 25.8695
1.1369e-12 | 26.0395 | -1.0459e-11 | 26.2765
0 | 26.4435 | -1.0232e-11 | 26.7045
4.5475e-13 | 26.8465 | -7.5033e-12 | 27.1095
-2.2737e-13 | 27.2496 | -9.5497e-12 | 27.5156
0 | 27.6536 | -6.5938e-12 | 27.9206
-2.2737e-13 | 28.0566 | -7.5033e-12 | 28.3246
2.9559e-12 | 28.4616 | -8.6402e-12 | 28.7296
-1.3642e-12 | 28.8666 | -7.5033e-12 | 29.1347
-1.1369e-12 | 29.2717 | -6.3665e-12 | 29.5407
4.5475e-13 | 29.6767 | -7.9581e-12 | 29.9527
4.5475e-13 | 30.0797 | -7.9581e-12 | 30.3577
-4.5475e-13 | 30.4827 | -7.9581e-12 | 30.7668
0 | 30.8868 | -7.7307e-12 | 31.1728
-2.2737e-13 | 31.2898 | -7.5033e-12 | 31.5768
-2.2737e-13 | 31.6938 | -6.8212e-12 | 31.9808
2.2737e-12 | 32.0958 | -6.5938e-12 | 32.3869
-4.5475e-13 | 32.5019 | -8.6402e-12 | 32.7919
9.0949e-13 | 32.9059 | -6.8212e-12 | 33.1969
-1.1369e-12 | 33.3109 | -7.9581e-12 | 33.6029
1.3642e-12 | 33.7139 | -7.7307e-12 | 34.008
-2.2737e-13 | 34.12 | -8.6402e-12 | 34.413
9.0949e-13 | 34.524 | -8.6402e-12 | 34.82
0 | 34.928 | -9.3223e-12 | 35.226
2.2737e-13 | 35.331 | -7.5033e-12 | 35.631
0 | 35.735 | -8.6402e-12 | 36.0401
2.7285e-12 | 36.1381 | -8.1855e-12 | 36.4671
9.0949e-13 | 36.5441 | -7.5033e-12 | 36.8731
9.0949e-13 | 36.9481 | -9.0949e-12 | 37.2781
-2.2737e-13 | 37.3521 | -6.5938e-12 | 37.6822
-1.5916e-12 | 37.7552 | -8.1855e-12 | 38.0882
2.2737e-13 | 38.1602 | -6.8212e-12 | 38.4932
-1.3642e-12 | 38.5642 | -6.5938e-12 | 38.8992
1.3642e-12 | 38.9682 | -8.6402e-12 | 39.3032
-2.9559e-12 | 39.3732 | -7.7307e-12 | 39.7093
9.0949e-13 | 39.7783 | -8.8676e-12 | 40.1153
9.0949e-13 | 40.1813 | -1.0004e-11 | 40.5213
2.2737e-13 | 40.5853 | -7.9581e-12 | 40.9273
2.2737e-13 | 40.9893 | -7.7307e-12 | 41.3324
0 | 41.3934 | -1.1141e-11 | 41.7384
1.1369e-12 | 41.7964 | -8.1855e-12 | 42.1434
-2.2737e-13 | 42.2014 | -6.1391e-12 | 42.5494
1.1369e-12 | 42.6064 | -8.1855e-12 | 42.9555
-1.3642e-12 | 43.0115 | -7.0486e-12 | 43.3605
-4.5475e-13 | 43.4155 | -6.8212e-12 | 43.7755
1.3642e-12 | 43.8195 | -6.3665e-12 | 44.1835
1.3642e-12 | 44.2235 | -8.6402e-12 | 44.5876
4.5475e-13 | 44.6285 | -9.3223e-12 | 44.9936
-1.1369e-12 | 45.0316 | -8.1855e-12 | 45.3986
-2.2737e-13 | 45.4356 | -6.1391e-12 | 45.8056
2.2737e-13 | 45.8396 | -7.0486e-12 | 46.2136
2.7285e-12 | 46.2416 | -8.8676e-12 | 46.6177
1.1369e-12 | 46.6467 | -9.0949e-12 | 47.0247
1.3642e-12 | 47.0537 | -9.0949e-12 | 47.4317
-9.0949e-13 | 47.4577 | -7.0486e-12 | 47.8367
1.5916e-12 | 47.8607 | -5.457e-12 | 48.2448
-9.0949e-13 | 48.2668 | -6.1391e-12 | 48.6508
-9.0949e-13 | 48.6718 | -9.0949e-12 | 49.0558
-9.0949e-13 | 49.0758 | -6.8212e-12 | 49.4608
9.0949e-13 | 49.4808 | -7.5033e-12 | 49.8659
1.5916e-12 | 49.8839 | -6.1391e-12 | 50.2719
-2.2737e-13 | 50.2899 | -8.1855e-12 | 50.6779
2.2737e-13 | 50.6939 | -8.1855e-12 | 51.0819
9.0949e-13 | 51.0979 | -5.457e-12 | 51.487
2.9559e-12 | 51.502 | -6.8212e-12 | 51.892
1.5916e-12 | 51.907 | -8.1855e-12 | 52.308
4.5475e-13 | 52.313 | -6.1391e-12 | 52.712
1.3642e-12 | 52.717 | -6.1391e-12 | 53.117
2.0464e-12 | 53.122 | -7.5033e-12 | 53.5221
0 | 53.5261 | -7.0486e-12 | 53.9261
4.5475e-13 | 53.9281 | -6.5938e-12 | 54.3311
2.2737e-13 | 54.3311 | -5.457e-12 | 54.7371
-1.3642e-12 | 54.7351 | -1.0232e-11 | 55.1412
1.3642e-12 | 55.1401 | -7.7307e-12 | 55.5472
2.2737e-13 | 55.5442 | -6.3665e-12 | 55.9522
9.0949e-13 | 55.9482 | -1.0232e-11 | 56.3562
0 | 56.3532 | -7.5033e-12 | 56.7732
1.1369e-12 | 56.7563 | -8.8676e-12 | 57.1803
1.1369e-12 | 57.1603 | -8.1855e-12 | 57.5863
2.0464e-12 | 57.5653 | -6.5938e-12 | 57.9923
-1.3642e-12 | 57.9693 | -1.0004e-11 | 58.3993
-1.1369e-12 | 58.3733 | -6.5938e-12 | 58.8044
0 | 58.7774 | -7.0486e-12 | 59.2084
-9.0949e-13 | 59.1804 | -8.6402e-12 | 59.6134
-1.1369e-12 | 59.5834 | -9.5497e-12 | 60.0194
2.7285e-12 | 59.9884 | -9.3223e-12 | 60.4245
-1.3642e-12 | 60.3944 | -7.7307e-12 | 60.8305
-2.2737e-13 | 60.7995 | -6.5938e-12 | 61.2355
9.0949e-13 | 61.2045 | -1.0004e-11 | 61.6405
1.1369e-12 | 61.6085 | -9.3223e-12 | 62.0455
-2.7285e-12 | 62.0136 | -7.0486e-12 | 62.4506
-1.5916e-12 | 62.4196 | -7.5033e-12 | 62.8586
1.1369e-12 | 62.8246 | -7.7307e-12 | 63.2656
2.2737e-13 | 63.2296 | -7.0486e-12 | 63.6696
2.2737e-13 | 63.6336 | -7.7307e-12 | 64.0757
-1.5916e-12 | 64.0387 | -9.3223e-12 | 64.4807
-4.5475e-13 | 64.4437 | -5.457e-12 | 64.8877
1.1369e-12 | 64.8467 | -9.0949e-12 | 65.2957
0 | 65.2507 | -8.8676e-12 | 65.6998
1.1369e-12 | 65.6527 | -6.5938e-12 | 66.1048
1.3642e-12 | 66.0578 | -8.1855e-12 | 66.5088
-2.2737e-13 | 66.4638 | -9.3223e-12 | 66.9268
1.1369e-12 | 66.8688 | -7.5033e-12 | 67.3338
-1.1369e-12 | 67.2738 | -6.8212e-12 | 67.7379
-2.2737e-13 | 67.6779 | -6.1391e-12 | 68.2079
-1.1369e-12 | 68.0809 | -1.0004e-11 | 68.6229
2.2737e-13 | 68.4869 | -9.3223e-12 | 69.031
1.3642e-12 | 68.8899 | -5.457e-12 | 69.436
-1.3642e-12 | 69.294 | -7.9581e-12 | 69.842
-4.5475e-13 | 69.698 | -5.6843e-12 | 70.325
9.0949e-13 | 70.102 | -8.1855e-12 | 70.746
-1.1369e-12 | 70.506 | -8.6402e-12 | 71.1591
9.0949e-13 | 70.9101 | -9.5497e-12 | 71.5651
-2.2737e-13 | 71.3151 | -9.5497e-12 | 71.9701
-1.3642e-12 | 71.7211 | -6.1391e-12 | 72.3751
2.2737e-13 | 72.1261 | -9.0949e-12 | 72.7792
9.0949e-13 | 72.5292 | -6.3665e-12 | 73.1952
2.2737e-12 | 72.9332 | -8.1855e-12 | 73.6002
-9.0949e-13 | 73.3392 | -6.8212e-12 | 74.0062
-2.2737e-13 | 73.7432 | -8.8676e-12 | 74.4123
2.2737e-12 | 74.1472 | -6.3665e-12 | 74.8173
1.3642e-12 | 74.5533 | -7.9581e-12 | 75.2223
1.3642e-12 | 74.9583 | -7.5033e-12 | 75.6263
-2.2737e-12 | 75.3623 | -6.5938e-12 | 76.0314
4.5475e-13 | 75.7683 | -7.0486e-12 | 76.4364
4.5475e-13 | 76.1714 | -6.8212e-12 | 76.8414
-2.2737e-13 | 76.5764 | -9.5497e-12 | 77.2464
9.0949e-13 | 76.9804 | -8.1855e-12 | 77.6534
0 | 77.3844 | -7.7307e-12 | 78.0595
-1.3642e-12 | 77.7875 | -8.8676e-12 | 78.5105
4.5475e-13 | 78.1925 | -7.9581e-12 | 78.9165
-9.0949e-13 | 78.5975 | -7.0486e-12 | 79.3215
-1.819e-12 | 79.0015 | -6.1391e-12 | 79.7256
-4.5475e-13 | 79.4075 | -9.3223e-12 | 80.1296
0 | 79.8116 | -8.6402e-12 | 80.5356
2.0464e-12 | 80.2146 | -7.0486e-12 | 80.9416
1.5916e-12 | 80.6196 | -6.5938e-12 | 81.3496
-2.2737e-13 | 81.0236 | -7.5033e-12 | 81.7577
2.2737e-13 | 81.4277 | -8.6402e-12 | 82.1617
1.1369e-12 | 81.8317 | -7.9581e-12 | 82.5687
1.3642e-12 | 82.2357 | -8.1855e-12 | 82.9737
-1.5916e-12 | 82.6407 | -4.0927e-12 | 83.3778
0 | 83.0447 | -8.6402e-12 | 83.7828
2.2737e-13 | 83.4468 | -7.0486e-12 | 84.1888
9.0949e-13 | 83.8508 | -8.6402e-12 | 84.5928
2.2737e-13 | 84.2548 | -6.8212e-12 | 84.9979
2.2737e-13 | 84.6588
1.1369e-12 | 85.0629
-1.1369e-12 | 85.4679

## 570uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
2.3625e-13 | 1.2587e-12 | -1.1134e-11 | 1.3843e-12
GEMs On |  | GEMs On
1.5916e-12 | 0.31202 | -1.0687e-11 | 0.31102
1.3642e-12 | 0.99606 | -1.296e-11 | 0.99706
6.8212e-13 | 1.4011 | -1.2733e-11 | 1.4021
-1.5916e-12 | 1.8041 | -1.1823e-11 | 1.8071
-6.8212e-13 | 2.2081 | -1.1596e-11 | 2.2121
1.3642e-12 | 2.6122 | -1.0459e-11 | 2.6202
2.0464e-12 | 3.0182 | -9.5497e-12 | 3.0442
6.8212e-13 | 3.4232 | -1.0232e-11 | 3.4502
-4.5475e-13 | 3.8272 | -1.0004e-11 | 3.8562
-1.1369e-12 | 4.2312 | -7.5033e-12 | 4.2632
2.5011e-12 | 4.6373 | -1.2051e-11 | 4.6693
9.0949e-13 | 5.0423 | -1.1596e-11 | 5.0833
-6.8212e-13 | 5.4453 | -1.0687e-11 | 5.4913
-1.3642e-12 | 5.8503 | -1.0004e-11 | 5.8983
9.0949e-13 | 6.2554 | -1.0914e-11 | 6.3044
3.4106e-12 | 6.6584 | -1.0232e-11 | 6.7104
6.8212e-13 | 7.0634 | -1.0687e-11 | 7.1164
2.2737e-13 | 7.4684 | -1.0914e-11 | 7.5224
0 | 7.8715 | -1.1823e-11 | 7.9495
6.8212e-13 | 8.2735 | -1.3188e-11 | 8.3565
0 | 8.6785 | -1.2278e-11 | 8.7625
-2.2737e-13 | 9.0805 | -1.0914e-11 | 9.1685
-6.8212e-13 | 9.4845 | -1.0004e-11 | 9.5745
-1.5916e-12 | 9.8896 | -1.2278e-11 | 9.9806
-1.1369e-12 | 10.2936 | -1.0687e-11 | 10.3846
1.3642e-12 | 10.7206 | -8.1855e-12 | 10.7916
-4.5475e-13 | 11.1266 | -1.1596e-11 | 11.1996
-1.3642e-12 | 11.5327 | -1.296e-11 | 11.6067
1.5916e-12 | 11.9377 | -7.7307e-12 | 12.0137
9.0949e-13 | 12.3427 | -8.4128e-12 | 12.4187
9.0949e-13 | 12.7467 | -1.0459e-11 | 12.8237
2.7285e-12 | 13.1497 | -1.1823e-11 | 13.2298
-6.8212e-13 | 13.5538 | -9.5497e-12 | 13.6368
6.8212e-13 | 13.9578 | -1.0004e-11 | 14.0428
2.0464e-12 | 14.3628 | -1.1596e-11 | 14.4608
1.3642e-12 | 14.7678 | -1.2278e-11 | 14.8659
9.0949e-13 | 15.1739 | -1.3415e-11 | 15.2719
-6.8212e-13 | 15.5789 | -1.3188e-11 | 15.6799
1.5916e-12 | 15.9839 | -9.3223e-12 | 16.0939
9.0949e-13 | 16.3899 | -1.0459e-11 | 16.5149
-1.1369e-12 | 16.793 | -1.387e-11 | 16.921
-2.2737e-13 | 17.198 | -1.1596e-11 | 17.327
1.3642e-12 | 17.602 | -1.2051e-11 | 17.734
2.0464e-12 | 18.007 | -1.2278e-11 | 18.14
-1.1369e-12 | 18.411 | -1.2278e-11 | 18.5441
-6.8212e-13 | 18.8171 | -1.1596e-11 | 18.9511
-1.5916e-12 | 19.2221 | -1.4779e-11 | 19.4091
1.5916e-12 | 19.6261 | -9.5497e-12 | 19.8151
2.2737e-13 | 20.0331 | -1.0687e-11 | 20.2212
1.1369e-12 | 20.4372 | -1.1369e-11 | 20.6282
-2.2737e-13 | 20.8422 | -1.2733e-11 | 21.0332
9.0949e-13 | 21.2472 | -1.296e-11 | 21.4372
2.0464e-12 | 21.6502 | -1.1596e-11 | 21.8422
-2.2737e-13 | 22.0553 | -1.1596e-11 | 22.2503
-6.8212e-13 | 22.4583 | -8.4128e-12 | 22.6573
0 | 22.8633 | -1.1823e-11 | 23.0633
2.2737e-13 | 23.2653 | -1.0232e-11 | 23.4683
1.3642e-12 | 23.6703 | -1.1596e-11 | 23.8744
1.1369e-12 | 24.0754 | -1.0232e-11 | 24.2784
-1.819e-12 | 24.4814 | -1.296e-11 | 24.6904
2.2737e-13 | 24.8864 | -1.0914e-11 | 25.0954
2.7285e-12 | 25.2904 | -1.0004e-11 | 25.5035
1.1369e-12 | 25.6965 | -1.2051e-11 | 25.9085
-6.8212e-13 | 26.1005 | -1.1369e-11 | 26.3265
-2.9559e-12 | 26.5055 | -9.3223e-12 | 26.7375
6.8212e-13 | 26.9105 | -9.5497e-12 | 27.1425
-2.2737e-13 | 27.3156 | -1.0914e-11 | 27.5466
2.0464e-12 | 27.7196 | -1.1369e-11 | 27.9526
-2.2737e-12 | 28.1236 | -1.0004e-11 | 28.3576
-6.8212e-13 | 28.5306 | -1.0004e-11 | 28.9066
1.1369e-12 | 28.9347 | -1.2278e-11 | 29.3117
0 | 29.3397 | -1.1596e-11 | 29.7567
0 | 29.7417 | -1.0232e-11 | 30.1617
-1.3642e-12 | 30.1457 | -1.1596e-11 | 30.7268
-6.8212e-13 | 30.5508 | -1.2278e-11 | 31.1338
0 | 30.9558 | -1.0459e-11 | 31.5398
1.1369e-12 | 31.3588 | -1.1369e-11 | 31.9478
9.0949e-13 | 31.7618 | -1.2051e-11 | 32.3528
-3.638e-12 | 32.1668 | -1.387e-11 | 32.7599
-4.5475e-13 | 32.5719 | -9.5497e-12 | 33.1659
9.0949e-13 | 32.9769 | -9.5497e-12 | 33.5719
6.8212e-13 | 33.3809 | -1.2051e-11 | 33.9919
9.0949e-13 | 33.7839 | -1.0687e-11 | 34.398
-1.3642e-12 | 34.188 | -9.5497e-12 | 34.809
-2.2737e-12 | 34.593 | -1.0004e-11 | 35.215
0 | 34.999 | -1.2051e-11 | 35.62
6.8212e-13 | 35.401 | -1.0914e-11 | 36.0251
2.2737e-13 | 35.8051 | -1.1823e-11 | 36.4361
-1.3642e-12 | 36.2091 | -1.0687e-11 | 36.8421
-1.3642e-12 | 36.6151 | -1.1596e-11 | 37.2481
1.5916e-12 | 37.0201 | -1.0232e-11 | 37.6531
6.8212e-13 | 37.4251 | -1.1369e-11 | 38.0582
-4.5475e-13 | 37.8282 | -1.0687e-11 | 38.4642
0 | 38.2322 | -1.1596e-11 | 38.8692
2.2737e-13 | 38.6352 | -1.0687e-11 | 39.2752
1.3642e-12 | 39.0402 | -9.3223e-12 | 39.6803
9.0949e-13 | 39.4453 | -1.0232e-11 | 40.0863
-1.3642e-12 | 39.8493 | -9.3223e-12 | 40.4903
-1.3642e-12 | 40.2553 | -1.2051e-11 | 40.8943
1.1369e-12 | 40.6603 | -1.1596e-11 | 41.2994
2.0464e-12 | 41.0654 | -1.1596e-11 | 41.7064
2.2737e-13 | 41.4774 | -1.0459e-11 | 42.1124
9.0949e-13 | 41.8814 | -1.0004e-11 | 42.5184
9.0949e-13 | 42.2854 | -1.2278e-11 | 42.9245
2.2737e-13 | 42.6904 | -8.8676e-12 | 43.3305
-6.8212e-13 | 43.0935 | -9.0949e-12 | 43.7365
1.1369e-12 | 43.4985 | -1.1369e-11 | 44.1415
3.4106e-12 | 43.9045 | -1.2733e-11 | 44.5466
2.2737e-13 | 44.3085 | -9.3223e-12 | 44.9536
-6.8212e-13 | 44.7126 | -1.0004e-11 | 45.3596
-1.3642e-12 | 45.1176 | -8.8676e-12 | 45.7646
1.1369e-12 | 45.5226 | -1.2278e-11 | 46.1696
2.0464e-12 | 45.9276 | -7.9581e-12 | 46.5757
9.0949e-13 | 46.3336 | -1.2278e-11 | 46.9797
-6.8212e-13 | 46.7397 | -1.2278e-11 | 47.3847
-2.2737e-13 | 47.1427 | -1.0232e-11 | 47.7907
1.3642e-12 | 47.5467 | -1.296e-11 | 48.1948
6.8212e-13 | 47.9527 | -1.1596e-11 | 48.6008
0 | 48.3578 | -1.2733e-11 | 49.0058
1.3642e-12 | 48.7598 | -1.0687e-11 | 49.4108
-2.2737e-13 | 49.1718 | -1.0687e-11 | 49.8158
-4.5475e-13 | 49.5768 | -9.0949e-12 | 50.2219
0 | 49.9799 | -1.296e-11 | 50.6259
9.0949e-13 | 50.3829 | -1.0687e-11 | 51.0329
1.3642e-12 | 50.7869 | -1.0914e-11 | 51.4379
2.2737e-13 | 51.1929 | -1.0687e-11 | 51.844
0 | 51.5969 | -1.1823e-11 | 52.25
6.8212e-13 | 51.999 | -1.1596e-11 | 52.655
-1.3642e-12 | 52.402 | -1.0459e-11 | 53.061
-6.8212e-13 | 52.806 | -1.1369e-11 | 53.4671
2.0464e-12 | 53.21 | -1.1596e-11 | 53.8731
1.5916e-12 | 53.6161 | -8.8676e-12 | 54.2771
-1.1369e-12 | 54.0201 | -1.3415e-11 | 54.6831
-1.819e-12 | 54.4251 | -9.3223e-12 | 55.0881
0 | 54.8301 | -1.2051e-11 | 55.4922
1.3642e-12 | 55.2322 | -1.1596e-11 | 55.8982
9.0949e-13 | 55.6382 | -1.0232e-11 | 56.3052
-6.8212e-13 | 56.0422 | -1.1596e-11 | 56.7112
-1.5916e-12 | 56.4472 | -1.2278e-11 | 57.1163
1.1369e-12 | 56.8522 | -1.0232e-11 | 57.5203
-2.2737e-13 | 57.2583 | -1.3188e-11 | 57.9263
-6.8212e-13 | 57.6613 | -1.0459e-11 | 58.3333
0 | 58.0663 | -1.0459e-11 | 58.7374
-6.8212e-13 | 58.4693 | -1.1596e-11 | 59.1424
2.0464e-12 | 58.8734 | -1.1596e-11 | 59.5464
2.0464e-12 | 59.2784 | -1.0914e-11 | 59.9504
9.0949e-13 | 59.6834 | -1.0687e-11 | 60.3554
-1.3642e-12 | 60.0874 | -1.0914e-11 | 60.7605
9.0949e-13 | 60.4935 | -1.2278e-11 | 61.1655
1.1369e-12 | 60.8975 | -1.0459e-11 | 61.5705
1.3642e-12 | 61.3035 | -9.3223e-12 | 61.9765
-1.1369e-12 | 61.7085 | -8.4128e-12 | 62.3806
6.8212e-13 | 62.1135 | -1.1823e-11 | 62.7856
1.1369e-12 | 62.5176 | -1.2278e-11 | 63.1906
0 | 62.9216 | -1.2733e-11 | 63.5966
-6.8212e-13 | 63.3236 | -1.0687e-11 | 64.0017
6.8212e-13 | 63.7276 | -1.1823e-11 | 64.4067
1.1369e-12 | 64.1317 | -1.2051e-11 | 64.8117
-1.1369e-12 | 64.5367 | -1.1369e-11 | 65.2177
1.1369e-12 | 64.9407 | -1.0914e-11 | 65.6248
1.3642e-12 | 65.3447 | -1.2733e-11 | 66.0288
-2.7285e-12 | 65.7508 | -1.2733e-11 | 66.4328
-1.3642e-12 | 66.1558 | -1.0232e-11 | 66.8388
2.5011e-12 | 66.5608 | -8.8676e-12 | 67.2438
6.8212e-13 | 66.9658 | -1.3188e-11 | 67.6489
0 | 67.3709 | -1.0459e-11 | 68.0549
-1.819e-12 | 67.7739 | -9.0949e-12 | 68.4629
-2.2737e-13 | 68.1779 | -1.2278e-11 | 68.8669
1.3642e-12 | 68.5819 | -1.0914e-11 | 69.273
-2.2737e-13 | 68.9869 | -1.0004e-11 | 69.678
2.2737e-13 | 69.392 | -1.0459e-11 | 70.085
-2.2737e-13 | 69.796 | -1.0687e-11 | 70.491
9.0949e-13 | 70.201 | -1.1596e-11 | 70.9071
2.2737e-13 | 70.605 | -1.0232e-11 | 71.3131
1.3642e-12 | 71.0091 | -1.0914e-11 | 71.7181
-4.5475e-13 | 71.4151 | -1.0687e-11 | 72.1241
-1.5916e-12 | 71.8191 | -1.387e-11 | 72.5292
6.8212e-13 | 72.2241 | -1.0687e-11 | 72.9722
3.8654e-12 | 72.6272 | -1.1596e-11 | 73.3972
9.0949e-13 | 73.0322 | -1.3188e-11 | 73.8032
-1.1369e-12 | 73.4352 | -1.296e-11 | 74.2082
9.0949e-13 | 73.8392 | -1.2051e-11 | 74.6143
1.5916e-12 | 74.2442 | -1.2733e-11 | 75.0193
1.5916e-12 | 74.6493 | -8.8676e-12 | 75.4253
1.1369e-12 | 75.0533 | -1.0004e-11 | 75.8293
-2.5011e-12 | 75.4583 | -1.296e-11 | 76.2374
-4.5475e-13 | 75.8623 | -1.1823e-11 | 76.6434
1.5916e-12 | 76.2664 | -1.0004e-11 | 77.0474
0 | 76.6704 | -1.2278e-11 | 77.4524
0 | 77.0724 | -1.2733e-11 | 77.8575
0 | 77.4754 | -1.0687e-11 | 78.2625
-1.819e-12 | 77.8774 | -1.0232e-11 | 78.6685
-1.1369e-12 | 78.2825 | -1.1823e-11 | 79.0735
1.3642e-12 | 78.6885 | -1.3415e-11 | 79.4796
6.8212e-13 | 79.0935 | -9.0949e-12 | 79.8856
-2.2737e-12 | 79.4985 | -1.1596e-11 | 80.2906
-1.1369e-12 | 79.9046 | -1.4097e-11 | 80.6966
2.2737e-13 | 80.3096 | -1.1596e-11 | 81.1006
-4.5475e-13 | 80.7146 | -8.1855e-12 | 81.5067
-6.8212e-13 | 81.1196 | -1.2051e-11 | 81.9117
-1.5916e-12 | 81.5237 | -9.3223e-12 | 82.3167
1.3642e-12 | 81.9297 | -1.0687e-11 | 82.7207
1.1369e-12 | 82.3347 | -1.2051e-11 | 83.1248
1.3642e-12 | 82.7397 | -1.0459e-11 | 83.5298
2.2737e-13 | 83.1448 | -1.4552e-11 | 83.9358
|  | -1.4325e-11 | 84.3418
|  | -1.1369e-11 | 84.7478
|  | -8.4128e-12 | 85.1519
|  | -1.0687e-11 | 85.5569
|  | -1.1823e-11 | 85.9629

## 580uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-3.8283e-13 | 1.0057e-12 | -1.5735e-11 | 1.5844e-12
GEMs On |  | GEMs On
-2.2737e-13 | 0.31302 | -1.387e-11 | 0.31102
-6.8212e-13 | 0.99606 | -1.5689e-11 | 0.99506
-2.2737e-13 | 1.4001 | -1.5689e-11 | 1.4011
6.8212e-13 | 1.8051 | -1.7508e-11 | 1.8051
-4.5475e-13 | 2.2091 | -1.3642e-11 | 2.2111
0 | 2.6111 | -1.7053e-11 | 2.6272
1.3642e-12 | 3.0142 | -1.6144e-11 | 3.0332
9.0949e-13 | 3.4192 | -1.5916e-11 | 3.4382
-2.2737e-13 | 3.8242 | -1.8872e-11 | 3.8462
4.5475e-13 | 4.2282 | -1.6371e-11 | 4.2532
4.5475e-13 | 4.6343 | -1.7053e-11 | 4.6573
9.0949e-13 | 5.0373 | -1.3642e-11 | 5.0643
-1.819e-12 | 5.4413 | -1.6371e-11 | 5.4703
-4.5475e-13 | 5.8453 | -1.8645e-11 | 5.8763
-2.2737e-13 | 6.2494 | -1.5916e-11 | 6.2834
-2.2737e-13 | 6.6534 | -1.819e-11 | 6.6884
-4.5475e-13 | 7.0584 | -1.4779e-11 | 7.0944
-3.4106e-12 | 7.4634 | -1.7053e-11 | 7.5324
6.8212e-13 | 7.8685 | -1.7053e-11 | 7.9385
1.1369e-12 | 8.2725 | -1.4779e-11 | 8.3435
-1.1369e-12 | 8.6775 | -1.4779e-11 | 8.7485
-4.5475e-13 | 9.0825 | -1.4552e-11 | 9.1535
-6.8212e-13 | 9.4855 | -1.5007e-11 | 9.5595
-4.5475e-13 | 9.8906 | -1.7735e-11 | 9.9656
6.8212e-13 | 10.2956 | -1.5007e-11 | 10.3706
-2.0464e-12 | 10.6996 | -1.6598e-11 | 10.7766
-1.1369e-12 | 11.1056 | -1.5916e-11 | 11.1826
0 | 11.5107 | -1.387e-11 | 11.5897
4.5475e-13 | 11.9127 | -1.5916e-11 | 11.9957
-2.2737e-13 | 12.3177 | -1.5916e-11 | 12.4017
-2.2737e-13 | 12.7207 | -1.5234e-11 | 12.8247
1.1369e-12 | 13.1248 | -1.387e-11 | 13.2328
-1.3642e-12 | 13.5298 | -1.387e-11 | 13.6408
-1.5916e-12 | 13.9328 | -1.5007e-11 | 14.0608
-6.8212e-13 | 14.3388 | -1.6371e-11 | 14.4758
-4.5475e-13 | 14.7428 | -1.3415e-11 | 14.8828
4.5475e-13 | 15.1479 | -1.5689e-11 | 15.2879
1.3642e-12 | 15.5519 | -1.5234e-11 | 15.6949
2.0464e-12 | 15.9559 | -1.387e-11 | 16.1039
0 | 16.3609 | -1.8645e-11 | 16.5109
-6.8212e-13 | 16.765 | -1.4325e-11 | 16.916
-2.2737e-13 | 17.17 | -1.4779e-11 | 17.321
1.3642e-12 | 17.574 | -1.3415e-11 | 17.726
0 | 17.98 | -1.5689e-11 | 18.132
-1.3642e-12 | 18.382 | -1.8417e-11 | 18.5381
-2.2737e-13 | 18.7881 | -1.6598e-11 | 18.9551
4.5475e-13 | 19.1921 | -1.5916e-11 | 19.3671
-1.819e-12 | 19.5971 | -1.7735e-11 | 19.7721
-1.5916e-12 | 20.0011 | -1.4779e-11 | 20.1781
-6.8212e-13 | 20.4052 | -1.7508e-11 | 20.6192
-2.2737e-13 | 20.8082 | -1.5689e-11 | 21.0252
4.5475e-13 | 21.2122 | -1.5007e-11 | 21.4322
-1.819e-12 | 21.6172 | -1.6371e-11 | 21.8382
-1.819e-12 | 22.0223 | -1.5689e-11 | 22.2463
-4.5475e-13 | 22.4273 | -1.5234e-11 | 22.6523
9.0949e-13 | 22.8313 | -1.3188e-11 | 23.0593
-1.819e-12 | 23.2343 | -1.6371e-11 | 23.4633
-4.5475e-13 | 23.6394 | -1.6598e-11 | 23.8704
-1.3642e-12 | 24.0424 | -1.4552e-11 | 24.2754
2.0464e-12 | 24.4474 | -1.4325e-11 | 24.6814
4.5475e-13 | 24.8524 | -2.0009e-11 | 25.0874
-6.8212e-13 | 25.2564 | -1.4552e-11 | 25.4935
-6.8212e-13 | 25.6605 | -1.8417e-11 | 25.8985
-4.5475e-13 | 26.0655 | -1.5007e-11 | 26.3045
-4.5475e-13 | 26.4685 | -1.7735e-11 | 26.7115
0 | 26.8735 | -1.9099e-11 | 27.1176
-2.0464e-12 | 27.2756 | -1.3642e-11 | 27.5236
-4.5475e-13 | 27.6796 | -1.728e-11 | 27.9306
-4.5475e-13 | 28.0836 | -1.728e-11 | 28.3366
4.5475e-13 | 28.4876 | -1.6371e-11 | 28.7436
-4.5475e-13 | 28.8926 | -1.5234e-11 | 29.1517
-1.5916e-12 | 29.2967 | -1.3188e-11 | 29.6087
1.1369e-12 | 29.7017 | -1.387e-11 | 30.0157
-1.3642e-12 | 30.1067 | -1.6371e-11 | 30.4207
-4.5475e-13 | 30.5117 | -1.6144e-11 | 30.8268
4.5475e-13 | 30.9158 | -1.5689e-11 | 31.2308
-1.5916e-12 | 31.3198 | -1.387e-11 | 31.6378
4.5475e-13 | 31.7238 | -1.5689e-11 | 32.0438
4.5475e-13 | 32.1278 | -1.7053e-11 | 32.4499
0 | 32.5329 | -1.5007e-11 | 32.8569
-1.1369e-12 | 32.9349 | -1.6144e-11 | 33.2629
-2.2737e-13 | 33.3389 | -1.4779e-11 | 33.6709
-1.5916e-12 | 33.7439 | -1.2733e-11 | 34.0759
-2.2737e-13 | 34.1489 | -1.5916e-11 | 34.484
0 | 34.554 | -1.387e-11 | 34.889
6.8212e-13 | 34.956 | -1.5689e-11 | 35.295
-2.2737e-13 | 35.359 | -1.5916e-11 | 35.701
-1.819e-12 | 35.763 | -1.3415e-11 | 36.1071
0 | 36.1681 | -1.5007e-11 | 36.5121
-6.8212e-13 | 36.5701 | -1.5689e-11 | 36.9181
4.5475e-13 | 36.9741 | -1.7735e-11 | 37.3361
-2.0464e-12 | 37.3781 | -1.5007e-11 | 37.7442
-6.8212e-13 | 37.7832 | -1.7735e-11 | 38.1512
-1.5916e-12 | 38.1872 | -1.4779e-11 | 38.5632
-2.2737e-13 | 38.5922 | -1.5234e-11 | 38.9702
-4.5475e-13 | 38.9972 | -1.6371e-11 | 39.3753
-1.1369e-12 | 39.4012 | -1.6598e-11 | 39.7843
0 | 39.8063 | -1.5689e-11 | 40.1903
-4.5475e-13 | 40.2083 | -1.3642e-11 | 40.5953
-1.1369e-12 | 40.6123 | -1.5234e-11 | 41.0014
6.8212e-13 | 41.0174 | -1.3642e-11 | 41.4074
6.8212e-13 | 41.4224 | -1.6371e-11 | 41.8134
0 | 41.8264 | -1.2051e-11 | 42.2194
-1.3642e-12 | 42.2274 | -1.4552e-11 | 42.6254
-4.5475e-13 | 42.6324 | -1.7053e-11 | 43.0315
-4.5475e-13 | 43.0355 | -1.5689e-11 | 43.4415
-4.5475e-13 | 43.4395 | -1.5007e-11 | 43.8465
1.1369e-12 | 43.8435 | -1.5689e-11 | 44.2525
-6.8212e-13 | 44.2485 | -1.5689e-11 | 44.6585
-4.5475e-13 | 44.6525 | -1.7735e-11 | 45.0636
-1.819e-12 | 45.0566 | -1.5234e-11 | 45.4686
-2.2737e-13 | 45.4626 | -1.6144e-11 | 45.8846
-6.8212e-13 | 45.8666 | -1.5234e-11 | 46.2906
9.0949e-13 | 46.2707 | -1.7053e-11 | 46.6997
-6.8212e-13 | 46.6757 | -1.8417e-11 | 47.1067
-6.8212e-13 | 47.0807 | -1.819e-11 | 47.5137
-1.1369e-12 | 47.4847 | -1.819e-11 | 47.9187
-1.819e-12 | 47.8897 | -1.6371e-11 | 48.3248
-6.8212e-13 | 48.2948 | -1.5007e-11 | 48.7308
4.5475e-13 | 48.6988 | -1.819e-11 | 49.1368
0 | 49.1038 | -1.5007e-11 | 49.5418
-1.3642e-12 | 49.5068 | -1.5007e-11 | 49.9459
0 | 49.9128 | -1.8417e-11 | 50.3519
0 | 50.3149 | -1.2506e-11 | 50.7619
9.0949e-13 | 50.7169 | -1.728e-11 | 51.1679
0 | 51.1209 | -1.6371e-11 | 51.574
-6.8212e-13 | 51.523 | -1.6598e-11 | 51.98
-2.2737e-13 | 51.927 | -1.5689e-11 | 52.395
-1.5916e-12 | 52.332 | -1.5916e-11 | 52.805
6.8212e-13 | 52.736 | -1.5916e-11 | 53.21
-1.3642e-12 | 53.139 | -1.6598e-11 | 53.6161
4.5475e-13 | 53.5441 | -1.6144e-11 | 54.0211
2.0464e-12 | 53.9471 | -1.5916e-11 | 54.4271
-4.5475e-13 | 54.3511 | -1.9782e-11 | 54.8351
-1.819e-12 | 54.7561 | -1.6371e-11 | 55.2432
-4.5475e-13 | 55.1611 | -1.7053e-11 | 55.6512
-6.8212e-13 | 55.5642 | -1.6598e-11 | 56.0572
1.3642e-12 | 55.9692 | -1.4552e-11 | 56.4622
0 | 56.3722 | -1.8417e-11 | 56.8672
-1.819e-12 | 56.7743 | -1.3415e-11 | 57.2853
-1.3642e-12 | 57.1783 | -1.7053e-11 | 57.6933
-1.819e-12 | 57.5823 | -1.4779e-11 | 58.1013
-1.3642e-12 | 57.9873 | -1.5689e-11 | 58.5104
-4.5475e-13 | 58.3913 | -1.4325e-11 | 58.9154
-2.2737e-13 | 58.7954 | -1.7508e-11 | 59.3214
-1.3642e-12 | 59.1994 | -1.7053e-11 | 59.7314
-4.5475e-13 | 59.6054 | -1.4779e-11 | 60.1384
-1.3642e-12 | 60.0094 | -1.7053e-11 | 60.5455
-6.8212e-13 | 60.4145 | -1.5234e-11 | 60.9515
-4.5475e-13 | 60.8185 | -1.5689e-11 | 61.3605
-1.5916e-12 | 61.2225 | -1.4779e-11 | 61.7665
-2.0464e-12 | 61.6265 | -1.6598e-11 | 62.1756
1.1369e-12 | 62.0305 | -1.8645e-11 | 62.5876
-4.5475e-13 | 62.4366 | -1.4779e-11 | 63.0036
-2.2737e-13 | 62.8396 | -1.5007e-11 | 63.4086
-1.1369e-12 | 63.2426 | -1.2051e-11 | 63.8137
1.1369e-12 | 63.6496 | -1.5689e-11 | 64.2197
-1.819e-12 | 64.0547 | -1.9554e-11 | 64.6287
1.819e-12 | 64.4587 | -1.5689e-11 | 65.0367
-1.819e-12 | 64.8627 | -1.4779e-11 | 65.4427
6.8212e-13 | 65.2667 | -1.5689e-11 | 65.8488
1.3642e-12 | 65.6728 | -1.6371e-11 | 66.2548
-6.8212e-13 | 66.0768 | -1.387e-11 | 66.6628
-1.1369e-12 | 66.4798 | -1.7508e-11 | 67.0678
-1.1369e-12 | 66.8838 | -1.4552e-11 | 67.4749
-6.8212e-13 | 67.2929 | -1.0914e-11 | 67.8839
-2.5011e-12 | 67.6979 | -1.4325e-11 | 68.2899
-4.5475e-13 | 68.1019 | -1.3415e-11 | 68.6979
-1.5916e-12 | 68.5069 | -1.6371e-11 | 69.1069
6.8212e-13 | 68.9109 | -1.6598e-11 | 69.512
4.5475e-13 | 69.316 | -1.2278e-11 | 69.919
-1.819e-12 | 69.721 | -1.728e-11 | 70.326
1.3642e-12 | 70.127 | -1.5916e-11 | 70.731
4.5475e-13 | 70.531 | -1.5007e-11 | 71.1381
9.0949e-13 | 70.9361 | -1.5007e-11 | 71.5441
-6.8212e-13 | 71.3391 | -1.5007e-11 | 71.9491
1.3642e-12 | 71.7441 | -1.728e-11 | 72.3551
-1.5916e-12 | 72.1511 | -1.6371e-11 | 72.7592
-1.1369e-12 | 72.5551 | -1.6144e-11 | 73.1732
-1.5916e-12 | 72.9602 | -1.6598e-11 | 73.5782
-6.8212e-13 | 73.3642 | -1.6144e-11 | 73.9852
-6.8212e-13 | 73.7682 | -1.5007e-11 | 74.3912
-4.5475e-13 | 74.1722 | -1.6371e-11 | 74.8003
-2.5011e-12 | 74.5763 | -1.6371e-11 | 75.2153
6.8212e-13 | 74.9813 | -1.6144e-11 | 75.6203
-1.5916e-12 | 75.3853 | -1.6598e-11 | 76.0254
9.0949e-13 | 75.7903 | -1.6371e-11 | 76.4334
-2.5011e-12 | 76.1954 | -1.3415e-11 | 76.8394
-1.5916e-12 | 76.5994 | -1.6371e-11 | 77.2514
-4.5475e-13 | 77.0044 | -1.7053e-11 | 77.6564
-1.5916e-12 | 77.4074 | -1.4552e-11 | 78.0615
0 | 77.8114 | -1.4779e-11 | 78.4695
6.8212e-13 | 78.2145 | -1.4325e-11 | 78.8765
-2.2737e-13 | 78.6185 | -1.4325e-11 | 79.2825
9.0949e-13 | 79.0215 | -1.3188e-11 | 79.6906
-1.3642e-12 | 79.4265 | -1.5689e-11 | 80.1016
-2.2737e-13 | 79.8306 | -1.7735e-11 | 80.5066
2.2737e-12 | 80.2336 | -1.4552e-11 | 80.9126
-4.5475e-13 | 80.6386 | -1.2506e-11 | 81.3187
0 | 81.0426 | -1.5234e-11 | 81.7237
-6.8212e-13 | 81.4447 | -1.6371e-11 | 82.1297
4.5475e-13 | 81.8477 | -1.6144e-11 | 82.5357
6.8212e-13 | 82.2527 | -1.3188e-11 | 82.9407
0 | 82.6587 | -1.2733e-11 | 83.3478
-6.8212e-13 | 83.0607 | -1.819e-11 | 83.7528
-2.2737e-13 | 83.4638 | -1.5007e-11 | 84.1588
9.0949e-13 | 83.8678 | -1.6371e-11 | 84.5638
-1.3642e-12 | 84.2718 | -1.819e-11 | 84.9699
-2.2737e-13 | 84.6768
-1.3642e-12 | 85.0809
-2.2737e-13 | 85.4859
9.0949e-13 | 85.8889
-4.5475e-13 | 86.2919
-1.3642e-12 | 86.696
4.5475e-13 | 87.102

## 590uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-6.2351e-13 | 1.0642e-12 | -2.2181e-11 | 2.0045e-12
GEMs On |  | GEMs On
-1.819e-12 | 0.31002 | -1.6826e-11 | 0.31302
-9.0949e-13 | 0.99506 | -1.7508e-11 | 0.99906
1.3642e-12 | 1.3991 | -2.0009e-11 | 1.4041
1.3642e-12 | 1.8031 | -2.0009e-11 | 1.8091
2.2737e-13 | 2.2081 | -2.4784e-11 | 2.2131
-2.0464e-12 | 2.6121 | -2.3647e-11 | 2.6191
1.1369e-12 | 3.0162 | -2.2737e-11 | 3.0252
-1.1369e-12 | 3.4202 | -2.4556e-11 | 3.4302
-9.0949e-13 | 3.8252 | -2.1828e-11 | 3.8352
-9.0949e-13 | 4.2292 | -2.1373e-11 | 4.2402
-1.1369e-12 | 4.6333 | -2.4556e-11 | 4.6463
4.5475e-13 | 5.0373 | -2.3419e-11 | 5.0513
-4.5475e-13 | 5.4413 | -2.0691e-11 | 5.4563
-1.3642e-12 | 5.8463 | -2.3192e-11 | 5.8613
6.8212e-13 | 6.2504 | -2.0918e-11 | 6.2664
-2.7285e-12 | 6.6544 | -2.0236e-11 | 6.6724
-9.0949e-13 | 7.0594 | -1.8645e-11 | 7.0784
-1.1369e-12 | 7.4644 | -2.2283e-11 | 7.4854
4.5475e-13 | 7.8694 | -2.5011e-11 | 7.8905
0 | 8.2745 | -2.0009e-11 | 8.2955
-1.819e-12 | 8.6775 | -2.3647e-11 | 8.7015
2.2737e-13 | 9.0825 | -2.251e-11 | 9.1075
0 | 9.4865 | -2.4102e-11 | 9.5125
-1.1369e-12 | 9.8896 | -2.0691e-11 | 9.9186
-1.3642e-12 | 10.2946 | -2.0691e-11 | 10.3236
0 | 10.7006 | -2.1373e-11 | 10.7296
6.8212e-13 | 11.1026 | -2.3419e-11 | 11.1356
-2.0464e-12 | 11.5077 | -2.6603e-11 | 11.5407
2.2737e-13 | 11.9127 | -2.2055e-11 | 11.9477
0 | 12.3167 | -2.3192e-11 | 12.3517
-6.8212e-13 | 12.7197 | -2.0009e-11 | 12.7577
0 | 13.1237 | -2.3192e-11 | 13.1638
-6.8212e-13 | 13.5258 | -2.3874e-11 | 13.5698
-1.3642e-12 | 13.9318 | -2.251e-11 | 13.9758
-6.8212e-13 | 14.3368 | -1.8645e-11 | 14.3808
-9.0949e-13 | 14.7408 | -2.0009e-11 | 14.7849
-6.8212e-13 | 15.1449 | -2.1828e-11 | 15.1899
6.8212e-13 | 15.5489 | -1.9782e-11 | 15.5959
-4.5475e-13 | 15.9539 | -2.0009e-11 | 16.0019
6.8212e-13 | 16.3569 | -2.0009e-11 | 16.4079
-1.819e-12 | 16.761 | -2.5011e-11 | 16.812
0 | 17.166 | -2.6148e-11 | 17.218
1.5916e-12 | 17.569 | -2.1373e-11 | 17.624
-2.5011e-12 | 17.974 | -2.0236e-11 | 18.029
2.2737e-13 | 18.379 | -1.9782e-11 | 18.434
-1.3642e-12 | 18.7831 | -2.0009e-11 | 18.8401
1.1369e-12 | 19.1871 | -2.2737e-11 | 19.2461
-6.8212e-13 | 19.5921 | -2.5693e-11 | 19.6521
-4.5475e-13 | 19.9961 | -2.1146e-11 | 20.0572
1.1369e-12 | 20.4002 | -2.0009e-11 | 20.4622
-6.8212e-13 | 20.8042 | -1.9554e-11 | 20.8682
-2.0464e-12 | 21.2072 | -2.1828e-11 | 21.2732
0 | 21.6132 | -2.251e-11 | 21.6782
-4.5475e-13 | 22.0153 | -2.0918e-11 | 22.0843
4.5475e-13 | 22.4203 | -2.3419e-11 | 22.4903
1.3642e-12 | 22.8243 | -2.3192e-11 | 22.8953
-1.819e-12 | 23.2293 | -2.3874e-11 | 23.3003
-2.2737e-12 | 23.6344 | -2.3419e-11 | 23.7044
0 | 24.0394 | -2.0691e-11 | 24.1114
2.2737e-13 | 24.4414 | -2.3647e-11 | 24.5174
0 | 24.8444 | -2.0009e-11 | 24.9244
2.2737e-13 | 25.2474 | -2.1828e-11 | 25.3295
-4.5475e-13 | 25.6505 | -2.0009e-11 | 25.7345
2.2737e-13 | 26.0535 | -2.2055e-11 | 26.1405
-4.5475e-13 | 26.4585 | -2.2283e-11 | 26.5465
-1.3642e-12 | 26.8635 | -2.1373e-11 | 26.9525
1.5916e-12 | 27.2676 | -2.2737e-11 | 27.3576
-1.3642e-12 | 27.6716 | -2.2055e-11 | 27.7636
0 | 28.0756 | -2.2737e-11 | 28.1686
1.5916e-12 | 28.4776 | -2.2737e-11 | 28.5746
-1.1369e-12 | 28.8827 | -2.3419e-11 | 28.9807
-9.0949e-13 | 29.2867 | -2.2737e-11 | 29.3857
-9.0949e-13 | 29.6907 | -2.1828e-11 | 29.7907
-4.5475e-13 | 30.0947 | -2.4784e-11 | 30.1967
-1.819e-12 | 30.4987 | -2.1828e-11 | 30.6028
-2.0464e-12 | 30.9038 | -2.0236e-11 | 31.0098
2.2737e-13 | 31.3088 | -2.3647e-11 | 31.4158
-4.5475e-13 | 31.7118 | -2.1828e-11 | 31.8208
-6.8212e-13 | 32.1148 | -2.4556e-11 | 32.2258
2.2737e-13 | 32.5189 | -2.4556e-11 | 32.6329
-1.3642e-12 | 32.9219 | -1.9782e-11 | 33.0379
6.8212e-13 | 33.3269 | -2.5011e-11 | 33.4429
-4.5475e-13 | 33.7309 | -2.2055e-11 | 33.8489
-4.5475e-13 | 34.136 | -1.8645e-11 | 34.255
-1.819e-12 | 34.539 | -2.1373e-11 | 34.66
-4.5475e-13 | 34.945 | -2.7967e-11 | 35.066
-4.5475e-13 | 35.348 | -1.8872e-11 | 35.472
-2.0464e-12 | 35.751 | -1.9327e-11 | 35.877
-1.819e-12 | 36.1561 | -2.5693e-11 | 36.2821
-1.1369e-12 | 36.5601 | -2.4102e-11 | 36.6881
-9.0949e-13 | 36.9661 | -2.3419e-11 | 37.0941
1.5916e-12 | 37.3691 | -2.3192e-11 | 37.5001
-1.3642e-12 | 37.7752 | -2.5238e-11 | 37.9052
-1.819e-12 | 38.1802 | -2.2283e-11 | 38.3112
0 | 38.5862 | -2.3419e-11 | 38.7162
-4.5475e-13 | 38.9882 | -2.2283e-11 | 39.1202
6.8212e-13 | 39.3922 | -2.1146e-11 | 39.5263
-2.2737e-12 | 39.7963 | -2.3647e-11 | 39.9313
-9.0949e-13 | 40.2003 | -2.1146e-11 | 40.3363
6.8212e-13 | 40.6053 | -2.0918e-11 | 40.7423
2.0464e-12 | 41.0093 | -2.3647e-11 | 41.1474
-1.3642e-12 | 41.4144 | -2.2737e-11 | 41.5514
-2.0464e-12 | 41.8184 | -2.251e-11 | 41.9574
-2.0464e-12 | 42.2244 | -2.2737e-11 | 42.3634
-1.3642e-12 | 42.6294 | -2.1828e-11 | 42.7685
-1.3642e-12 | 43.0335 | -1.7963e-11 | 43.1745
-1.1369e-12 | 43.4385 | -2.0918e-11 | 43.5815
-6.8212e-13 | 43.8435 | -2.2283e-11 | 43.9875
1.1369e-12 | 44.2485 | -2.3192e-11 | 44.3935
-9.0949e-13 | 44.6536 | -1.9782e-11 | 44.7996
-2.0464e-12 | 45.0576 | -2.3647e-11 | 45.2046
2.2737e-13 | 45.4616 | -2.251e-11 | 45.6096
-6.8212e-13 | 45.8656 | -2.2283e-11 | 46.0146
1.3642e-12 | 46.2696 | -2.251e-11 | 46.4207
-9.0949e-13 | 46.6757 | -2.3192e-11 | 46.8257
-1.1369e-12 | 47.0797 | -2.0691e-11 | 47.2317
0 | 47.4847 | -2.0009e-11 | 47.6377
-1.3642e-12 | 47.8877 | -2.251e-11 | 48.0427
2.2737e-13 | 48.2928 | -2.2737e-11 | 48.4468
-4.5475e-13 | 48.6958 | -2.1146e-11 | 48.8518
-2.7285e-12 | 49.1008 | -1.9782e-11 | 49.2568
0 | 49.5058 | -2.6148e-11 | 49.6618
-9.0949e-13 | 49.9089 | -2.3874e-11 | 50.0669
6.8212e-13 | 50.3139 | -2.6603e-11 | 50.4719
0 | 50.7179 | -1.9327e-11 | 50.8779
-2.2737e-12 | 51.1199 | -2.251e-11 | 51.2829
-3.4106e-12 | 51.5249 | -2.1146e-11 | 51.688
-4.5475e-13 | 51.93 | -2.1373e-11 | 52.093
2.2737e-13 | 52.333 | -2.2283e-11 | 52.499
-6.8212e-13 | 52.737 | -2.1373e-11 | 52.904
-2.2737e-12 | 53.142 | -2.3647e-11 | 53.31
-2.2737e-12 | 53.5461 | -2.2737e-11 | 53.7161
-9.0949e-13 | 53.9511 | -2.251e-11 | 54.1211
0 | 54.3561 | -2.0918e-11 | 54.5251
-2.0464e-12 | 54.7591 | -2.251e-11 | 54.9311
-3.1832e-12 | 55.1651 | -2.3874e-11 | 55.3372
1.819e-12 | 55.5722 | -2.251e-11 | 55.7432
-6.8212e-13 | 55.9782 | -2.4784e-11 | 56.1482
2.2737e-13 | 56.3832 | -2.1828e-11 | 56.5532
2.2737e-13 | 56.7873 | -2.1373e-11 | 56.9593
0 | 57.1913 | -2.1373e-11 | 57.3653
6.8212e-13 | 57.5933 | -2.251e-11 | 57.7713
-3.4106e-12 | 57.9973 | -2.4784e-11 | 58.1773
-2.2737e-12 | 58.4013 | -2.2737e-11 | 58.5823
2.2737e-13 | 58.8064 | -2.2283e-11 | 58.9874
-1.1369e-12 | 59.2094 | -2.2283e-11 | 59.3934
-6.8212e-13 | 59.6144 | -2.2737e-11 | 59.7994
-1.819e-12 | 60.0204 | -1.8645e-11 | 60.2044
-4.5475e-13 | 60.4245 | -2.1373e-11 | 60.6085
-2.0464e-12 | 60.8285 | -2.2283e-11 | 61.0145
-4.5475e-13 | 61.2335 | -2.2737e-11 | 61.4195
-9.0949e-13 | 61.6385 | -2.1828e-11 | 61.8255
-1.1369e-12 | 62.0436 | -2.3419e-11 | 62.2316
0 | 62.4476 | -2.2283e-11 | 62.6386
0 | 62.8496 | -1.819e-11 | 63.0446
2.2737e-13 | 63.2526 | -2.0236e-11 | 63.4506
0 | 63.6556 | -2.4102e-11 | 63.8546
-6.8212e-13 | 64.0597 | -2.3192e-11 | 64.2617
-9.0949e-13 | 64.4637 | -2.4784e-11 | 64.6657
-2.5011e-12 | 64.8677 | -2.0918e-11 | 65.0707
-4.5475e-13 | 65.2717 | -2.2283e-11 | 65.4758
2.2737e-13 | 65.6758 | -2.2283e-11 | 65.8808
0 | 66.0808 | -2.0009e-11 | 66.2868
2.2737e-13 | 66.4828 | -1.9782e-11 | 66.6938
2.2737e-13 | 66.8868 | -2.3874e-11 | 67.0988
-1.3642e-12 | 67.2909 | -1.8872e-11 | 67.5039
0 | 67.6959 | -2.3647e-11 | 67.9079
-1.1369e-12 | 68.0989 | -2.2737e-11 | 68.3159
-1.1369e-12 | 68.5039 | -2.8649e-11 | 68.7199
-1.1369e-12 | 68.9089 | -2.3192e-11 | 69.1269
-2.0464e-12 | 69.314 | -2.1373e-11 | 69.531
-9.0949e-13 | 69.743 | -2.1146e-11 | 69.937
0 | 70.156 | -2.0918e-11 | 70.343
-6.8212e-13 | 70.56 | -1.9782e-11 | 70.749
-1.3642e-12 | 70.9641 | -2.1373e-11 | 71.1541
-2.2737e-12 | 71.3681 | -2.0691e-11 | 71.5581
-2.7285e-12 | 71.7731 | -2.1146e-11 | 71.9641
1.3642e-12 | 72.1781 | -1.819e-11 | 72.3691
-1.819e-12 | 72.5832 | -2.7512e-11 | 72.7762
-9.0949e-13 | 72.9872 | -2.2737e-11 | 73.1802
0 | 73.3912 | -2.3647e-11 | 73.5842
2.2737e-13 | 73.7942 | -2.2283e-11 | 73.9892
6.8212e-13 | 74.1982 | -2.251e-11 | 74.3953
2.2737e-13 | 74.6023 | -2.2283e-11 | 74.8023
-6.8212e-13 | 75.0073 | -2.4784e-11 | 75.2073
-1.819e-12 | 75.4103 | -1.728e-11 | 75.6123
6.8212e-13 | 75.8153 | -2.3192e-11 | 76.0173
-1.819e-12 | 76.2194 | -2.2055e-11 | 76.4224
0 | 76.6254 | -2.0691e-11 | 76.8284
0 | 77.0274 | -1.8417e-11 | 77.2344
-6.8212e-13 | 77.4474 | -2.5921e-11 | 77.6394
-2.2737e-12 | 77.8684 | -2.3419e-11 | 78.0445
-4.5475e-13 | 78.2735 | -1.9327e-11 | 78.4495
1.3642e-12 | 78.6775 | -2.5238e-11 | 78.8555
-9.0949e-13 | 79.0825 | -2.1828e-11 | 79.2595
-6.8212e-13 | 79.4856 | -2.3647e-11 | 79.6666
-1.1369e-12 | 79.8896 | -2.2737e-11 | 80.0736
2.2737e-13 | 80.3316 | -2.1373e-11 | 80.4776
-9.0949e-13 | 80.7426 | -2.4556e-11 | 80.8836
-6.8212e-13 | 81.1466 | -2.0236e-11 | 81.2887
-9.0949e-13 | 81.5507 | -2.3192e-11 | 81.6947
-4.5475e-13 | 81.9657 | -2.3419e-11 | 82.1017
0 | 82.3957 | -2.1828e-11 | 82.5057
-6.8212e-13 | 82.8017 | -2.4102e-11 | 82.9107
-4.5475e-13 | 83.2128 | -2.3874e-11 | 83.3178
1.3642e-12 | 83.6368
-2.2737e-12 | 84.0428
-9.0949e-13 | 84.4488
-6.8212e-13 | 84.8548
-2.0464e-12 | 85.2979
2.2737e-13 | 85.7309
-2.0464e-12 | 86.1409
-1.3642e-12 | 86.564
-6.8212e-13 | 86.975
-6.8212e-13 | 87.378
0 | 87.825
-1.1369e-12 | 88.269
0 | 88.6731
-6.8212e-13 | 89.0761
-1.1369e-12 | 89.4811
-1.1369e-12 | 89.8921
0 | 90.3042
-2.5011e-12 | 90.7142
-6.8212e-13 | 91.1192
6.8212e-13 | 91.5232

## 600uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.2445e-12 | 1.1404e-12 | -3.0576e-11 | 2.4799e-12
GEMs On |  | GEMs On
-1.5916e-12 | 0.31302 | -2.8422e-11 | 0.31102
-2.5011e-12 | 0.99706 | -3.115e-11 | 0.99506
-2.2737e-13 | 1.4021 | -2.6148e-11 | 1.4011
9.0949e-13 | 1.8061 | -3.5016e-11 | 1.8061
0 | 2.2091 | -3.0695e-11 | 2.2121
-1.1369e-12 | 2.6132 | -2.9559e-11 | 2.6171
-9.0949e-13 | 3.0182 | -3.0013e-11 | 3.0222
-2.5011e-12 | 3.4212 | -3.1378e-11 | 3.4262
2.2737e-13 | 3.8252 | -3.4561e-11 | 3.8312
1.5916e-12 | 4.2302 | -2.4784e-11 | 4.2372
-2.2737e-12 | 4.6353 | -3.1378e-11 | 4.6433
-9.0949e-13 | 5.0403 | -3.0695e-11 | 5.0483
2.2737e-13 | 5.4443 | -3.1832e-11 | 5.4533
0 | 5.8483 | -3.0695e-11 | 5.8583
1.5916e-12 | 6.2504 | -3.4106e-11 | 6.2654
-2.2737e-12 | 6.6554 | -2.9331e-11 | 6.6714
-1.5916e-12 | 7.0594 | -3.2514e-11 | 7.0774
2.2737e-13 | 7.4644 | -2.683e-11 | 7.4824
0 | 7.8694 | -2.8649e-11 | 7.8875
-4.5475e-13 | 8.2735 | -2.7512e-11 | 8.2935
0 | 8.6775 | -3.2514e-11 | 8.6995
-2.9559e-12 | 9.0795 | -3.3879e-11 | 9.1045
-1.3642e-12 | 9.4845 | -3.3424e-11 | 9.5115
-2.7285e-12 | 9.8896 | -2.6375e-11 | 9.9166
-2.2737e-12 | 10.2936 | -2.9559e-11 | 10.3216
-1.5916e-12 | 10.6976 | -3.3651e-11 | 10.7266
-1.5916e-12 | 11.1016 | -2.9559e-11 | 11.1326
-2.2737e-13 | 11.5057 | -2.7285e-11 | 11.5397
-1.1369e-12 | 11.9107 | -2.8422e-11 | 11.9447
-1.1369e-12 | 12.3157 | -3.206e-11 | 12.3497
-1.819e-12 | 12.7207 | -3.0695e-11 | 12.7557
-9.0949e-13 | 13.1267 | -2.9331e-11 | 13.1608
2.2737e-13 | 13.5318 | -2.9331e-11 | 13.5678
-4.5475e-13 | 13.9358 | -2.9331e-11 | 13.9728
-2.2737e-13 | 14.3398 | -2.8422e-11 | 14.3778
-1.1369e-12 | 14.7438 | -3.1832e-11 | 14.7838
-1.5916e-12 | 15.1499 | -3.206e-11 | 15.1879
-1.3642e-12 | 15.5539 | -3.206e-11 | 15.5929
-1.1369e-12 | 15.9589 | -2.8876e-11 | 15.9989
-2.2737e-13 | 16.3629 | -2.8194e-11 | 16.4029
-1.1369e-12 | 16.767 | -3.3651e-11 | 16.809
-2.5011e-12 | 17.171 | -2.5693e-11 | 17.215
-1.5916e-12 | 17.576 | -2.8422e-11 | 17.621
-2.7285e-12 | 17.982 | -3.2287e-11 | 18.026
-1.5916e-12 | 18.388 | -2.8876e-11 | 18.4321
-2.2737e-12 | 18.7941 | -2.9786e-11 | 18.8391
-1.1369e-12 | 19.2001 | -3.3879e-11 | 19.2441
-3.4106e-12 | 19.6051 | -3.2742e-11 | 19.6501
-9.0949e-13 | 20.0111 | -3.1378e-11 | 20.0552
-1.1369e-12 | 20.4142 | -2.7512e-11 | 20.4602
-1.3642e-12 | 20.8192 | -2.8876e-11 | 20.8652
-1.1369e-12 | 21.2242 | -3.7062e-11 | 21.2702
-1.1369e-12 | 21.6272 | -3.0695e-11 | 21.6752
0 | 22.0333 | -3.0695e-11 | 22.0803
2.2737e-13 | 22.4343 | -2.5693e-11 | 22.4853
-1.819e-12 | 22.8383 | -3.1832e-11 | 22.8903
-2.7285e-12 | 23.2433 | -3.3879e-11 | 23.2953
-2.9559e-12 | 23.6473 | -3.1832e-11 | 23.7004
-4.5475e-13 | 24.0524 | -3.115e-11 | 24.1064
-2.2737e-13 | 24.4574 | -2.9331e-11 | 24.5114
-9.0949e-13 | 24.8604 | -2.9559e-11 | 24.9174
-2.5011e-12 | 25.2655 | -2.8649e-11 | 25.3224
-9.0949e-13 | 25.6705 | -2.9786e-11 | 25.7285
-9.0949e-13 | 26.0755 | -3.0923e-11 | 26.1325
-2.2737e-12 | 26.4795 | -3.3879e-11 | 26.5395
-1.819e-12 | 26.8855 | -3.2742e-11 | 26.9445
-1.3642e-12 | 27.2896 | -3.4106e-11 | 27.3496
4.5475e-13 | 27.6946 | -2.8194e-11 | 27.7556
4.5475e-13 | 28.0976 | -2.9559e-11 | 28.1616
-1.1369e-12 | 28.5006 | -3.3424e-11 | 28.5666
-1.819e-12 | 28.9057 | -3.4106e-11 | 28.9717
-1.1369e-12 | 29.3097 | -3.0013e-11 | 29.3767
2.2737e-13 | 29.7157 | -2.9786e-11 | 29.7817
0 | 30.1187 | -3.0013e-11 | 30.1867
-1.3642e-12 | 30.5207 | -3.0241e-11 | 30.5918
-4.5475e-13 | 30.9258 | -2.8422e-11 | 30.9988
-1.819e-12 | 31.3308 | -2.683e-11 | 31.4028
2.2737e-13 | 31.7358 | -2.8194e-11 | 31.8078
0 | 32.1398 | -3.1378e-11 | 32.2148
-1.5916e-12 | 32.5429 | -3.0013e-11 | 32.6219
-3.4106e-12 | 32.9489 | -2.6148e-11 | 33.0269
-2.7285e-12 | 33.3539 | -3.115e-11 | 33.4319
4.5475e-13 | 33.7579 | -3.115e-11 | 33.8379
2.2737e-13 | 34.1619 | -2.8649e-11 | 34.245
-1.1369e-12 | 34.567 | -3.115e-11 | 34.652
-1.3642e-12 | 34.972 | -3.3197e-11 | 35.058
2.2737e-13 | 35.377 | -2.683e-11 | 35.462
0 | 35.7801 | -3.0013e-11 | 35.868
-3.638e-12 | 36.1841 | -3.1378e-11 | 36.2731
-9.0949e-13 | 36.5891 | -2.9559e-11 | 36.6781
-1.1369e-12 | 36.9921 | -2.5693e-11 | 37.0851
4.5475e-13 | 37.3971 | -3.0241e-11 | 37.4901
9.0949e-13 | 37.8012 | -3.206e-11 | 37.8962
-4.5475e-13 | 38.2052 | -2.7285e-11 | 38.3012
-2.5011e-12 | 38.6092 | -2.9559e-11 | 38.7072
-1.819e-12 | 39.0132 | -2.9786e-11 | 39.1122
-1.1369e-12 | 39.4192 | -3.0013e-11 | 39.5183
0 | 39.8243 | -3.2742e-11 | 39.9233
-1.3642e-12 | 40.2263 | -2.7967e-11 | 40.3273
-2.5011e-12 | 40.6303 | -3.1832e-11 | 40.7333
-9.0949e-13 | 41.0354 | -3.5016e-11 | 41.1384
-1.819e-12 | 41.4394 | -3.5016e-11 | 41.5444
-1.5916e-12 | 41.8434 | -2.8194e-11 | 41.9484
-1.819e-12 | 42.2494 | -2.7967e-11 | 42.3534
-2.2737e-12 | 42.6534 | -3.3879e-11 | 42.7595
-1.3642e-12 | 43.0595 | -3.1832e-11 | 43.1645
2.2737e-13 | 43.4655 | -3.2514e-11 | 43.5695
-1.819e-12 | 43.8685 | -3.2287e-11 | 43.9755
-1.819e-12 | 44.2735 | -3.0013e-11 | 44.3805
-2.2737e-12 | 44.6786 | -3.0241e-11 | 44.7856
9.0949e-13 | 45.0826 | -2.6148e-11 | 45.1916
-2.5011e-12 | 45.4866 | -2.7967e-11 | 45.5956
-4.5475e-13 | 45.8916 | -3.0695e-11 | 46.0006
-2.2737e-12 | 46.2966 | -2.9331e-11 | 46.4057
-9.0949e-13 | 46.7027 | -2.9331e-11 | 46.8097
-1.1369e-12 | 47.1057 | -3.2514e-11 | 47.2147
-2.5011e-12 | 47.5117 | -2.3419e-11 | 47.6197
-1.3642e-12 | 47.9147 | -3.3197e-11 | 48.0258
-4.5475e-13 | 48.3208 | -2.8194e-11 | 48.4318
-1.819e-12 | 48.7238 | -3.5016e-11 | 48.8378
-1.819e-12 | 49.1308 | -3.3197e-11 | 49.2438
-2.5011e-12 | 49.5358 | -3.1832e-11 | 49.6498
-1.5916e-12 | 49.9409 | -2.9559e-11 | 50.0549
-2.2737e-13 | 50.3459 | -2.7512e-11 | 50.4609
-9.0949e-13 | 50.7499 | -3.115e-11 | 50.8679
-3.638e-12 | 51.1539 | -2.8876e-11 | 51.2729
-2.2737e-13 | 51.559 | -3.2742e-11 | 51.678
-4.5475e-13 | 51.962 | -3.5698e-11 | 52.082
2.2737e-13 | 52.366 | -2.8649e-11 | 52.488
9.0949e-13 | 52.772 | -3.2514e-11 | 52.893
2.2737e-13 | 53.176 | -2.7967e-11 | 53.2991
-2.5011e-12 | 53.5811 | -3.3651e-11 | 53.7041
-1.3642e-12 | 53.9841 | -3.0695e-11 | 54.1091
-1.5916e-12 | 54.3891 | -2.9559e-11 | 54.5141
-1.819e-12 | 54.7931 | -3.1832e-11 | 54.9211
-1.819e-12 | 55.1982 | -2.5921e-11 | 55.3262
-1.819e-12 | 55.6042 | -3.0695e-11 | 55.7322
-2.9559e-12 | 56.0092 | -3.1378e-11 | 56.1382
-2.2737e-12 | 56.4142 | -3.0241e-11 | 56.5432
-2.7285e-12 | 56.8192 | -3.1832e-11 | 56.9473
-9.0949e-13 | 57.2253 | -2.7285e-11 | 57.3533
-2.2737e-13 | 57.6283 | -3.3197e-11 | 57.7603
-2.5011e-12 | 58.0323 | -2.6375e-11 | 58.1663
-1.1369e-12 | 58.4363 | -3.115e-11 | 58.5724
-1.1369e-12 | 58.8414 | -2.8194e-11 | 58.9774
-1.3642e-12 | 59.2474 | -2.9786e-11 | 59.3834
1.5916e-12 | 59.6514 | -3.3197e-11 | 59.7884
-1.5916e-12 | 60.0554 | -3.0923e-11 | 60.1944
-2.2737e-13 | 60.4605 | -2.7285e-11 | 60.5995
-1.3642e-12 | 60.8645 | -3.4788e-11 | 61.0065
-1.1369e-12 | 61.2675 | -2.8649e-11 | 61.4115
-1.5916e-12 | 61.6725 | -3.0241e-11 | 61.8175
-2.9559e-12 | 62.0776 | -3.1832e-11 | 62.2236
-1.3642e-12 | 62.4826 | -3.206e-11 | 62.6306
-1.3642e-12 | 62.8876 | -3.2742e-11 | 63.0356
-4.5475e-13 | 63.2926 | -3.1832e-11 | 63.4426
-2.7285e-12 | 63.6966 | -3.2514e-11 | 63.8486
-3.4106e-12 | 64.1007 | -2.7285e-11 | 64.2537
-9.0949e-13 | 64.5047 | -3.7062e-11 | 64.6597
-2.2737e-13 | 64.9087 | -2.7057e-11 | 65.0647
-1.5916e-12 | 65.3137 | -2.9786e-11 | 65.4707
-2.5011e-12 | 65.7188 | -2.5011e-11 | 65.8768
-1.3642e-12 | 66.1238 | -2.8876e-11 | 66.2828
4.5475e-13 | 66.5288 | -3.1832e-11 | 66.6888
-4.3201e-12 | 66.9328 | -3.1378e-11 | 67.0948
-2.2737e-12 | 67.3368 | -2.9786e-11 | 67.4999
-1.3642e-12 | 67.7409 | -3.5925e-11 | 67.9059
-2.2737e-12 | 68.1459 | -2.8876e-11 | 68.3109
-4.5475e-13 | 68.5499 | -3.1378e-11 | 68.7169
-9.0949e-13 | 68.9529 | -2.8422e-11 | 69.1209
-3.638e-12 | 69.358 | -3.5925e-11 | 69.527
-4.5475e-13 | 69.763 | -3.2287e-11 | 69.933
-3.4106e-12 | 70.167 | -3.115e-11 | 70.338
-2.2737e-13 | 70.572 | -3.0695e-11 | 70.7441
-9.0949e-13 | 70.9751 | -2.9786e-11 | 71.1491
9.0949e-13 | 71.3801 | -3.115e-11 | 71.5551
-1.3642e-12 | 71.7841 | -3.1378e-11 | 71.9621
-1.5916e-12 | 72.1891 | -2.7512e-11 | 72.3671
-9.0949e-13 | 72.5931 | -2.8649e-11 | 72.7732
-1.1369e-12 | 72.9982 | -3.1832e-11 | 73.1772
-1.5916e-12 | 73.4032 | -3.3424e-11 | 73.5842
-1.1369e-12 | 73.8072 | -3.3197e-11 | 73.9892
-2.7285e-12 | 74.2122 | -3.1378e-11 | 74.3953
2.2737e-13 | 74.6163 | -3.2514e-11 | 74.8013
-2.2737e-13 | 75.0203 | -3.5243e-11 | 75.2073
-1.3642e-12 | 75.4233 | -3.3197e-11 | 75.6123
-1.1369e-12 | 75.8273 | -3.2287e-11 | 76.0183
-4.0927e-12 | 76.2324 | -3.2514e-11 | 76.4244
0 | 76.6384 | -3.1378e-11 | 76.8304
-1.819e-12 | 77.0404 | -2.8876e-11 | 77.2354
-1.819e-12 | 77.4474 | -3.2514e-11 | 77.6414
9.0949e-13 | 77.8534 | -3.115e-11 | 78.0475
-2.2737e-12 | 78.2565 | -3.115e-11 | 78.4525
-9.0949e-13 | 78.6605 | -3.0923e-11 | 78.8575
-3.4106e-12 | 79.0645 | -2.8194e-11 | 79.2625
-1.5916e-12 | 79.4695 | -2.9559e-11 | 79.6686
-1.1369e-12 | 79.8746 | -3.0923e-11 | 80.0746
-2.2737e-12 | 80.2796 | -3.2287e-11 | 80.4796
-2.5011e-12 | 80.6846 | -3.3879e-11 | 80.8836
2.2737e-13 | 81.0906 | -3.2287e-11 | 81.2896
-1.1369e-12 | 81.4947 | -3.0013e-11 | 81.6957
-2.9559e-12 | 81.8997 | -3.115e-11 | 82.1007
-4.5475e-13 | 82.3047 | -2.9786e-11 | 82.5057
-1.819e-12 | 82.7097 | -2.8876e-11 | 82.9117
-2.2737e-12 | 83.1148 | -2.5011e-11 | 83.3178
-9.0949e-13 | 83.5188 | -2.8422e-11 | 83.7248
-2.5011e-12 | 83.9228 | -3.1832e-11 | 84.1308
|  | -3.115e-11 | 84.5368
|  | -3.206e-11 | 84.9409
|  | -2.8876e-11 | 85.3469

## 610uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.4536e-12 | 1.2518e-12 | -4.1791e-11 | 3.0499e-12
GEMs On |  | GEMs On
-4.5475e-13 | 0.31402 | -3.979e-11 | 0.31002
-3.4106e-12 | 1.0001 | -4.3201e-11 | 0.99506
-1.1369e-12 | 1.4051 | -4.07e-11 | 1.4001
-1.5916e-12 | 1.8101 | -4.2064e-11 | 1.8071
-2.2737e-13 | 2.2151 | -3.9336e-11 | 2.2161
-2.2737e-13 | 2.6191 | -3.9563e-11 | 2.6221
0 | 3.0232 | -4.0927e-11 | 3.0272
-2.2737e-12 | 3.4272 | -4.6157e-11 | 3.4382
-1.3642e-12 | 3.8342 | -3.8426e-11 | 3.8452
-4.5475e-13 | 4.2382 | -4.7976e-11 | 4.2492
-4.5475e-13 | 4.6443 | -4.8658e-11 | 4.6553
-4.3201e-12 | 5.0493 | -4.07e-11 | 5.0743
1.1369e-12 | 5.4543 | -4.2291e-11 | 5.4803
2.2737e-13 | 5.8623 | -3.6152e-11 | 5.8853
0 | 6.2684 | -3.7744e-11 | 6.2914
2.2737e-13 | 6.6724 | -3.9563e-11 | 6.6964
-1.5916e-12 | 7.0774 | -4.4338e-11 | 7.1024
-1.1369e-12 | 7.4854 | -4.0473e-11 | 7.5094
4.5475e-13 | 7.8905 | -4.4338e-11 | 7.9155
-1.3642e-12 | 8.2945 | -4.2064e-11 | 8.3225
-2.2737e-13 | 8.6995 | -4.934e-11 | 8.7285
-1.1369e-12 | 9.1045 | -4.1837e-11 | 9.1335
-2.5011e-12 | 9.5105 | -4.2974e-11 | 9.5395
-1.3642e-12 | 9.9196 | -4.1609e-11 | 9.9456
-1.1369e-12 | 10.3336 | -4.4338e-11 | 10.3506
-1.819e-12 | 10.7376 | -4.5702e-11 | 10.7556
-4.5475e-13 | 11.1446 | -3.9336e-11 | 11.1766
-2.7285e-12 | 11.5487 | -4.2974e-11 | 11.5837
0 | 11.9547 | -3.9563e-11 | 11.9987
-1.1369e-12 | 12.3807 | -4.4793e-11 | 12.4067
-1.3642e-12 | 12.7877 | -3.9336e-11 | 12.8137
-2.2737e-13 | 13.1937 | -4.0245e-11 | 13.2198
2.2737e-13 | 13.6018 | -4.1155e-11 | 13.6258
-1.819e-12 | 14.0048 | -3.7289e-11 | 14.0338
-1.5916e-12 | 14.5198 | -4.5475e-11 | 14.4388
-2.2737e-12 | 14.9319 | -4.8203e-11 | 14.8619
-2.2737e-13 | 15.3779 | -4.3201e-11 | 15.2669
-1.5916e-12 | 15.8039 | -4.4793e-11 | 15.6759
-4.5475e-13 | 16.2389 | -4.3201e-11 | 16.0819
-1.1369e-12 | 16.6429 | -4.4565e-11 | 16.4869
-1.3642e-12 | 17.049 | -3.9108e-11 | 16.892
-3.4106e-12 | 17.453 | -4.1155e-11 | 17.304
-1.5916e-12 | 17.858 | -4.4565e-11 | 17.727
-9.0949e-13 | 18.264 | -3.8199e-11 | 18.134
-1.819e-12 | 18.6681 | -4.2519e-11 | 18.5411
-1.3642e-12 | 19.0721 | -4.0245e-11 | 18.9471
-2.2737e-12 | 19.4761 | -4.1155e-11 | 19.3521
-2.2737e-12 | 19.8801 | -4.1609e-11 | 19.7591
1.1369e-12 | 20.2842 | -3.8199e-11 | 20.1641
-4.5475e-13 | 20.6882 | -4.0927e-11 | 20.5802
-2.5011e-12 | 21.0922 | -4.0473e-11 | 20.9872
-1.3642e-12 | 21.4982 | -3.7744e-11 | 21.3942
-1.819e-12 | 21.9052 | -3.979e-11 | 21.8002
-2.2737e-12 | 22.3163 | -4.5702e-11 | 22.2073
-1.5916e-12 | 22.7233 | -4.07e-11 | 22.6123
2.2737e-13 | 23.1283 | -4.4338e-11 | 23.0183
-1.5916e-12 | 23.5333 | -4.0927e-11 | 23.4223
-1.3642e-12 | 23.9384 | -4.1837e-11 | 23.8294
-9.0949e-13 | 24.3434 | -4.3201e-11 | 24.2554
0 | 24.7474 | -4.5929e-11 | 24.6624
-4.5475e-13 | 25.1654 | -3.8654e-11 | 25.0714
-1.1369e-12 | 25.5705 | -4.1609e-11 | 25.4765
-2.5011e-12 | 25.9775 | -4.3428e-11 | 25.8995
-9.0949e-13 | 26.3855 | -3.4788e-11 | 26.3045
-2.2737e-13 | 26.7915 | -4.1155e-11 | 26.7095
-1.5916e-12 | 27.1956 | -4.1837e-11 | 27.1166
-1.1369e-12 | 27.5996 | -4.4338e-11 | 27.5226
-1.5916e-12 | 28.0056 | -4.5475e-11 | 27.9286
-9.0949e-13 | 28.4096 | -4.5702e-11 | 28.3346
-2.2737e-12 | 28.8136 | -3.9563e-11 | 28.7426
-2.5011e-12 | 29.2187 | -3.5925e-11 | 29.1497
-9.0949e-13 | 29.6237 | -4.2974e-11 | 29.5567
0 | 30.0277 | -4.07e-11 | 29.9617
-9.0949e-13 | 30.4297 | -3.5925e-11 | 30.3737
-9.0949e-13 | 30.8338 | -3.7744e-11 | 30.7808
-1.819e-12 | 31.2368 | -3.7289e-11 | 31.1858
-1.3642e-12 | 31.6418 | -4.07e-11 | 31.5938
-1.1369e-12 | 32.0458 | -4.2064e-11 | 31.9988
-1.3642e-12 | 32.4509 | -4.1609e-11 | 32.4038
-2.7285e-12 | 32.8549 | -4.0245e-11 | 32.8159
-3.8654e-12 | 33.2609 | -4.1155e-11 | 33.2229
-1.819e-12 | 33.6659 | -3.8426e-11 | 33.6299
-1.3642e-12 | 34.0699 | -3.8199e-11 | 34.045
-3.8654e-12 | 34.475 | -3.8426e-11 | 34.454
-1.5916e-12 | 34.881 | -4.2064e-11 | 34.86
-2.2737e-13 | 35.285 | -4.411e-11 | 35.266
2.2737e-13 | 35.689 | -3.7289e-11 | 35.673
-1.5916e-12 | 36.0921 | -4.2974e-11 | 36.0811
-9.0949e-13 | 36.4971 | -3.6152e-11 | 36.4861
-3.638e-12 | 36.9001 | -4.502e-11 | 36.8911
-1.5916e-12 | 37.3071 | -4.5929e-11 | 37.2971
-2.2737e-12 | 37.7132 | -4.2519e-11 | 37.7042
-1.3642e-12 | 38.1192 | -3.6152e-11 | 38.1092
-2.5011e-12 | 38.5242 | -3.979e-11 | 38.5152
2.2737e-13 | 38.9292 | -4.2519e-11 | 38.9212
2.2737e-13 | 39.3332 | -4.411e-11 | 39.3263
-1.5916e-12 | 39.7373 | -4.502e-11 | 39.7323
-3.638e-12 | 40.1443 | -4.1837e-11 | 40.1383
-1.819e-12 | 40.5493 | -4.411e-11 | 40.5433
2.2737e-13 | 40.9543 | -3.547e-11 | 40.9533
-2.7285e-12 | 41.3584 | -3.9108e-11 | 41.3604
-1.5916e-12 | 41.7644 | -4.1609e-11 | 41.7654
-2.7285e-12 | 42.1684 | -3.9108e-11 | 42.1744
-1.1369e-12 | 42.5734 | -3.7971e-11 | 42.5824
1.1369e-12 | 42.9775 | -4.2519e-11 | 42.9875
-1.5916e-12 | 43.3815 | -4.3656e-11 | 43.3985
-2.5011e-12 | 43.7855 | -3.547e-11 | 43.8035
-1.1369e-12 | 44.1895 | -4.7066e-11 | 44.2095
0 | 44.5936 | -4.4338e-11 | 44.6185
-1.1369e-12 | 44.9956 | -4.2291e-11 | 45.0236
-1.5916e-12 | 45.4006 | -4.07e-11 | 45.4446
-5.457e-12 | 45.8066 | -4.3656e-11 | 45.8546
-1.1369e-12 | 46.2116 | -3.9108e-11 | 46.2597
-1.1369e-12 | 46.6167 | -4.2519e-11 | 46.6647
-9.0949e-13 | 47.0217 | -4.502e-11 | 47.0697
-1.1369e-12 | 47.4247 | -4.3428e-11 | 47.4767
-2.9559e-12 | 47.8307 | -3.8654e-11 | 47.8817
0 | 48.2358 | -4.07e-11 | 48.2858
9.0949e-13 | 48.6388 | -4.07e-11 | 48.6928
-4.5475e-13 | 49.0418 | -4.3428e-11 | 49.0978
-2.9559e-12 | 49.4468 | -4.1837e-11 | 49.5038
-4.0927e-12 | 49.8528 | -3.9108e-11 | 49.9119
-1.3642e-12 | 50.2569 | -3.5925e-11 | 50.3169
1.5916e-12 | 50.6619 | -4.6157e-11 | 50.7279
-9.0949e-13 | 51.0659 | -3.7289e-11 | 51.1349
-1.819e-12 | 51.4719 | -4.1609e-11 | 51.5429
-2.5011e-12 | 51.877 | -3.9108e-11 | 51.949
2.2737e-13 | 52.283 | -4.4338e-11 | 52.353
2.2737e-13 | 52.686 | -4.1155e-11 | 52.773
-1.5916e-12 | 53.09 | -4.4338e-11 | 53.179
-2.2737e-12 | 53.4951 | -4.3428e-11 | 53.5881
-9.0949e-13 | 53.9001 | -3.7744e-11 | 53.9951
-2.2737e-13 | 54.3051 | -4.5702e-11 | 54.4001
-4.5475e-13 | 54.7091 | -4.8431e-11 | 54.8091
-1.3642e-12 | 55.1142 | -4.1609e-11 | 55.2212
4.5475e-13 | 55.5202 | -4.5475e-11 | 55.6262
-1.5916e-12 | 55.9242 | -4.0245e-11 | 56.0312
-1.5916e-12 | 56.3292 | -4.07e-11 | 56.4382
-5.0022e-12 | 56.7353 | -4.2291e-11 | 56.8492
-2.5011e-12 | 57.1403 | -4.1837e-11 | 57.2553
-1.1369e-12 | 57.5463 | -4.5929e-11 | 57.6613
-1.3642e-12 | 57.9513 | -4.3428e-11 | 58.0673
-3.4106e-12 | 58.3543 | -4.3656e-11 | 58.4733
-3.8654e-12 | 58.7584 | -3.7744e-11 | 58.8794
9.0949e-13 | 59.1634 | -4.4338e-11 | 59.2854
-4.5475e-13 | 59.5664 | -4.5929e-11 | 59.7064
-2.7285e-12 | 59.9714 | -4.6384e-11 | 60.1144
-3.638e-12 | 60.3775 | -3.9108e-11 | 60.5285
-2.5011e-12 | 60.7825 | -3.8654e-11 | 60.9365
1.1369e-12 | 61.1865 | -4.4565e-11 | 61.3425
-1.3642e-12 | 61.5905 | -4.0473e-11 | 61.7485
-2.7285e-12 | 61.9956 | -3.9336e-11 | 62.1546
-3.638e-12 | 62.4016 | -4.5929e-11 | 62.5696
9.0949e-13 | 62.8076 | -4.5475e-11 | 62.9786
-2.2737e-13 | 63.2126 | -4.07e-11 | 63.5196
-2.9559e-12 | 63.6146 | -4.3201e-11 | 63.9257
-2.2737e-12 | 64.0197 | -4.3656e-11 | 64.3317
-1.819e-12 | 64.4247 | -4.4565e-11 | 64.7397
-1.5916e-12 | 64.8287 | -3.9108e-11 | 65.2017
9.0949e-13 | 65.2327 | -3.5243e-11 | 65.6178
-2.7285e-12 | 65.6377 | -3.5925e-11 | 66.0288
-2.7285e-12 | 66.0418 | -4.6384e-11 | 66.4358
-1.1369e-12 | 66.4458 | -3.8654e-11 | 66.8408
-1.3642e-12 | 66.8508 | -3.4788e-11 | 67.2478
-1.3642e-12 | 67.2538 | -4.2064e-11 | 67.6579
-1.5916e-12 | 67.6589 | -4.7066e-11 | 68.0639
-2.2737e-12 | 68.0629 | -4.0245e-11 | 68.4719
-2.2737e-12 | 68.4689 | -3.9563e-11 | 68.8799
-1.3642e-12 | 68.8729 | -4.1155e-11 | 69.286
-2.2737e-12 | 69.284 | -4.3201e-11 | 69.708
-9.0949e-13 | 69.688 | -4.4338e-11 | 70.118
-1.5916e-12 | 70.093 | -4.4793e-11 | 70.524
-9.0949e-13 | 70.498 | -3.7289e-11 | 70.9351
-3.8654e-12 | 70.9021 | -4.2974e-11 | 71.3421
-2.2737e-12 | 71.3071 | -4.3656e-11 | 71.7521
-1.1369e-12 | 71.7131 | -3.9108e-11 | 72.1591
-9.0949e-13 | 72.1181 | -4.2291e-11 | 72.5641
-1.1369e-12 | 72.5352 | -4.1837e-11 | 72.9692
-1.1369e-12 | 72.9392 | -4.2064e-11 | 73.3742
-2.5011e-12 | 73.3442 | -4.2974e-11 | 73.7802
-1.5916e-12 | 73.7512 | -3.979e-11 | 74.1852
-2.2737e-13 | 74.1552 | -4.07e-11 | 74.5903
-4.0927e-12 | 74.5793 | -4.4338e-11 | 74.9963
-2.7285e-12 | 75.0123 | -4.1155e-11 | 75.4013
2.2737e-13 | 75.4173 | -4.2974e-11 | 75.8053
-1.3642e-12 | 75.8213 | -4.2974e-11 | 76.2114
-3.4106e-12 | 76.2254 | -4.3656e-11 | 76.6174
-2.2737e-12 | 76.6304 | -3.6152e-11 | 77.0234
-2.2737e-13 | 77.0354 | -4.1837e-11 | 77.4284
-9.0949e-13 | 77.4404 | -4.7294e-11 | 77.8354
-2.7285e-12 | 77.8454 | -4.3428e-11 | 78.2405
-2.9559e-12 | 78.2495 | -4.5475e-11 | 78.6605
-1.3642e-12 | 78.6535 | -4.2974e-11 | 79.0655
-1.5916e-12 | 79.0585 | -4.0245e-11 | 79.4695
-1.3642e-12 | 79.4635 | -4.3201e-11 | 79.8756
-1.5916e-12 | 79.8686 | -4.0927e-11 | 80.2816
-2.9559e-12 | 80.2726 | -3.7744e-11 | 80.6886
-4.5475e-13 | 80.6786 | -4.3656e-11 | 81.0936
-4.5475e-13 | 81.0826 | -4.5475e-11 | 81.5007
-2.7285e-12 | 81.4857 | -4.411e-11 | 81.9057
-1.1369e-12 | 81.8907 | -4.0473e-11 | 82.3107
-4.3201e-12 | 82.2947 | -4.7521e-11 | 82.7157
-2.2737e-13 | 82.6997 | -4.2974e-11 | 83.1238
-2.7285e-12 | 83.1038 | -4.6157e-11 | 83.5288
-2.9559e-12 | 83.5078 | -4.2064e-11 | 83.9338
-2.5011e-12 | 83.9138 | -3.979e-11 | 84.3388
|  | -4.07e-11 | 84.7448

## 620uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.991e-12 | 2.5218e-12 | -5.8737e-11 | 4.2557e-12
GEMs On |  | GEMs On
-1.819e-12 | 0.31302 | -5.8435e-11 | 0.31002
0 | 0.99806 | -6.2528e-11 | 0.99406
-1.3642e-12 | 1.4001 | -6.1164e-11 | 1.4021
-2.7285e-12 | 1.8041 | -5.5252e-11 | 1.8061
-1.819e-12 | 2.2091 | -6.5029e-11 | 2.2121
-2.5011e-12 | 2.6141 | -6.0481e-11 | 2.6181
2.16e-11 | 3.0182 | -5.9117e-11 | 3.0242
2.2737e-12 | 3.4242 | -5.9117e-11 | 3.4302
-1.819e-12 | 3.8292 | -6.7075e-11 | 3.8362
-2.2737e-12 | 4.2342 | -5.6616e-11 | 4.2412
-2.7285e-12 | 4.6383 | -5.6616e-11 | 4.6473
-6.8212e-13 | 5.0433 | -4.9795e-11 | 5.0523
-3.8654e-12 | 5.4483 | -5.7298e-11 | 5.4583
-6.8212e-13 | 5.8533 | -5.4115e-11 | 5.8643
2.1373e-11 | 6.2574 | -5.4115e-11 | 6.2694
-6.8212e-13 | 6.6634 | -4.7748e-11 | 6.6754
-3.4106e-12 | 7.0674 | -6.0709e-11 | 7.0814
-3.1832e-12 | 7.4724 | -5.9799e-11 | 7.4884
-1.819e-12 | 7.8774 | -6.0481e-11 | 7.8935
-3.4106e-12 | 8.2815 | -6.0936e-11 | 8.2985
-2.2737e-12 | 8.6875 | -6.3665e-11 | 8.7035
-2.2737e-12 | 9.0915 | -5.6161e-11 | 9.1105
-2.5011e-12 | 9.4975 | -5.8662e-11 | 9.5165
-2.2737e-12 | 9.9026 | -6.0027e-11 | 9.9226
-3.638e-12 | 10.3076 | -5.9572e-11 | 10.3276
-9.0949e-13 | 10.7136 | -5.2751e-11 | 10.7336
-3.1832e-12 | 11.1166 | -6.5711e-11 | 11.1386
0 | 11.5217 | -6.321e-11 | 11.5447
-9.0949e-13 | 11.9237 | -6.1164e-11 | 11.9507
-3.1832e-12 | 12.3277 | -6.1164e-11 | 12.3567
-2.2737e-12 | 12.7327 | -5.7071e-11 | 12.7627
-2.2737e-12 | 13.1378 | -5.9345e-11 | 13.1668
-1.3642e-12 | 13.5438 | -5.1841e-11 | 13.5728
-2.5011e-12 | 13.9488 | -5.6616e-11 | 13.9788
0 | 14.3518 | -5.5707e-11 | 14.3848
4.5475e-13 | 14.7538 | -6.4574e-11 | 14.7889
-2.5011e-12 | 15.1569 | -5.8662e-11 | 15.1959
-2.5011e-12 | 15.5619 | -5.8662e-11 | 15.6019
-1.819e-12 | 15.9669 | -6.6848e-11 | 16.0069
-2.7285e-12 | 16.3729 | -5.2978e-11 | 16.4119
-2.5011e-12 | 16.777 | -5.7071e-11 | 16.819
-2.5011e-12 | 17.182 | -5.5479e-11 | 17.225
-9.0949e-13 | 17.586 | -6.8439e-11 | 17.631
-9.0949e-13 | 17.989 | -5.4115e-11 | 18.036
-2.2737e-12 | 18.393 | -5.7298e-11 | 18.442
-2.2737e-13 | 18.8001 | -6.1618e-11 | 18.8471
-1.3642e-12 | 19.2031 | -5.457e-11 | 19.2531
-2.2737e-12 | 19.6091 | -6.321e-11 | 19.6571
-1.1369e-12 | 20.0141 | -5.1614e-11 | 20.0632
-2.0464e-12 | 20.4182 | -6.0936e-11 | 20.4682
-2.0464e-12 | 20.8232 | -5.1841e-11 | 20.8732
-3.1832e-12 | 21.2292 | -5.798e-11 | 21.2782
-2.2737e-12 | 21.6342 | -6.1164e-11 | 21.6832
-1.819e-12 | 22.0403 | -6.4801e-11 | 22.0883
-2.5011e-12 | 22.4463 | -6.0936e-11 | 22.4953
-2.2737e-12 | 22.8503 | -5.6616e-11 | 22.9013
-1.1369e-12 | 23.2553 | -5.5707e-11 | 23.3053
-2.2737e-12 | 23.6593 | -5.7071e-11 | 23.7104
2.2737e-13 | 24.0644 | -5.5707e-11 | 24.1154
-1.819e-12 | 24.4694 | -5.9799e-11 | 24.5224
-9.0949e-13 | 24.8744 | -5.8208e-11 | 24.9274
-3.638e-12 | 25.2795 | -6.2073e-11 | 25.3325
-4.7748e-12 | 25.6835 | -6.2073e-11 | 25.7365
1.1369e-12 | 26.0885 | -5.9117e-11 | 26.1425
-6.8212e-13 | 26.4925 | -6.3892e-11 | 26.5465
-1.1369e-12 | 26.8975 | -6.6393e-11 | 26.9515
-3.1832e-12 | 27.3026 | -5.5479e-11 | 27.3566
-2.0464e-12 | 27.7076 | -6.0481e-11 | 27.7636
-2.2737e-12 | 28.1126 | -6.0481e-11 | 28.1686
-1.1369e-12 | 28.5176 | -5.4342e-11 | 28.5736
-2.0464e-12 | 28.9216 | -6.1846e-11 | 28.9787
-2.0464e-12 | 29.3267 | -5.3433e-11 | 29.3847
-1.1369e-12 | 29.7317 | -5.457e-11 | 29.7887
-2.5011e-12 | 30.1377 | -6.1164e-11 | 30.1937
-1.1369e-12 | 30.5418 | -6.3665e-11 | 30.5998
-2.2737e-12 | 30.9458 | -5.8662e-11 | 31.0068
-1.819e-12 | 31.3518 | -5.7298e-11 | 31.4118
0 | 31.7558 | -6.0709e-11 | 31.8178
-3.1832e-12 | 32.1588 | -6.2073e-11 | 32.2238
-2.2737e-12 | 32.5639 | -5.9117e-11 | 32.6279
-1.3642e-12 | 32.9689 | -5.8208e-11 | 33.0329
-3.8654e-12 | 33.3739 | -4.8203e-11 | 33.4379
-1.819e-12 | 33.7779 | -7.185e-11 | 33.8429
-1.3642e-12 | 34.183 | -5.4342e-11 | 34.25
-6.8212e-13 | 34.587 | -6.1618e-11 | 34.656
-1.819e-12 | 34.99 | -5.9799e-11 | 35.061
-2.5011e-12 | 35.394 | -6.1846e-11 | 35.467
-9.0949e-13 | 35.7981 | -5.4797e-11 | 35.8721
-6.8212e-13 | 36.2021 | -5.4797e-11 | 36.2781
-3.1832e-12 | 36.6061 | -6.753e-11 | 36.6831
-1.1369e-12 | 37.0121 | -6.7303e-11 | 37.0881
-1.1369e-12 | 37.4171 | -5.5252e-11 | 37.4941
-2.7285e-12 | 37.8222 | -5.5934e-11 | 37.9012
-1.819e-12 | 38.2262 | -5.8435e-11 | 38.3062
-1.1369e-12 | 38.6312 | -5.8435e-11 | 38.7112
-3.4106e-12 | 39.0362 | -5.7298e-11 | 39.1162
-2.7285e-12 | 39.4423 | -6.3665e-11 | 39.5223
-1.819e-12 | 39.8473 | -6.1618e-11 | 39.9273
-3.638e-12 | 40.2523 | -6.0481e-11 | 40.3323
-2.2737e-12 | 40.6573 | -5.457e-11 | 40.7373
-2.7285e-12 | 41.0643 | -5.6843e-11 | 41.1414
-2.2737e-12 | 41.4694 | -5.6616e-11 | 41.5474
-2.7285e-12 | 41.8744 | -5.457e-11 | 41.9524
-2.7285e-12 | 42.2784 | -5.9345e-11 | 42.3574
-6.8212e-13 | 42.6844 | -5.8435e-11 | 42.7625
-9.0949e-13 | 43.0885 | -5.6616e-11 | 43.1685
-6.8212e-13 | 43.4915 | -6.4574e-11 | 43.5735
-2.0464e-12 | 43.8955 | -5.9345e-11 | 43.9805
-2.2737e-12 | 44.3015 | -6.3665e-11 | 44.3855
-4.0927e-12 | 44.7056 | -5.4342e-11 | 44.7916
-1.1369e-12 | 45.1106 | -5.5479e-11 | 45.1966
-2.2737e-12 | 45.5156 | -4.8885e-11 | 45.6016
-1.3642e-12 | 45.9216 | -5.9572e-11 | 46.0056
-2.5011e-12 | 46.3276 | -5.7071e-11 | 46.4096
-1.819e-12 | 46.7327 | -6.321e-11 | 46.8157
-3.8654e-12 | 47.1367 | -5.0704e-11 | 47.2207
-2.2737e-12 | 47.5417 | -6.2073e-11 | 47.6267
-2.0464e-12 | 47.9467 | -5.5934e-11 | 48.0318
-6.8212e-13 | 48.3508 | -6.3665e-11 | 48.4378
-2.5011e-12 | 48.7558 | -5.7753e-11 | 48.8428
-2.0464e-12 | 49.1608 | -6.7075e-11 | 49.2488
-9.0949e-13 | 49.5658 | -5.9572e-11 | 49.6548
-3.4106e-12 | 49.9689 | -5.8208e-11 | 50.0609
-3.4106e-12 | 50.3729 | -5.7298e-11 | 50.4659
-2.5011e-12 | 50.7779 | -5.8435e-11 | 50.8719
-3.4106e-12 | 51.1839 | -6.3892e-11 | 51.2769
-1.819e-12 | 51.5889 | -6.5484e-11 | 51.683
-2.2737e-12 | 51.995 | -6.0481e-11 | 52.087
-9.0949e-13 | 52.399 | -5.6843e-11 | 52.494
-6.8212e-13 | 52.803 | -6.3665e-11 | 52.899
-3.4106e-12 | 53.207 | -5.457e-11 | 53.304
-1.3642e-12 | 53.6131 | -5.8662e-11 | 53.7081
-2.2737e-12 | 54.0191 | -5.798e-11 | 54.1131
-2.5011e-12 | 54.4231 | -6.0936e-11 | 54.5191
-2.2737e-12 | 54.8281 | -5.5252e-11 | 54.9251
-1.3642e-12 | 55.2332 | -5.8435e-11 | 55.3302
0 | 55.6372 | -6.5938e-11 | 55.7372
-1.819e-12 | 56.0412 | -4.7521e-11 | 56.1422
-2.5011e-12 | 56.4462 | -6.2528e-11 | 56.5472
-3.4106e-12 | 56.8503 | -5.8208e-11 | 56.9513
-2.7285e-12 | 57.2553 | -6.1164e-11 | 57.3573
-6.8212e-13 | 57.6603 | -6.1164e-11 | 57.7633
-3.1832e-12 | 58.0623 | -6.1618e-11 | 58.1673
-1.1369e-12 | 58.4693 | -5.5934e-11 | 58.5724
-2.0464e-12 | 58.8744 | -5.9117e-11 | 58.9784
-2.2737e-12 | 59.2794 | -5.6843e-11 | 59.3844
-6.8212e-13 | 59.6834 | -6.23e-11 | 59.7904
-4.5475e-12 | 60.0874 | -6.5938e-11 | 60.1954
-2.0464e-12 | 60.4925 | -4.9113e-11 | 60.6005
4.5475e-13 | 60.8975 | -5.9572e-11 | 61.0065
-1.3642e-12 | 61.3015 | -5.6161e-11 | 61.4105
-1.3642e-12 | 61.7065 | -5.8435e-11 | 61.8155
-3.4106e-12 | 62.1106 | -6.5029e-11 | 62.2206
-2.2737e-13 | 62.5156 | -5.7753e-11 | 62.6266
-2.0464e-12 | 62.9186 | -6.8212e-11 | 63.0326
-2.5011e-12 | 63.3246 | -5.2978e-11 | 63.4386
-1.3642e-12 | 63.7317 | -6.23e-11 | 63.8426
-3.1832e-12 | 64.1367 | -5.5252e-11 | 64.2477
-1.819e-12 | 64.5417 | -5.7298e-11 | 64.6537
-4.0927e-12 | 64.9467 | -6.0936e-11 | 65.0587
-2.5011e-12 | 65.3517 | -4.934e-11 | 65.4647
-2.7285e-12 | 65.7558 | -5.1614e-11 | 65.8698
-9.0949e-13 | 66.1608 | -5.798e-11 | 66.2748
-3.1832e-12 | 66.5648 | -5.7071e-11 | 66.6808
-3.1832e-12 | 66.9698 | -6.2073e-11 | 67.0858
-1.819e-12 | 67.3739 | -5.9345e-11 | 67.4909
-1.3642e-12 | 67.7799 | -6.0936e-11 | 67.8959
-2.7285e-12 | 68.1859 | -5.6616e-11 | 68.3019
-6.8212e-13 | 68.5899 | -5.7071e-11 | 68.7089
-1.819e-12 | 68.9939 | -5.9117e-11 | 69.114
-1.819e-12 | 69.398 | -5.5252e-11 | 69.519
-1.1369e-12 | 69.802 | -5.9345e-11 | 69.925
-6.8212e-13 | 70.209 | -5.457e-11 | 70.33
-1.3642e-12 | 70.612 | -6.3437e-11 | 70.7351
2.2737e-13 | 71.0171 | -5.5934e-11 | 71.1411
-9.0949e-13 | 71.4221 | -6.0027e-11 | 71.5481
-4.0927e-12 | 71.8251 | -6.2528e-11 | 71.9541
-1.819e-12 | 72.2291 | -6.0027e-11 | 72.3591
-4.0927e-12 | 72.6342 | -4.7066e-11 | 72.7672
-2.5011e-12 | 73.0392 | -5.5934e-11 | 73.1742
-3.8654e-12 | 73.4442 | -5.366e-11 | 73.5802
-2.2737e-12 | 73.8492 | -5.5934e-11 | 73.9852
-2.0464e-12 | 74.2532 | -6.0027e-11 | 74.3912
-1.3642e-12 | 74.6583 | -5.4342e-11 | 74.7963
-9.0949e-13 | 75.0633 | -5.7753e-11 | 75.2023
-1.3642e-12 | 75.4673 | -6.1164e-11 | 75.6073
-3.1832e-12 | 75.8723 | -6.321e-11 | 76.0153
-2.5011e-12 | 76.2774 | -6.1618e-11 | 76.4194
-2.2737e-12 | 76.6824 | -5.6616e-11 | 76.8244
-1.819e-12 | 77.0874 | -6.2528e-11 | 77.2294
-2.5011e-12 | 77.4924 | -5.8435e-11 | 77.6364
-1.1369e-12 | 77.8975 | -5.5252e-11 | 78.0435
-1.1369e-12 | 78.3025 | -5.7071e-11 | 78.4485
-3.1832e-12 | 78.7065 | -5.798e-11 | 78.8545
-2.5011e-12 | 79.1115 | -6.3665e-11 | 79.2595
-2.2737e-12 | 79.5165 | -5.5252e-11 | 79.6656
-1.1369e-12 | 79.9196 | -5.6161e-11 | 80.0706
-3.1832e-12 | 80.3246 | -6.0027e-11 | 80.4776
-2.0464e-12 | 80.7296 | -6.1618e-11 | 80.8836
-2.5011e-12 | 81.1346 | -6.3665e-11 | 81.2887
-6.8212e-13 | 81.5407 | -5.3433e-11 | 81.6947
-1.1369e-12 | 81.9447 | -5.9799e-11 | 82.0997
-2.2737e-13 | 82.3497 | -6.23e-11 | 82.5047
-2.0464e-12 | 82.7537 | -5.6616e-11 | 82.9107
-2.2737e-12 | 83.1588 | -5.6843e-11 | 83.3158
-1.3642e-12 | 83.5648 | -5.5707e-11 | 83.7218
-2.5011e-12 | 83.9698 | -5.798e-11 | 84.1268
-2.5011e-12 | 84.3748
-2.2737e-12 | 84.7798
-6.8212e-13 | 85.1839
-4.0927e-12 | 85.5889
-2.0464e-12 | 85.9949
-1.3642e-12 | 86.4009
-1.819e-12 | 86.805
-3.4106e-12 | 87.21
-2.7285e-12 | 87.615
-6.8212e-13 | 88.02
-2.0464e-12 | 88.4231
-4.0927e-12 | 88.8271
-2.7285e-12 | 89.2321
-1.819e-12 | 89.6371
-2.0464e-12 | 90.0422
-4.7748e-12 | 90.4472
-1.819e-12 | 90.8512
-5.0022e-12 | 91.2572
-6.8212e-13 | 91.6622
-1.819e-12 | 92.0663
-2.0464e-12 | 92.4703
-2.2737e-12 | 92.8753
-1.819e-12 | 93.2793
-2.0464e-12 | 93.6824
-6.8212e-13 | 94.0874
-2.0464e-12 | 94.4914
-1.819e-12 | 94.8964
-2.5011e-12 | 95.3015
2.2737e-13 | 95.7065
-1.819e-12 | 96.1105
-3.1832e-12 | 96.5155
-3.4106e-12 | 96.9195
-1.819e-12 | 97.3246
-2.2737e-12 | 97.7296
-2.2737e-12 | 98.1336
-3.8654e-12 | 98.5396
-1.3642e-12 | 98.9447
-4.0927e-12 | 99.3497
-6.8212e-13 | 99.7547
-4.0927e-12 | 100.1577
-1.3642e-12 | 100.5628
-1.1369e-12 | 100.9668
-2.0464e-12 | 101.3708
-1.3642e-12 | 101.7768
-3.4106e-12 | 102.1818
-2.7285e-12 | 102.5869
-2.7285e-12 | 102.9919
-3.1832e-12 | 103.3959
-2.5011e-12 | 103.8009
6.8212e-13 | 104.207
-2.2737e-12 | 104.61
-2.0464e-12 | 105.015
-2.5011e-12 | 105.42
-2.0464e-12 | 105.8241
-2.7285e-12 | 106.2311
-3.1832e-12 | 106.6351
-1.1369e-12 | 107.0391
-2.0464e-12 | 107.4431
-1.819e-12 | 107.8492
-1.1369e-12 | 108.2552
-2.7285e-12 | 108.6592
-2.2737e-12 | 109.0632
-2.0464e-12 | 109.4703
-1.3642e-12 | 109.8753
-2.2737e-12 | 110.2803
-1.1369e-12 | 110.6853
-1.819e-12 | 111.0904
-2.5011e-12 | 111.4954
-2.0464e-12 | 111.8994
-2.2737e-12 | 112.3044
-2.7285e-12 | 112.7094
-2.2737e-12 | 113.1145
-2.7285e-12 | 113.5195
-2.5011e-12 | 113.9245
-2.2737e-12 | 114.3275
1.5916e-12 | 114.7316
-3.4106e-12 | 115.1366
-2.2737e-12 | 115.5416
-1.3642e-12 | 115.9466
-1.1369e-12 | 116.3507
-3.4106e-12 | 116.7567
-3.4106e-12 | 117.1607
-2.5011e-12 | 117.5667
-2.2737e-12 | 117.9697
1.819e-12 | 118.3758
-2.16e-11 | 118.7818
-6.3665e-12 | 119.1858
-2.2737e-12 | 119.5908
-2.2737e-12 | 119.9969
-2.2737e-12 | 120.4029
-1.3642e-12 | 120.8079
-2.7285e-12 | 121.2139
-1.819e-12 | 121.62
-5.6843e-12 | 122.025
-8.8676e-12 | 122.43
-6.3665e-12 | 122.835
-6.3665e-12 | 123.238
-4.0927e-12 | 123.6421
-9.0949e-13 | 124.0481
-2.0464e-12 | 124.4521
-1.819e-12 | 124.8581
-1.3642e-12 | 125.2632
-2.5011e-12 | 125.6682
2.2737e-13 | 126.0732
-1.1369e-12 | 126.4772

## 630uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-2.5416e-12 | 1.1449e-12 | -8.23e-11 | 5.8498e-12
GEMs On |  | GEMs On
-4.0927e-12 | 0.31002 | -1.005e-10 | 0.31202
-5.0022e-12 | 0.99406 | -8.913e-11 | 0.99706
-1.1369e-12 | 1.3991 | -8.1172e-11 | 1.4021
-1.1369e-12 | 1.8031 | -8.5947e-11 | 1.8081
-1.5916e-12 | 2.2081 | -8.3901e-11 | 2.2131
-2.5011e-12 | 2.6141 | -7.7307e-11 | 2.6181
-1.819e-12 | 3.0192 | -8.4128e-11 | 3.0242
-4.5475e-12 | 3.4232 | -8.9813e-11 | 3.4312
-1.819e-12 | 3.8272 | -8.4583e-11 | 3.8362
-2.9559e-12 | 4.2332 | -7.8899e-11 | 4.2412
-3.1832e-12 | 4.6443 | -7.7534e-11 | 4.6473
-9.0949e-13 | 5.0493 | -8.7539e-11 | 5.0533
-2.2737e-12 | 5.4533 | -8.5265e-11 | 5.4583
-2.2737e-12 | 5.8573 | -7.8671e-11 | 5.8633
-9.0949e-13 | 6.2654 | -8.0263e-11 | 6.2694
-2.5011e-12 | 6.6674 | -8.2991e-11 | 6.6744
-2.9559e-12 | 7.0724 | -7.8671e-11 | 7.0794
-1.819e-12 | 7.4784 | -8.7994e-11 | 7.4844
-2.2737e-12 | 7.8845 | -8.2991e-11 | 7.8905
-4.3201e-12 | 8.2895 | -7.8444e-11 | 8.2965
-1.1369e-12 | 8.6935 | -8.6175e-11 | 8.7025
-5.6843e-12 | 9.0985 | -8.8676e-11 | 9.1085
-2.9559e-12 | 9.5025 | -7.6398e-11 | 9.5135
-2.5011e-12 | 9.9066 | -7.9808e-11 | 9.9186
-3.1832e-12 | 10.3106 | -8.6175e-11 | 10.3236
-1.819e-12 | 10.7166 | -8.2309e-11 | 10.7286
-1.5916e-12 | 11.1216 | -9.6406e-11 | 11.1336
-2.2737e-12 | 11.5257 | -9.1632e-11 | 11.5387
-3.638e-12 | 11.9297 | -8.4583e-11 | 11.9447
-3.1832e-12 | 12.3337 | -8.6175e-11 | 12.3507
-4.5475e-13 | 12.7377 | -8.5947e-11 | 12.8157
-2.2737e-12 | 13.1408 | -7.0713e-11 | 13.2228
-3.1832e-12 | 13.5458 | -7.5715e-11 | 13.6278
-2.0464e-12 | 13.9498 | -9.0267e-11 | 14.0348
-2.2737e-12 | 14.3538 | -7.276e-11 | 14.4408
-3.1832e-12 | 14.7608 | -7.6398e-11 | 14.8489
-1.819e-12 | 15.1669 | -7.8899e-11 | 15.2929
-3.1832e-12 | 15.5719 | -8.0263e-11 | 15.6999
-2.2737e-12 | 15.9769 | -7.8899e-11 | 16.1039
-2.2737e-12 | 16.3819 | -7.8899e-11 | 16.5109
-1.819e-12 | 16.786 | -8.0263e-11 | 16.917
-1.819e-12 | 17.191 | -7.6852e-11 | 17.325
-2.0464e-12 | 17.596 | -8.1627e-11 | 17.73
-1.819e-12 | 18 | -8.2764e-11 | 18.136
-2.2737e-13 | 18.405 | -8.7766e-11 | 18.5411
-2.2737e-12 | 18.8091 | -8.4128e-11 | 18.9481
-1.819e-12 | 19.2151 | -7.5715e-11 | 19.3541
-3.1832e-12 | 19.6191 | -8.3901e-11 | 19.7601
-9.0949e-13 | 20.0241 | -7.7307e-11 | 20.1671
-6.1391e-12 | 20.4282 | -9.1177e-11 | 20.5722
-1.1369e-12 | 20.8332 | -7.7534e-11 | 20.9782
-2.9559e-12 | 21.2372 | -9.1859e-11 | 21.3832
-3.1832e-12 | 21.6422 | -8.0718e-11 | 21.7892
-2.9559e-12 | 22.0463 | -7.185e-11 | 22.1943
-2.9559e-12 | 22.4513 | -7.2077e-11 | 22.5993
-4.0927e-12 | 22.8563 | -8.3901e-11 | 23.0053
-1.5916e-12 | 23.2613 | -8.0263e-11 | 23.4103
-2.9559e-12 | 23.6654 | -8.6402e-11 | 23.8164
-2.2737e-13 | 24.0714 | -7.5715e-11 | 24.2214
-2.9559e-12 | 24.4754 | -9.1632e-11 | 24.6274
-2.2737e-12 | 24.8814 | -8.6629e-11 | 25.0344
-6.8212e-13 | 25.2844 | -8.0945e-11 | 25.4405
-2.9559e-12 | 25.6885 | -8.8448e-11 | 25.8535
-2.2737e-12 | 26.0935 | -7.7762e-11 | 26.2595
-2.5011e-12 | 26.4995 | -8.0036e-11 | 26.6665
-2.5011e-12 | 26.9045 | -7.8444e-11 | 27.0706
-1.1369e-12 | 27.3086 | -8.7766e-11 | 27.4776
-2.0464e-12 | 27.7156 | -8.8903e-11 | 27.8836
-2.2737e-12 | 28.1206 | -8.6857e-11 | 28.2886
-3.4106e-12 | 28.5246 | -6.9804e-11 | 28.6946
-2.5011e-12 | 28.9297 | -9.2996e-11 | 29.1007
-1.1369e-12 | 29.3347 | -7.7534e-11 | 29.5067
-3.1832e-12 | 29.7407 | -7.3896e-11 | 29.9117
-2.2737e-13 | 30.1447 | -7.9808e-11 | 30.3177
-2.9559e-12 | 30.5487 | -7.6398e-11 | 30.7228
-2.9559e-12 | 30.9528 | -8.5038e-11 | 31.1278
-2.5011e-12 | 31.3578 | -7.708e-11 | 31.5328
-2.5011e-12 | 31.7628 | -8.0263e-11 | 31.9388
-2.0464e-12 | 32.1678 | -7.9808e-11 | 32.3458
-3.1832e-12 | 32.5729 | -8.3674e-11 | 32.7519
-3.638e-12 | 32.9789 | -7.2077e-11 | 33.1589
-2.9559e-12 | 33.3829 | -8.0718e-11 | 33.5639
-1.5916e-12 | 33.7879 | -8.481e-11 | 33.9699
-3.4106e-12 | 34.194 | -8.2764e-11 | 34.375
-9.0949e-13 | 34.599 | -8.0036e-11 | 34.781
-2.0464e-12 | 35.004 | -8.1172e-11 | 35.186
-9.0949e-13 | 35.408 | -8.8448e-11 | 35.592
-2.2737e-12 | 35.813 | -8.4128e-11 | 35.9981
-4.0927e-12 | 36.2181 | -7.5033e-11 | 36.4031
-6.8212e-13 | 36.6241 | -9.0949e-11 | 36.8081
-3.638e-12 | 37.0281 | -9.0495e-11 | 37.2141
-3.1832e-12 | 37.4331 | -7.2305e-11 | 37.6191
-3.1832e-12 | 37.8372 | -7.185e-11 | 38.0242
-2.9559e-12 | 38.2422 | -9.1177e-11 | 38.4312
-2.5011e-12 | 38.6472 | -8.2764e-11 | 38.8372
-5.457e-12 | 39.0522 | -7.708e-11 | 39.2422
-1.819e-12 | 39.4573 | -8.2082e-11 | 39.6483
-3.1832e-12 | 39.8613 | -8.0036e-11 | 40.0553
-2.0464e-12 | 40.2663 | -7.9126e-11 | 40.4603
-2.2737e-12 | 40.6733 | -8.481e-11 | 40.8653
-2.0464e-12 | 41.0774 | -7.8444e-11 | 41.2714
-3.4106e-12 | 41.4814 | -8.2309e-11 | 41.6784
-2.5011e-12 | 41.8874 | -8.5038e-11 | 42.0834
-3.638e-12 | 42.2914 | -7.9126e-11 | 42.4894
-3.1832e-12 | 42.6964 | -8.2309e-11 | 42.8954
-6.8212e-13 | 43.1015 | -8.4583e-11 | 43.3015
-1.5916e-12 | 43.5055 | -6.753e-11 | 43.7075
-3.4106e-12 | 43.9105 | -9.1177e-11 | 44.1125
-4.0927e-12 | 44.3165 | -8.1627e-11 | 44.5175
-3.1832e-12 | 44.7216 | -8.6629e-11 | 44.9226
-3.1832e-12 | 45.1256 | -8.7311e-11 | 45.3276
-2.9559e-12 | 45.5316 | -8.2991e-11 | 45.7316
-9.0949e-13 | 45.9356 | -8.2537e-11 | 46.1376
-2.9559e-12 | 46.3406 | -8.1627e-11 | 46.5427
-1.819e-12 | 46.7447 | -8.913e-11 | 46.9487
-3.1832e-12 | 47.1497 | -8.7311e-11 | 47.3557
-1.819e-12 | 47.5557 | -6.5029e-11 | 47.7617
-3.1832e-12 | 47.9607 | -8.4583e-11 | 48.1668
-2.9559e-12 | 48.3648 | -7.9581e-11 | 48.5728
-1.819e-12 | 48.7688 | -8.2764e-11 | 48.9778
-6.8212e-13 | 49.1738 | -8.0718e-11 | 49.3838
-9.0949e-13 | 49.5768 | -9.6406e-11 | 49.7888
-3.4106e-12 | 49.9809 | -8.6857e-11 | 50.1929
-1.1369e-12 | 50.3869 | -7.617e-11 | 50.5969
-4.5475e-12 | 50.7909 | -8.8903e-11 | 51.0039
-4.3201e-12 | 51.1959 | -7.4579e-11 | 51.4089
-3.4106e-12 | 51.6009 | -7.9581e-11 | 51.814
-2.2737e-12 | 52.006 | -8.1172e-11 | 52.219
-1.1369e-12 | 52.411 | -8.8448e-11 | 52.624
-2.2737e-12 | 52.816 | -8.3901e-11 | 53.029
-3.1832e-12 | 53.221 | -8.6175e-11 | 53.4351
-4.5475e-12 | 53.6251 | -8.5947e-11 | 53.8411
-1.819e-12 | 54.0291 | -8.7766e-11 | 54.2461
-2.2737e-12 | 54.4341 | -8.2537e-11 | 54.6521
-2.2737e-12 | 54.8391 | -8.3446e-11 | 55.0581
-4.0927e-12 | 55.2442 | -8.6629e-11 | 55.4632
-3.4106e-12 | 55.6502 | -8.4128e-11 | 55.8692
-2.5011e-12 | 56.0532 | -8.5493e-11 | 56.2742
-2.2737e-12 | 56.4582 | -7.1395e-11 | 56.6782
-3.638e-12 | 56.8633 | -7.708e-11 | 57.0833
-2.5011e-12 | 57.2673 | -8.5493e-11 | 57.4913
-3.638e-12 | 57.6733 | -8.1172e-11 | 57.8973
-1.819e-12 | 58.0773 | -8.0263e-11 | 58.3023
-2.9559e-12 | 58.4823 | -7.7762e-11 | 58.7074
-2.2737e-13 | 58.8874 | -8.6175e-11 | 59.1124
-4.5475e-12 | 59.2904 | -7.7307e-11 | 59.5184
-1.5916e-12 | 59.6964 | -7.6398e-11 | 59.9234
-3.4106e-12 | 60.1004 | -8.1172e-11 | 60.3284
-2.0464e-12 | 60.5045 | -8.8903e-11 | 60.7335
-3.1832e-12 | 60.9095 | -7.6398e-11 | 61.1385
-2.2737e-12 | 61.3145 | -8.3446e-11 | 61.5445
-2.0464e-12 | 61.7195 | -7.8671e-11 | 61.9495
-3.4106e-12 | 62.1245 | -8.7311e-11 | 62.3546
-2.2737e-12 | 62.5306 | -9.1859e-11 | 62.7606
-3.4106e-12 | 62.9346 | -8.2991e-11 | 63.1666
-3.4106e-12 | 63.3396 | -8.5038e-11 | 63.5716
-4.3201e-12 | 63.7437 | -9.2314e-11 | 63.9777
-3.4106e-12 | 64.1487 | -8.6175e-11 | 64.3827
-1.1369e-12 | 64.5537 | -7.5943e-11 | 64.7887
-4.5475e-12 | 64.9577 | -8.3901e-11 | 65.1947
-2.9559e-12 | 65.3617 | -8.3446e-11 | 65.6008
2.2737e-13 | 65.7668 | -7.1168e-11 | 66.0058
-9.0949e-13 | 66.1698 | -8.481e-11 | 66.4098
-2.0464e-12 | 66.5748 | -8.4583e-11 | 66.8148
-2.2737e-12 | 66.9798 | -9.0495e-11 | 67.2198
-1.1369e-12 | 67.3838 | -7.8899e-11 | 67.6259
-2.9559e-12 | 67.7879 | -8.8676e-11 | 68.0309
-4.7748e-12 | 68.1919 | -8.6402e-11 | 68.4379
-2.2737e-12 | 68.5979 | -8.3901e-11 | 68.8419
-3.1832e-12 | 69.0019 | -6.9349e-11 | 69.247
-1.819e-12 | 69.407 | -8.5038e-11 | 69.653
-1.1369e-12 | 69.811 | -8.7539e-11 | 70.058
-3.4106e-12 | 70.217 | -8.5947e-11 | 70.463
-2.5011e-12 | 70.621 | -7.8671e-11 | 70.8671
-2.0464e-12 | 71.0261 | -7.6398e-11 | 71.2731
-4.5475e-13 | 71.4301 | -7.3669e-11 | 71.6791
-2.9559e-12 | 71.8331 | -7.8444e-11 | 72.0851
-3.638e-12 | 72.2381 | -9.868e-11 | 72.4962
-3.638e-12 | 72.6432 | -7.8444e-11 | 72.9012
-4.0927e-12 | 73.0482 | -8.2309e-11 | 73.3082
-2.5011e-12 | 73.4532 | -8.2309e-11 | 73.7132
-3.1832e-12 | 73.8592 | -7.7762e-11 | 74.1192
-2.9559e-12 | 74.2632 | -7.9808e-11 | 74.5243
-1.819e-12 | 74.6703 | -8.0263e-11 | 74.9293
-1.5916e-12 | 75.0743 | -8.8676e-11 | 75.3343
-4.7748e-12 | 75.4803 | -8.5493e-11 | 75.7403
-3.1832e-12 | 75.8853 | -8.3446e-11 | 76.1464
-2.9559e-12 | 76.2914 | -7.5715e-11 | 76.5514
-1.819e-12 | 76.6964 | -7.6852e-11 | 76.9574
-3.4106e-12 | 77.1004 | -8.0263e-11 | 77.3634
-3.1832e-12 | 77.5044 | -7.5033e-11 | 77.7715
-1.1369e-12 | 77.9095 | -7.5943e-11 | 78.1775
-2.9559e-12 | 78.3135 | -8.5947e-11 | 78.5835
-9.0949e-13 | 78.7185 | -9.1632e-11 | 78.9895
-1.1369e-12 | 79.1225 | -8.5265e-11 | 79.3955
-3.4106e-12 | 79.5276 | -8.6175e-11 | 79.7996
-1.819e-12 | 79.9326 | -7.3442e-11 | 80.2056
-3.4106e-12 | 80.3376 | -8.2537e-11 | 80.6106
-3.638e-12 | 80.7416 | -7.8899e-11 | 81.0166
-2.9559e-12 | 81.1466 | -8.913e-11 | 81.4227
4.5475e-13 | 81.5527 | -9.3905e-11 | 81.8277
-4.3201e-12 | 81.9567 | -8.5493e-11 | 82.2327
-4.0927e-12 | 82.3617 | -7.8444e-11 | 82.6377
-1.5916e-12 | 82.7737 | -8.5493e-11 | 83.0427
-1.1369e-12 | 83.1798 | -7.4806e-11 | 83.4468
-2.2737e-12 | 83.5858 | -7.9808e-11 | 83.8518
-1.819e-12 | 83.9908 | -8.1627e-11 | 84.2578
-4.3201e-12 | 84.3958 | -8.4583e-11 | 84.6628
|  | -7.9581e-11 | 85.0689
|  | -7.2532e-11 | 85.4739
|  | -7.9581e-11 | 85.8799

## 640uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-3.7816e-12 | 1.2229e-12 | -1.1421e-10 | 7.8612e-12
GEMs On |  | GEMs On
-2.9559e-12 | 0.31002 | -1.0164e-10 | 0.31402
-2.9559e-12 | 0.99506 | -1.1346e-10 | 0.99906
1.3642e-12 | 1.4021 | -1.1323e-10 | 1.4051
-3.638e-12 | 1.8071 | -1.0687e-10 | 1.8121
-4.0927e-12 | 2.2121 | -1.0914e-10 | 2.2171
-2.7285e-12 | 2.6171 | -1.1528e-10 | 2.6212
-3.638e-12 | 3.0222 | -1.1391e-10 | 3.0262
-3.638e-12 | 3.4262 | -1.1573e-10 | 3.4312
-4.7748e-12 | 3.8302 | -1.1005e-10 | 3.8352
-4.7748e-12 | 4.2352 | -1.2187e-10 | 4.2412
-1.5916e-12 | 4.6413 | -1.0687e-10 | 4.6463
-3.8654e-12 | 5.0463 | -1.0914e-10 | 5.0513
-4.3201e-12 | 5.4493 | -1.1028e-10 | 5.4553
-4.3201e-12 | 5.8533 | -1.337e-10 | 5.8593
-2.9559e-12 | 6.2594 | -1.2164e-10 | 6.2654
-2.2737e-12 | 6.6654 | -1.1278e-10 | 6.6704
-3.8654e-12 | 7.0744 | -1.1141e-10 | 7.0754
-3.4106e-12 | 7.4814 | -1.1005e-10 | 7.4804
-4.3201e-12 | 7.8875 | -1.2255e-10 | 7.8855
-4.0927e-12 | 8.2955 | -1.0436e-10 | 8.2905
-6.1391e-12 | 8.7005 | -1.1141e-10 | 8.6945
-5.2296e-12 | 9.1145 | -1.2506e-10 | 9.0985
-3.8654e-12 | 9.5245 | -1.1323e-10 | 9.5035
-2.9559e-12 | 9.9296 | -1.1073e-10 | 9.9076
-2.9559e-12 | 10.3346 | -1.146e-10 | 10.3126
-2.7285e-12 | 10.7416 | -1.1437e-10 | 10.7156
-2.2737e-12 | 11.1456 | -1.2346e-10 | 11.1206
-4.0927e-12 | 11.5507 | -1.2437e-10 | 11.5237
-2.7285e-12 | 11.9577 | -1.1823e-10 | 11.9347
-4.7748e-12 | 12.3657 | -1.1937e-10 | 12.3397
-4.0927e-12 | 12.7737 | -1.1073e-10 | 12.7447
-1.819e-12 | 13.1777 | -1.2824e-10 | 13.1487
-6.1391e-12 | 13.6318 | -1.1687e-10 | 13.5548
-2.5011e-12 | 14.0358 | -1.1278e-10 | 13.9588
-1.3642e-12 | 14.4398 | -1.171e-10 | 14.3648
-6.3665e-12 | 14.8459 | -1.171e-10 | 14.7698
-4.3201e-12 | 15.2499 | -1.2619e-10 | 15.1749
-4.7748e-12 | 15.6549 | -1.0778e-10 | 15.5799
-3.8654e-12 | 16.0609 | -1.2437e-10 | 15.9839
-5.0022e-12 | 16.4659 | -1.0778e-10 | 16.3879
-3.4106e-12 | 16.873 | -1.1823e-10 | 16.793
-5.0022e-12 | 17.279 | -1.0778e-10 | 17.197
-2.5011e-12 | 17.683 | -1.0755e-10 | 17.602
-5.0022e-12 | 18.088 | -1.1323e-10 | 18.006
-2.9559e-12 | 18.4931 | -1.0459e-10 | 18.411
-3.8654e-12 | 18.8981 | -1.0959e-10 | 18.8151
-2.7285e-12 | 19.3031 | -1.1141e-10 | 19.2191
-3.4106e-12 | 19.7071 | -1.0778e-10 | 19.6231
-1.5916e-12 | 20.1111 | -1.1619e-10 | 20.0282
-4.0927e-12 | 20.5172 | -1.0959e-10 | 20.4332
-3.638e-12 | 20.9242 | -1.0527e-10 | 20.8372
-4.7748e-12 | 21.3512 | -1.0436e-10 | 21.2412
-3.638e-12 | 21.7763 | -1.08e-10 | 21.6462
-3.8654e-12 | 22.1913 | -9.527e-11 | 22.0503
-3.638e-12 | 22.6013 | -1.146e-10 | 22.4623
-5.457e-12 | 23.0063 | -1.0914e-10 | 22.8663
-3.8654e-12 | 23.4123 | -1.1346e-10 | 23.2713
-4.3201e-12 | 23.8214 | -1.0118e-10 | 23.6793
-4.0927e-12 | 24.2264 | -1.2369e-10 | 24.0834
-4.0927e-12 | 24.6314 | -1.1914e-10 | 24.4874
-4.3201e-12 | 25.0414 | -1.0755e-10 | 24.8924
-4.0927e-12 | 25.4495 | -1.1687e-10 | 25.2965
-2.5011e-12 | 25.8555 | -1.03e-10 | 25.7005
-4.3201e-12 | 26.2635 | -1.1596e-10 | 26.1055
-4.0927e-12 | 26.6695 | -1.1141e-10 | 26.5115
-2.9559e-12 | 27.1045 | -1.0459e-10 | 26.9155
-1.1369e-12 | 27.5116 | -1.0937e-10 | 27.3196
-4.7748e-12 | 27.9156 | -1.1005e-10 | 27.7246
-2.2737e-12 | 28.3216 | -1.196e-10 | 28.1416
-2.9559e-12 | 28.7276 | -1.0959e-10 | 28.5466
-2.7285e-12 | 29.1347 | -1.055e-10 | 28.9527
-2.2737e-12 | 29.5417 | -1.1278e-10 | 29.3587
-3.8654e-12 | 29.9507 | -1.171e-10 | 29.7627
-3.8654e-12 | 30.3557 | -1.1778e-10 | 30.2017
-6.1391e-12 | 30.7618 | -1.0755e-10 | 30.6268
-1.3642e-12 | 31.1658 | -1.1528e-10 | 31.0328
-1.3642e-12 | 31.5958 | -1.2619e-10 | 31.4378
-3.638e-12 | 32.0008 | -1.08e-10 | 31.8428
-2.7285e-12 | 32.4068 | -1.1232e-10 | 32.2488
-3.8654e-12 | 32.8229 | -1.3256e-10 | 32.6529
-1.3642e-12 | 33.2379 | -1.0959e-10 | 33.0579
-4.0927e-12 | 33.6429 | -1.171e-10 | 33.4619
-5.2296e-12 | 34.0609 | -1.2437e-10 | 33.8659
-2.5011e-12 | 34.466 | -1.2619e-10 | 34.271
-4.7748e-12 | 34.872 | -9.6634e-11 | 34.676
-3.8654e-12 | 35.279 | -9.7771e-11 | 35.082
-4.0927e-12 | 35.686 | -1.1551e-10 | 35.488
-2.5011e-12 | 36.0921 | -1.146e-10 | 35.8931
-2.9559e-12 | 36.4961 | -1.2119e-10 | 36.2991
-5.2296e-12 | 36.9011 | -1.1073e-10 | 36.7031
-5.9117e-12 | 37.3071 | -1.0959e-10 | 37.1071
-4.0927e-12 | 37.7122 | -1.0641e-10 | 37.5112
-2.5011e-12 | 38.1162 | -9.5497e-11 | 37.9152
-6.1391e-12 | 38.5222 | -1.2164e-10 | 38.3252
-3.638e-12 | 38.9272 | -1.196e-10 | 38.7292
-4.0927e-12 | 39.3312 | -1.0755e-10 | 39.1392
-1.5916e-12 | 39.7363 | -1.0891e-10 | 39.5473
-5.0022e-12 | 40.1413 | -1.08e-10 | 39.9523
-4.0927e-12 | 40.5493 | -1.105e-10 | 40.3573
-5.2296e-12 | 40.9833 | -1.105e-10 | 40.7623
-4.7748e-12 | 41.4024 | -1.105e-10 | 41.1764
-3.8654e-12 | 41.8084 | -1.0505e-10 | 41.5824
-2.7285e-12 | 42.2144 | -1.0846e-10 | 41.9874
-4.7748e-12 | 42.6194 | -1.296e-10 | 42.3944
-5.9117e-12 | 43.0235 | -1.0687e-10 | 42.7995
-4.7748e-12 | 43.4455 | -1.0778e-10 | 43.2095
-3.4106e-12 | 43.8505 | -1.2506e-10 | 43.6145
-3.8654e-12 | 44.2555 | -1.2164e-10 | 44.0195
-4.0927e-12 | 44.6615 | -1.13e-10 | 44.4285
-6.1391e-12 | 45.0686 | -1.1528e-10 | 44.8336
-6.8212e-12 | 45.4736 | -1.1414e-10 | 45.2396
-3.638e-12 | 45.8856 | -1.2005e-10 | 45.6436
-5.457e-12 | 46.2946 | -1.2506e-10 | 46.0496
-4.3201e-12 | 46.6997 | -1.121e-10 | 46.4557
-2.5011e-12 | 47.1047 | -1.0209e-10 | 46.8597
-3.638e-12 | 47.5097 | -1.1983e-10 | 47.2647
-4.0927e-12 | 47.9147 | -1.1482e-10 | 47.6697
-5.2296e-12 | 48.3208 | -9.9362e-11 | 48.0737
-3.4106e-12 | 48.7278 | -1.2574e-10 | 48.4788
-3.638e-12 | 49.1368 | -1.2051e-10 | 48.8838
-3.4106e-12 | 49.5418 | -1.1937e-10 | 49.2888
-3.4106e-12 | 49.9459 | -1.3029e-10 | 49.6928
-4.0927e-12 | 50.3499 | -1.0414e-10 | 50.0979
-2.7285e-12 | 50.7549 | -1.1528e-10 | 50.5039
-2.9559e-12 | 51.1599 | -1.1437e-10 | 50.9089
-3.4106e-12 | 51.565 | -9.5952e-11 | 51.3119
-5.457e-12 | 51.968 | -1.146e-10 | 51.718
-3.4106e-12 | 52.373 | -1.08e-10 | 52.123
-5.2296e-12 | 52.778 | -1.1391e-10 | 52.527
-2.9559e-12 | 53.184 | -1.1278e-10 | 52.932
-2.7285e-12 | 53.5891 | -1.0823e-10 | 53.337
-3.4106e-12 | 53.9961 | -1.0687e-10 | 53.7411
-3.4106e-12 | 54.4031 | -1.0778e-10 | 54.1481
-5.9117e-12 | 54.8071 | -1.3006e-10 | 54.5521
-2.9559e-12 | 55.2122 | -1.2506e-10 | 54.9561
-2.5011e-12 | 55.6162 | -1.1164e-10 | 55.3592
-4.0927e-12 | 56.0262 | -1.1437e-10 | 55.7642
-4.3201e-12 | 56.4322 | -1.1005e-10 | 56.1692
-3.8654e-12 | 56.8383 | -1.1573e-10 | 56.5732
-1.5916e-12 | 57.2433 | -1.1437e-10 | 56.9783
-4.3201e-12 | 57.6483 | -1.2506e-10 | 57.3833
-1.1369e-12 | 58.0603 | -1.1687e-10 | 57.7883
-5.0022e-12 | 58.4663 | -1.1164e-10 | 58.1933
-4.0927e-12 | 58.8714 | -1.1914e-10 | 58.5984
-6.1391e-12 | 59.2764 | -1.2233e-10 | 59.0034
-3.638e-12 | 59.6804 | -1.296e-10 | 59.4084
-4.0927e-12 | 60.0864 | -1.1869e-10 | 59.8144
-2.9559e-12 | 60.4915 | -1.2051e-10 | 60.2194
-4.7748e-12 | 60.9025 | -1.1846e-10 | 60.6235
-5.0022e-12 | 61.3075 | -1.171e-10 | 61.0275
-2.5011e-12 | 61.7255 | -1.2983e-10 | 61.4325
-5.0022e-12 | 62.1315 | -1.1073e-10 | 61.8385
-3.8654e-12 | 62.5366 | -1.2233e-10 | 62.2426
-4.7748e-12 | 62.9406 | -1.2324e-10 | 62.6476
-3.8654e-12 | 63.3456 | -1.1528e-10 | 63.0526
-3.8654e-12 | 63.7516 | -1.1073e-10 | 63.4576
-3.8654e-12 | 64.1557 | -1.1005e-10 | 63.8627
-5.457e-12 | 64.5627 | -1.1551e-10 | 64.2687
-4.0927e-12 | 64.9677 | -1.1732e-10 | 64.6747
-4.7748e-12 | 65.3737 | -1.2369e-10 | 65.0807
-3.4106e-12 | 65.7788 | -1.0573e-10 | 65.4848
-1.5916e-12 | 66.1838 | -1.1187e-10 | 65.8908
-3.4106e-12 | 66.5888 | -1.03e-10 | 66.2968
-5.2296e-12 | 66.9938 | -1.1414e-10 | 66.7118
-4.3201e-12 | 67.3988 | -1.221e-10 | 67.1558
-6.1391e-12 | 67.8039 | -9.7998e-11 | 67.5609
-2.5011e-12 | 68.2169 | -1.1232e-10 | 67.9659
-2.9559e-12 | 68.6249 | -1.2483e-10 | 68.3709
-4.3201e-12 | 69.0349 | -1.1096e-10 | 68.7769
-2.7285e-12 | 69.443 | -1.2119e-10 | 69.182
-4.7748e-12 | 69.848 | -1.1096e-10 | 69.585
-5.457e-12 | 70.253 | -1.1278e-10 | 69.99
-5.0022e-12 | 70.659 | -1.1937e-10 | 70.396
-2.7285e-12 | 71.0711 | -1.1437e-10 | 70.8
-4.0927e-12 | 71.4801 | -1.055e-10 | 71.2051
-2.5011e-12 | 71.8861 | -1.0846e-10 | 71.6101
-4.3201e-12 | 72.2941 | -1.1914e-10 | 72.0141
-3.638e-12 | 72.6982 | -1.0959e-10 | 72.4191
-4.0927e-12 | 73.1022 | -1.121e-10 | 72.8242
-5.457e-12 | 73.5202 | -1.3574e-10 | 73.2282
-4.3201e-12 | 73.9252 | -1.0778e-10 | 73.6332
-1.3642e-12 | 74.3303 | -1.2164e-10 | 74.0382
-3.638e-12 | 74.7363 | -1.0209e-10 | 74.4493
-5.0022e-12 | 75.1433 | -1.0368e-10 | 74.8543
-2.5011e-12 | 75.5503 | -1.2756e-10 | 75.2593
-2.9559e-12 | 75.9543 | -1.2096e-10 | 75.6643
-4.7748e-12 | 76.3584 | -1.1687e-10 | 76.0694
-5.0022e-12 | 76.7614 | -1.1346e-10 | 76.4764
-4.0927e-12 | 77.1674 | -1.1323e-10 | 76.8824
-2.9559e-12 | 77.5714 | -1.0505e-10 | 77.2884
-3.8654e-12 | 77.9765 | -1.246e-10 | 77.6934
-4.3201e-12 | 78.3865 | -1.0914e-10 | 78.0965
-2.9559e-12 | 78.8135 | -1.2119e-10 | 78.5025
-2.9559e-12 | 79.2175 | -1.1801e-10 | 78.9075
-2.9559e-12 | 79.6226 | -1.105e-10 | 79.3135
-3.638e-12 | 80.0316 | -1.1664e-10 | 79.7176
-2.9559e-12 | 80.4366 | -1.0959e-10 | 80.1226
-6.5938e-12 | 80.9096 | -1.2187e-10 | 80.5276
-5.2296e-12 | 81.3897 | -1.3097e-10 | 80.9326
-4.7748e-12 | 81.7957 | -1.08e-10 | 81.3367
-3.638e-12 | 82.2447 | -1.1596e-10 | 81.7417
-3.4106e-12 | 82.6507 | -1.1937e-10 | 82.1457
-2.5011e-12 | 83.0648 | -1.1437e-10 | 82.5517
-3.4106e-12 | 83.4778 | -1.0436e-10 | 82.9557
-2.2737e-12 | 83.8838 | -1.3029e-10 | 83.3608
-4.0927e-12 | 84.2898 | -1.2506e-10 | 83.7648
-3.4106e-12 | 84.6948 | -1.055e-10 | 84.1708
-1.5916e-12 | 85.0999
-3.4106e-12 | 85.5059

## 650uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-4.2993e-12 | 5.4629e-12 | -1.612e-10 | 1.0926e-11
GEMs On |  | GEMs On
-5.0022e-12 | 0.31202 | -1.5234e-10 | 0.31102
-5.0022e-12 | 0.99606 | -1.4757e-10 | 0.99306
-3.8654e-12 | 1.4021 | -1.5621e-10 | 1.3991
-4.5475e-12 | 1.8061 | -1.8053e-10 | 1.8061
-5.2296e-12 | 2.2121 | -1.6826e-10 | 2.2111
-2.7285e-12 | 2.6161 | -1.5302e-10 | 2.6412
-3.4106e-12 | 3.0202 | -1.678e-10 | 3.0832
-2.7285e-12 | 3.4242 | -1.4688e-10 | 3.4982
-4.3201e-12 | 3.8292 | -1.8485e-10 | 3.9042
-4.5475e-12 | 4.2472 | -1.6485e-10 | 4.3142
-4.5475e-12 | 4.6533 | -1.4552e-10 | 4.7233
-6.1391e-12 | 5.0583 | -1.6416e-10 | 5.1373
-3.1832e-12 | 5.4623 | -1.5757e-10 | 5.5443
-4.5475e-12 | 5.8673 | -1.4757e-10 | 5.9513
-4.3201e-12 | 6.2794 | -1.6553e-10 | 6.3564
-4.3201e-12 | 6.6944 | -1.6007e-10 | 6.7754
-5.9117e-12 | 7.0984 | -1.5893e-10 | 7.1804
-5.0022e-12 | 7.5034 | -1.6394e-10 | 7.5904
-5.0022e-12 | 7.9095 | -1.6826e-10 | 7.9955
-4.7748e-12 | 8.3155 | -1.528e-10 | 8.4005
-3.8654e-12 | 8.7195 | -1.4688e-10 | 8.8055
-5.6843e-12 | 9.1245 | -1.7076e-10 | 9.2115
-2.0464e-12 | 9.5285 | -1.3847e-10 | 9.6145
-5.6843e-12 | 9.9326 | -1.6394e-10 | 10.0196
-9.0949e-12 | 10.3386 | -1.6689e-10 | 10.4266
-2.2737e-12 | 10.7436 | -1.6007e-10 | 10.8306
-3.1832e-12 | 11.1486 | -1.6098e-10 | 11.2356
-4.7748e-12 | 11.5537 | -1.7076e-10 | 11.6407
-4.5475e-12 | 11.9587 | -1.628e-10 | 12.0467
-5.2296e-12 | 12.3637 | -1.7462e-10 | 12.4517
-5.0022e-12 | 12.7697 | -1.7462e-10 | 12.8567
-3.638e-12 | 13.1738 | -1.8326e-10 | 13.2618
-4.7748e-12 | 13.5798 | -1.4597e-10 | 13.6668
-5.0022e-12 | 13.9838 | -1.487e-10 | 14.0708
-3.4106e-12 | 14.3888 | -1.6689e-10 | 14.4758
-3.8654e-12 | 14.7928 | -1.6462e-10 | 14.8819
-8.6402e-12 | 15.1979 | -1.4802e-10 | 15.2859
-2.0464e-12 | 15.6049 | -1.6166e-10 | 15.6909
-3.8654e-12 | 16.0109 | -1.6212e-10 | 16.0959
-3.8654e-12 | 16.4159 | -1.5893e-10 | 16.5029
-4.3201e-12 | 16.824 | -1.6348e-10 | 16.908
-2.2737e-12 | 17.237 | -1.528e-10 | 17.315
-6.1391e-12 | 17.643 | -1.4757e-10 | 17.721
-4.7748e-12 | 18.062 | -1.528e-10 | 18.127
-2.2737e-12 | 18.4691 | -1.7212e-10 | 18.5321
-3.638e-12 | 18.8741 | -1.6917e-10 | 18.9371
-4.3201e-12 | 19.2801 | -1.5575e-10 | 19.3421
-3.4106e-12 | 19.6861 | -1.4347e-10 | 19.7471
-4.7748e-12 | 20.0922 | -1.5507e-10 | 20.1532
-3.638e-12 | 20.4972 | -1.5416e-10 | 20.5582
-3.1832e-12 | 20.9052 | -1.678e-10 | 20.9692
-1.1369e-12 | 21.3122 | -1.6189e-10 | 21.3742
-5.2296e-12 | 21.7172 | -1.6098e-10 | 21.7793
-3.638e-12 | 22.1253 | -1.6644e-10 | 22.1853
-3.4106e-12 | 22.5303 | -1.5689e-10 | 22.5913
-3.4106e-12 | 22.9363 | -1.6439e-10 | 22.9973
-4.5475e-12 | 23.3433 | -1.4916e-10 | 23.4023
-3.4106e-12 | 23.7484 | -1.4779e-10 | 23.8074
-4.5475e-12 | 24.1654 | -1.5962e-10 | 24.2134
-4.5475e-12 | 24.5714 | -1.6098e-10 | 24.6194
-5.9117e-12 | 24.9774 | -1.6644e-10 | 25.0234
-3.638e-12 | 25.3824 | -1.7963e-10 | 25.4285
-4.3201e-12 | 25.7875 | -1.6894e-10 | 25.8335
-3.8654e-12 | 26.1935 | -1.7212e-10 | 26.2455
-2.0464e-12 | 26.6075 | -1.5893e-10 | 26.6505
-3.638e-12 | 27.0135 | -1.703e-10 | 27.0556
-4.5475e-12 | 27.4186 | -1.8099e-10 | 27.4596
-2.2737e-12 | 27.8236 | -1.6598e-10 | 27.8666
-3.8654e-12 | 28.2366 | -1.5689e-10 | 28.2726
-4.7748e-12 | 28.6426 | -1.5393e-10 | 28.6776
-3.1832e-12 | 29.0477 | -1.4757e-10 | 29.0837
-4.3201e-12 | 29.4717 | -1.5325e-10 | 29.4917
-4.3201e-12 | 29.8777 | -1.5439e-10 | 29.8967
-6.3665e-12 | 30.2837 | -1.5893e-10 | 30.3027
-3.4106e-12 | 30.6898 | -1.4984e-10 | 30.7078
-4.7748e-12 | 31.0948 | -1.3893e-10 | 31.1118
-7.7307e-12 | 31.5008 | -1.5325e-10 | 31.5178
-2.2737e-12 | 31.9088 | -1.6394e-10 | 31.9218
-6.1391e-12 | 32.3169 | -1.6303e-10 | 32.3278
-4.5475e-12 | 32.7219 | -1.7462e-10 | 32.7329
-5.2296e-12 | 33.1269 | -1.6416e-10 | 33.1369
-2.5011e-12 | 33.5309 | -1.8304e-10 | 33.5419
-5.0022e-12 | 33.9379 | -1.5416e-10 | 33.9469
-4.5475e-12 | 34.344 | -1.437e-10 | 34.352
-1.3642e-12 | 34.751 | -1.653e-10 | 34.757
-2.7285e-12 | 35.157 | -1.4847e-10 | 35.161
-4.7748e-12 | 35.564 | -1.7167e-10 | 35.567
-3.638e-12 | 35.9731 | -1.7872e-10 | 35.9721
-5.2296e-12 | 36.3781 | -1.8895e-10 | 36.3771
-2.2737e-12 | 36.7851 | -1.6144e-10 | 36.7811
-3.4106e-12 | 37.1901 | -1.6962e-10 | 37.1861
-2.2737e-12 | 37.5962 | -1.5916e-10 | 37.5911
-3.1832e-12 | 38.0012 | -1.5302e-10 | 37.9972
-4.3201e-12 | 38.4062 | -1.7349e-10 | 38.4032
-3.8654e-12 | 38.8292 | -1.4916e-10 | 38.8082
-2.0464e-12 | 39.2342 | -1.5325e-10 | 39.2132
-3.8654e-12 | 39.6383 | -1.7417e-10 | 39.6643
-3.638e-12 | 40.0423 | -1.6144e-10 | 40.0793
-5.9117e-12 | 40.4483 | -1.5621e-10 | 40.4843
-3.4106e-12 | 40.8553 | -1.703e-10 | 40.8903
-4.3201e-12 | 41.2604 | -1.6439e-10 | 41.2954
-3.4106e-12 | 41.6654 | -1.5757e-10 | 41.7004
-3.4106e-12 | 42.0704 | -1.5643e-10 | 42.1064
-2.5011e-12 | 42.4744 | -1.5325e-10 | 42.5104
-3.4106e-12 | 42.8805 | -1.5234e-10 | 42.9154
-1.819e-12 | 43.2855 | -1.7849e-10 | 43.3265
-3.638e-12 | 43.6925 | -1.6189e-10 | 43.7315
-4.3201e-12 | 44.0985 | -1.5871e-10 | 44.1365
-3.4106e-12 | 44.5065 | -1.6325e-10 | 44.5406
-4.3201e-12 | 44.9236 | -1.628e-10 | 44.9456
-2.7285e-12 | 45.3306 | -1.8053e-10 | 45.3506
-2.5011e-12 | 45.7356 | -1.5802e-10 | 45.7566
-3.4106e-12 | 46.1416 | -1.7826e-10 | 46.1616
-4.7748e-12 | 46.5597 | -1.6416e-10 | 46.5677
-7.7307e-12 | 46.9667 | -1.6712e-10 | 46.9707
-3.8654e-12 | 47.3767 | -1.4211e-10 | 47.3767
-3.4106e-12 | 47.7857 | -1.487e-10 | 47.7807
-2.7285e-12 | 48.1928 | -1.6257e-10 | 48.1858
-3.638e-12 | 48.5988 | -1.4279e-10 | 48.5908
-4.5475e-12 | 49.0108 | -1.4734e-10 | 48.9948
-2.7285e-12 | 49.4348 | -1.7872e-10 | 49.3998
-4.5475e-12 | 49.8428 | -1.4552e-10 | 49.8038
-5.2296e-12 | 50.2469 | -1.5643e-10 | 50.2119
-4.5475e-12 | 50.6519 | -1.7099e-10 | 50.6159
-6.1391e-12 | 51.0569 | -1.5052e-10 | 51.0209
-2.7285e-12 | 51.4619 | -1.6939e-10 | 51.4259
-5.9117e-12 | 51.87 | -1.703e-10 | 51.831
-3.1832e-12 | 52.276 | -1.6553e-10 | 52.235
-2.7285e-12 | 52.682 | -1.5393e-10 | 52.64
-4.3201e-12 | 53.087 | -1.6598e-10 | 53.045
-2.7285e-12 | 53.4931 | -1.5416e-10 | 53.4491
-3.4106e-12 | 53.8971 | -1.728e-10 | 53.8541
-2.0464e-12 | 54.3031 | -1.5507e-10 | 54.2601
-4.3201e-12 | 54.7291 | -1.5439e-10 | 54.6641
-7.276e-12 | 55.1352 | -1.5189e-10 | 55.0692
-3.1832e-12 | 55.5452 | -1.6894e-10 | 55.4762
-5.6843e-12 | 55.9622 | -1.8736e-10 | 55.8812
-8.049e-11 | 56.3692 | -1.7303e-10 | 56.2862
-3.4106e-12 | 56.8813 | -1.6485e-10 | 56.7362
-6.8212e-12 | 57.2873 | -1.6325e-10 | 57.1443
-3.8654e-12 | 57.6923 | -1.7189e-10 | 57.5503
-6.1391e-12 | 58.0973 | -1.3779e-10 | 57.9543
-4.7748e-12 | 58.5053 | -1.6666e-10 | 58.3593
-3.638e-12 | 59.0414 | -1.6894e-10 | 58.8334
-4.3201e-12 | 59.4974 | -1.7849e-10 | 59.2424
-3.1832e-12 | 59.9024 | -1.6212e-10 | 59.6674
-4.7748e-12 | 60.3075 | -1.5393e-10 | 60.0724
-4.5475e-12 | 60.7135 | -1.7189e-10 | 60.4775
-2.7285e-12 | 61.1195 | -1.728e-10 | 60.8825
-3.4106e-12 | 61.5255 | -1.5893e-10 | 61.2875
-2.5011e-12 | 61.9295 | -1.6962e-10 | 61.6955
-3.1832e-12 | 62.3346 | -1.8395e-10 | 62.1026
-4.3201e-12 | 62.7396 | -1.7553e-10 | 62.5076
-3.4106e-12 | 63.1446 | -1.628e-10 | 62.9146
-3.638e-12 | 63.5486 | -1.7303e-10 | 63.3186
-2.2737e-12 | 63.9547 | -1.5143e-10 | 63.7236
-3.638e-12 | 64.3607 | -1.462e-10 | 64.1287
-3.4106e-12 | 64.7657 | -1.6735e-10 | 64.5337
-4.7748e-12 | 65.1707 | -1.5439e-10 | 64.9377
-3.1832e-12 | 65.5757 | -1.7189e-10 | 65.3427
-5.2296e-12 | 65.9808 | -1.5234e-10 | 65.7478
-5.0022e-12 | 66.3858 | -1.7326e-10 | 66.1518
-4.3201e-12 | 66.7898 | -1.4074e-10 | 66.5558
-2.2737e-12 | 67.1968 | -1.5507e-10 | 66.9598
-4.5475e-12 | 67.6029 | -1.8372e-10 | 67.3658
-4.3201e-12 | 68.0089 | -1.4597e-10 | 67.7699
-3.4106e-12 | 68.4149 | -1.4688e-10 | 68.1749
-3.1832e-12 | 68.8199 | -1.6212e-10 | 68.5799
-4.5475e-12 | 69.224 | -1.528e-10 | 68.9849
-6.1391e-12 | 69.631 | -1.6735e-10 | 69.389
-3.4106e-12 | 70.036 | -1.6257e-10 | 69.794
-3.4106e-12 | 70.44 | -1.5325e-10 | 70.199
-5.6843e-12 | 70.8481 | -1.7849e-10 | 70.604
-1.1369e-12 | 71.2531 | -1.5507e-10 | 71.0091
-3.1832e-12 | 71.6581 | -1.4279e-10 | 71.4131
-3.4106e-12 | 72.0621 | -1.6098e-10 | 71.8191
-2.2737e-12 | 72.4671 | -1.6098e-10 | 72.2241
-2.5011e-12 | 72.8752 | -1.7485e-10 | 72.6282
-3.1832e-12 | 73.2822 | -1.6485e-10 | 73.0322
-3.1832e-12 | 73.6922 | -1.6712e-10 | 73.4372
-3.4106e-12 | 74.1002 | -1.5439e-10 | 73.8422
-3.1832e-12 | 74.5063 | -1.537e-10 | 74.2463
-4.3201e-12 | 74.9113 | -1.6962e-10 | 74.6503
-5.2296e-12 | 75.3163 | -1.7099e-10 | 75.0563
-2.2737e-12 | 75.7213 | -1.5302e-10 | 75.4623
-3.638e-12 | 76.1274 | -1.7076e-10 | 75.8673
-1.819e-12 | 76.5324 | -1.7462e-10 | 76.2724
-3.4106e-12 | 76.9374 | -1.5029e-10 | 76.6764
-2.0464e-12 | 77.3404 | -1.5712e-10 | 77.0814
-3.4106e-12 | 77.7455 | -1.728e-10 | 77.4854
-3.8654e-12 | 78.1515 | -1.4916e-10 | 77.8905
-3.4106e-12 | 78.5575 | -1.5189e-10 | 78.2965
-4.7748e-12 | 78.9705 | -1.3461e-10 | 78.7005
-2.2737e-12 | 79.3745 | -1.7349e-10 | 79.1055
-1.819e-12 | 79.7796 | -1.5712e-10 | 79.5096
-3.4106e-12 | 80.1846 | -1.4484e-10 | 79.9146
-2.2737e-12 | 80.5886 | -1.7076e-10 | 80.3196
-3.8654e-12 | 80.9936 | -1.4643e-10 | 80.7266
-3.4106e-12 | 81.3987 | -1.5916e-10 | 81.1316
-3.638e-12 | 81.8037 | -1.5962e-10 | 81.5357
-2.2737e-12 | 82.2087 | -1.4779e-10 | 81.9397
-3.4106e-12 | 82.6127 | -1.6303e-10 | 82.3447
-4.7748e-12 | 83.0178 | -1.4757e-10 | 82.7487
-3.638e-12 | 83.4208 | -1.6848e-10 | 83.1538
-4.3201e-12 | 83.8258 | -1.5871e-10 | 83.5588
-2.7285e-12 | 84.2308
-4.3201e-12 | 84.6358
-2.5011e-12 | 85.0409

## 660uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-5.61e-12 | 1.4536e-12 | -2.276e-10 | 1.555e-11
GEMs On |  | GEMs On
-6.3665e-12 | 0.30902 | -2.4602e-10 | 0.31002
-9.3223e-12 | 0.99406 | -2.3692e-10 | 0.99506
-5.0022e-12 | 1.4001 | -2.3192e-10 | 1.3991
-9.3223e-12 | 1.8061 | -2.135e-10 | 1.8081
-1.0687e-11 | 2.2121 | -2.2419e-10 | 2.2141
-6.3665e-12 | 2.6212 | -2.5625e-10 | 2.6202
-4.7748e-12 | 3.0682 | -2.3306e-10 | 3.0262
-4.5475e-12 | 3.4742 | -2.5875e-10 | 3.4322
-7.0486e-12 | 3.8792 | -2.4897e-10 | 3.8392
-4.5475e-12 | 4.2842 | -2.2965e-10 | 4.2442
-5.9117e-12 | 4.6913 | -2.2646e-10 | 4.6583
-4.3201e-12 | 5.1273 | -2.2283e-10 | 5.0683
-6.8212e-12 | 5.5333 | -2.3283e-10 | 5.4743
-5.6843e-12 | 5.9693 | -2.567e-10 | 5.8793
-4.7748e-12 | 6.4194 | -2.3397e-10 | 6.2834
-7.0486e-12 | 6.8254 | -2.0782e-10 | 6.6924
-5.9117e-12 | 7.2304 | -2.2828e-10 | 7.0964
-6.1391e-12 | 7.6444 | -2.3351e-10 | 7.5044
-4.5475e-12 | 8.1345 | -2.3874e-10 | 7.9105
-4.3201e-12 | 8.5405 | -2.1419e-10 | 8.3175
-3.8654e-12 | 8.9455 | -2.5079e-10 | 8.7325
-5.0022e-12 | 9.3515 | -2.4056e-10 | 9.1375
-4.5475e-12 | 9.7566 | -2.2919e-10 | 9.5425
-5.457e-12 | 10.1616 | -2.201e-10 | 9.9496
-4.3201e-12 | 10.5676 | -2.0464e-10 | 10.3536
-5.0022e-12 | 10.9716 | -2.2624e-10 | 10.7616
-5.457e-12 | 11.3776 | -2.3965e-10 | 11.1676
-8.4128e-12 | 11.7837 | -2.1214e-10 | 11.5727
-5.457e-12 | 12.1867 | -2.2897e-10 | 11.9797
-5.0022e-12 | 12.5927 | -2.3533e-10 | 12.3847
-4.7748e-12 | 12.9967 | -1.8667e-10 | 12.7907
-5.6843e-12 | 13.4028 | -2.01e-10 | 13.1958
-5.457e-12 | 13.8068 | -2.4761e-10 | 13.6028
-5.457e-12 | 14.2118 | -2.251e-10 | 14.0078
-8.4128e-12 | 14.6168 | -2.2851e-10 | 14.4138
-4.7748e-12 | 15.0209 | -2.1532e-10 | 14.8259
-6.8212e-12 | 15.4259 | -2.2897e-10 | 15.3269
-8.6402e-12 | 15.8309 | -2.4329e-10 | 15.7309
-5.0022e-12 | 16.2359 | -2.3169e-10 | 16.1399
-5.0022e-12 | 16.6399 | -2.1373e-10 | 16.95
-5.6843e-12 | 17.045 | -2.3488e-10 | 17.357
-4.5475e-12 | 17.449 | -2.1987e-10 | 17.762
-5.457e-12 | 17.855 | -2.5466e-10 | 18.167
-8.1855e-12 | 18.26 | -2.558e-10 | 18.7861
-5.6843e-12 | 18.6661 | -2.0759e-10 | 19.3801
-6.1391e-12 | 19.0711 | -2.442e-10 | 19.7851
-6.3665e-12 | 19.4751 | -2.2555e-10 | 20.1912
-5.6843e-12 | 19.8801 | -2.2533e-10 | 20.5962
-4.7748e-12 | 20.2852 | -2.2123e-10 | 21.0002
-5.0022e-12 | 20.6902 | -2.2396e-10 | 21.4102
-3.8654e-12 | 21.0962 | -2.1237e-10 | 21.8382
-5.0022e-12 | 21.5012 | -2.4102e-10 | 22.2443
-3.8654e-12 | 21.9062 | -2.4079e-10 | 22.6493
-6.3665e-12 | 22.3113 | -2.2237e-10 | 23.0533
-5.6843e-12 | 22.7173 | -2.3715e-10 | 23.4583
-7.0486e-12 | 23.1243 | -2.11e-10 | 23.8634
-5.457e-12 | 23.5294 | -2.1055e-10 | 24.2684
-3.638e-12 | 23.9344 | -2.1669e-10 | 24.6734
-4.7748e-12 | 24.3394 | -2.11e-10 | 25.0784
-4.5475e-12 | 24.7434 | -2.3874e-10 | 25.4825
-6.3665e-12 | 25.1484 | -2.251e-10 | 25.8865
-4.3201e-12 | 25.5535 | -2.6148e-10 | 26.3045
-6.3665e-12 | 25.9585 | -2.1555e-10 | 26.8475
-5.457e-12 | 26.3625 | -2.3715e-10 | 27.3266
-2.9559e-12 | 26.9055 | -2.4988e-10 | 27.7316
-4.5475e-12 | 27.3126 | -2.3988e-10 | 28.1386
-4.7748e-12 | 27.7196 | -2.1851e-10 | 28.5446
-6.8212e-12 | 28.1246 | -2.3829e-10 | 28.9487
-7.0486e-12 | 28.5286 | -2.7126e-10 | 29.3817
-6.8212e-12 | 29.1927 | -2.2851e-10 | 29.7867
-7.7307e-12 | 30.0257 | -2.367e-10 | 30.1927
-5.457e-12 | 30.4317 | -2.1646e-10 | 30.5978
-3.8654e-12 | 30.8378 | -2.1123e-10 | 31.0028
-6.1391e-12 | 31.2438 | -2.2033e-10 | 31.4088
-5.457e-12 | 31.7658 | -2.0896e-10 | 31.8208
-7.7307e-12 | 32.1708 | -2.16e-10 | 32.2258
-6.3665e-12 | 32.5769 | -2.0668e-10 | 32.6319
-1.0004e-11 | 32.9819 | -2.367e-10 | 33.0379
-6.8212e-12 | 33.3869 | -2.1419e-10 | 33.4439
-6.1391e-12 | 33.7989 | -2.2783e-10 | 33.8469
-8.1855e-12 | 34.204 | -2.151e-10 | 34.251
-7.5033e-12 | 34.615 | -2.1555e-10 | 34.656
-3.8654e-12 | 35.02 | -2.1259e-10 | 35.062
-3.638e-12 | 35.434 | -2.3215e-10 | 35.467
-3.638e-12 | 35.8411 | -2.3078e-10 | 35.873
-4.7748e-12 | 36.2461 | -2.1737e-10 | 36.3101
-6.8212e-12 | 36.6521 | -2.16e-10 | 36.7171
-5.9117e-12 | 37.0591 | -2.1396e-10 | 37.1221
-5.0022e-12 | 37.4641 | -2.3147e-10 | 37.5262
-4.5475e-12 | 37.8682 | -2.0509e-10 | 37.9492
-6.3665e-12 | 38.2732 | -2.1305e-10 | 38.3552
-3.8654e-12 | 38.6782 | -2.467e-10 | 38.7602
-4.7748e-12 | 39.0832 | -2.417e-10 | 39.1652
-4.7748e-12 | 39.4863 | -2.201e-10 | 39.5693
-6.3665e-12 | 39.8933 | -2.1123e-10 | 39.9893
-4.5475e-12 | 40.2983 | -2.3306e-10 | 40.3943
-3.638e-12 | 40.7033 | -2.3192e-10 | 40.8003
-5.9117e-12 | 41.1084 | -2.1646e-10 | 41.2054
-4.7748e-12 | 41.5134 | -2.4988e-10 | 41.6104
-5.457e-12 | 41.9194 | -2.1487e-10 | 42.0154
-5.9117e-12 | 42.3244 | -2.0259e-10 | 42.4214
-3.8654e-12 | 42.7294 | -2.5602e-10 | 42.8304
-4.5475e-12 | 43.1355 | -2.3351e-10 | 43.2365
-6.8212e-12 | 43.5415 | -2.1737e-10 | 43.6435
-2.2737e-12 | 43.9455 | -2.0918e-10 | 44.0485
-9.3223e-12 | 44.3505 | -2.135e-10 | 44.4555
-5.457e-12 | 44.7566 | -2.2555e-10 | 44.8646
-5.457e-12 | 45.1626 | -2.276e-10 | 45.2726
-6.1391e-12 | 45.5746 | -2.417e-10 | 45.6786
-5.457e-12 | 45.9786 | -2.2396e-10 | 46.0826
-2.9559e-12 | 46.3826 | -2.2033e-10 | 46.4917
-3.8654e-12 | 46.7917 | -2.5148e-10 | 46.8967
-9.5497e-12 | 47.1957 | -2.4374e-10 | 47.3017
-8.1855e-12 | 47.6017 | -2.3488e-10 | 47.7077
-4.5475e-12 | 48.0067 | -2.3033e-10 | 48.1127
-4.7748e-12 | 48.4128 | -2.1896e-10 | 48.5168
-5.9117e-12 | 48.8168 | -2.3397e-10 | 48.9218
-4.5475e-12 | 49.2218 | -2.392e-10 | 49.3278
-6.1391e-12 | 49.6268 | -2.2692e-10 | 49.7328
-3.8654e-12 | 50.0319 | -2.1851e-10 | 50.1389
-6.3665e-12 | 50.4369 | -2.2192e-10 | 50.5439
-4.5475e-12 | 50.8429 | -2.1896e-10 | 50.9479
-4.7748e-12 | 51.2489 | -2.4806e-10 | 51.3549
-4.7748e-12 | 51.654 | -2.3738e-10 | 51.76
-5.9117e-12 | 52.059 | -2.2783e-10 | 52.165
-7.0486e-12 | 52.468 | -2.0145e-10 | 52.573
-6.1391e-12 | 52.873 | -2.0009e-10 | 52.981
-6.3665e-12 | 53.2781 | -2.4716e-10 | 53.4391
-4.7748e-12 | 53.6831 | -2.4238e-10 | 53.8441
-6.1391e-12 | 54.0891 | -2.2192e-10 | 54.2491
-6.1391e-12 | 54.4961 | -2.0464e-10 | 54.6551
-6.8212e-12 | 54.9031 | -2.3692e-10 | 55.0602
-4.5475e-12 | 55.3092 | -1.9713e-10 | 55.4652
-4.5475e-12 | 55.7182 | -2.0714e-10 | 55.8692
-5.457e-12 | 56.1232 | -2.0373e-10 | 56.2752
-1.0004e-11 | 56.5302 | -2.3192e-10 | 56.6882
-8.1855e-12 | 56.9393 | -1.9486e-10 | 57.0943
-3.8654e-12 | 57.3463 | -2.6307e-10 | 57.5003
-3.638e-12 | 57.7513 | -2.085e-10 | 57.9063
-7.0486e-12 | 58.1563 | -2.2715e-10 | 58.3113
-6.1391e-12 | 58.5643 | -2.035e-10 | 58.7154
-2.9559e-12 | 59.0094 | -2.4261e-10 | 59.1254
-5.9117e-12 | 59.4134 | -2.467e-10 | 59.5334
-4.7748e-12 | 59.8194 | -2.2305e-10 | 59.9394
-5.6843e-12 | 60.2244 | -2.2237e-10 | 60.3445
-5.6843e-12 | 60.6295 | -2.2305e-10 | 60.7495
-4.3201e-12 | 61.0345 | -2.1419e-10 | 61.1565
-5.0022e-12 | 61.4405 | -2.2715e-10 | 61.5625
-3.1832e-12 | 61.8455 | -2.2965e-10 | 61.9675
-5.9117e-12 | 62.2506 | -2.1259e-10 | 62.3726
-4.3201e-12 | 62.6546 | -2.2078e-10 | 62.7776
-5.6843e-12 | 63.0596 | -2.2419e-10 | 63.1826
-5.457e-12 | 63.4676 | -2.4852e-10 | 63.5916
-4.3201e-12 | 63.8727 | -2.5261e-10 | 64.0007
-3.638e-12 | 64.2787 | -2.1646e-10 | 64.4057
-6.3665e-12 | 64.6847 | -2.3874e-10 | 64.8147
-5.6843e-12 | 65.0897 | -2.2396e-10 | 65.2237
-6.1391e-12 | 65.4958 | -2.1851e-10 | 65.6287
-5.6843e-12 | 65.9108 | -1.9895e-10 | 66.0338
-5.0022e-12 | 66.3158 | -2.2396e-10 | 66.4428
-8.4128e-12 | 66.7218 | -2.3101e-10 | 66.8528
-6.1391e-12 | 67.1268 | -2.6171e-10 | 67.2569
-7.276e-12 | 67.5309 | -2.4443e-10 | 67.6619
-4.3201e-12 | 67.9369 | -2.3715e-10 | 68.0659
-4.7748e-12 | 68.3419 | -2.0827e-10 | 68.6349
-7.5033e-12 | 68.7469 | -2.3647e-10 | 69.174
-4.3201e-12 | 69.153 | -2.0259e-10 | 69.587
-3.1832e-12 | 69.558 | -2.3147e-10 | 69.992
-7.0486e-12 | 69.965 | -2.4079e-10 | 70.401
-6.3665e-12 | 70.372 | -2.3988e-10 | 70.9081
-7.276e-12 | 70.778 | -2.2419e-10 | 71.3461
-6.3665e-12 | 71.1841 | -2.16e-10 | 71.7511
-6.3665e-12 | 71.5891 | -2.3942e-10 | 72.1571
-6.1391e-12 | 71.9951 | -2.1805e-10 | 72.5652
-5.457e-12 | 72.4001 | -2.2442e-10 | 72.9692
-5.9117e-12 | 72.8052 | -2.3465e-10 | 73.3742
-3.638e-12 | 73.2122 | -2.5716e-10 | 73.7802
-8.4128e-12 | 73.6182 | -1.8963e-10 | 74.1852
-5.9117e-12 | 74.0232 | -2.5284e-10 | 74.5893
-4.7748e-12 | 74.4293 | -2.1987e-10 | 74.9933
-4.5475e-12 | 74.8363 | -2.2578e-10 | 75.3983
-4.7748e-12 | 75.2423 | -2.3601e-10 | 75.8033
-4.3201e-12 | 75.6473 | -2.4988e-10 | 76.2084
-4.3201e-12 | 76.0524 | -2.5466e-10 | 76.6134
-4.3201e-12 | 76.4594 | -2.251e-10 | 77.0174
-5.6843e-12 | 76.8654 | -2.2715e-10 | 77.4214
-4.7748e-12 | 77.2704 | -2.301e-10 | 77.8264
-9.5497e-12 | 77.6774 | -2.2715e-10 | 78.2315
-5.9117e-12 | 78.0855 | -2.3647e-10 | 78.6375
-5.9117e-12 | 78.4905 | -2.0896e-10 | 79.0425
-5.9117e-12 | 78.8955 | -2.2078e-10 | 79.4465
-5.0022e-12 | 79.3055 | -2.1009e-10 | 79.8516
-6.1391e-12 | 79.7136 | -2.4875e-10 | 80.2566
-6.1391e-12 | 80.1196 | -2.2897e-10 | 80.6606
-5.457e-12 | 80.5266 | -2.4511e-10 | 81.0646
-5.9117e-12 | 80.9336 | -2.4193e-10 | 81.4697
-2.5011e-12 | 81.3387 | -2.1055e-10 | 81.8747
-5.457e-12 | 81.7437 | -2.0896e-10 | 82.2797
-2.9559e-12 | 82.1487 | -2.2669e-10 | 82.6837
-5.457e-12 | 82.5547 | -2.3715e-10 | 83.0888
-6.3665e-12 | 82.9608 | -2.1896e-10 | 83.4938
-6.1391e-12 | 83.3668 | -2.2783e-10 | 83.8988
-5.0022e-12 | 83.7718 | -2.2442e-10 | 84.3038
-6.8212e-12 | 84.1768 | -2.3806e-10 | 84.7088
-6.1391e-12 | 84.5808
-4.5475e-12 | 84.9859
-4.5475e-12 | 85.3899
-5.457e-12 | 85.7939
-4.5475e-12 | 86.1999
-5.9117e-12 | 86.6069
-5.0022e-12 | 87.013
-5.0022e-12 | 87.417
-6.8212e-12 | 87.822
-5.6843e-12 | 88.228

## 670uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-6.4082e-12 | 1.6449e-12 | -3.1701e-10 | 2.274e-11
GEMs On |  | GEMs On
-9.3223e-12 | 0.31602 | -3.0536e-10 | 0.31002
-5.9117e-12 | 1.0001 | -3.4083e-10 | 0.99306
-5.457e-12 | 1.4051 | -3.2696e-10 | 1.3981
-5.6843e-12 | 1.8091 | -3.0059e-10 | 1.8021
-7.0486e-12 | 2.2141 | -3.3333e-10 | 2.2081
-5.2296e-12 | 2.6191 | -3.6812e-10 | 2.6131
-5.6843e-12 | 3.0232 | -3.1014e-10 | 3.0192
-6.1391e-12 | 3.4282 | -3.3879e-10 | 3.4252
-5.2296e-12 | 3.8332 | -3.1969e-10 | 3.8292
-5.9117e-12 | 4.2392 | -2.8331e-10 | 4.2342
-8.1855e-12 | 4.6443 | -3.1287e-10 | 4.6403
-7.276e-12 | 5.0493 | -3.4493e-10 | 5.0453
-5.9117e-12 | 5.4533 | -3.1378e-10 | 5.4503
-4.3201e-12 | 5.8583 | -3.5402e-10 | 5.8553
-5.2296e-12 | 6.2634 | -2.7057e-10 | 6.2594
-5.457e-12 | 6.6684 | -3.3356e-10 | 6.6644
-9.5497e-12 | 7.0734 | -3.4629e-10 | 7.0684
-6.8212e-12 | 7.4784 | -3.5107e-10 | 7.4724
-5.6843e-12 | 7.8845 | -3.1946e-10 | 7.8765
-6.8212e-12 | 8.2885 | -3.3356e-10 | 8.2815
-1.1141e-11 | 8.6945 | -3.4197e-10 | 8.6855
-1.0459e-11 | 9.1005 | -3.3879e-10 | 9.0895
-7.0486e-12 | 9.5125 | -3.5561e-10 | 9.4935
-7.7307e-12 | 9.9186 | -2.9672e-10 | 9.8996
-5.2296e-12 | 10.3226 | -3.3833e-10 | 10.3056
-7.7307e-12 | 10.7316 | -3.015e-10 | 10.7106
-4.5475e-12 | 11.1386 | -3.3583e-10 | 11.1156
-5.457e-12 | 11.5447 | -3.2014e-10 | 11.5197
-5.457e-12 | 11.9517 | -2.9354e-10 | 11.9247
-5.2296e-12 | 12.3587 | -3.2924e-10 | 12.3297
-5.9117e-12 | 12.7627 | -3.0423e-10 | 12.7347
-5.2296e-12 | 13.1707 | -3.4515e-10 | 13.1408
-6.5938e-12 | 13.6158 | -3.09e-10 | 13.5458
-7.276e-12 | 14.0318 | -2.7535e-10 | 13.9498
-5.9117e-12 | 14.4368 | -3.165e-10 | 14.3538
-4.0927e-12 | 14.8459 | -3.256e-10 | 14.7578
-7.7307e-12 | 15.2509 | -2.9127e-10 | 15.1629
-5.457e-12 | 15.6569 | -2.9149e-10 | 15.5679
-6.1391e-12 | 16.0649 | -3.0559e-10 | 15.9719
-3.8654e-12 | 16.4709 | -3.0263e-10 | 16.3779
-7.276e-12 | 16.876 | -2.949e-10 | 16.782
-4.5475e-12 | 17.281 | -2.6444e-10 | 17.187
-8.6402e-12 | 17.686 | -3.0832e-10 | 17.592
-5.6843e-12 | 18.091 | -3.6425e-10 | 17.996
-6.1391e-12 | 18.4961 | -3.181e-10 | 18.4001
-4.7748e-12 | 18.9011 | -3.0946e-10 | 18.8071
-6.8212e-12 | 19.3071 | -2.9422e-10 | 19.2111
-7.276e-12 | 19.7131 | -3.1537e-10 | 19.6151
-1.0459e-11 | 20.1192 | -2.9218e-10 | 20.0201
-6.5938e-12 | 20.5262 | -3.3947e-10 | 20.4252
-1.1141e-11 | 20.9322 | -3.0991e-10 | 20.8302
-5.9117e-12 | 21.3392 | -3.1719e-10 | 21.2342
-7.9581e-12 | 21.7452 | -2.8581e-10 | 21.6382
-8.1855e-12 | 22.1513 | -3.663e-10 | 22.0423
-5.2296e-12 | 22.5553 | -3.1855e-10 | 22.4463
-5.2296e-12 | 22.9613 | -3.115e-10 | 22.8513
-4.3201e-12 | 23.3673 | -3.6061e-10 | 23.2563
-5.2296e-12 | 23.7734 | -2.8399e-10 | 23.6604
-8.6402e-12 | 24.1774 | -3.588e-10 | 24.0664
-7.7307e-12 | 24.5824 | -3.2423e-10 | 24.4714
-5.457e-12 | 24.9864 | -3.14e-10 | 24.8754
-6.5938e-12 | 25.3914 | -3.3106e-10 | 25.2795
-8.1855e-12 | 25.7965 | -2.7239e-10 | 25.6855
-9.3223e-12 | 26.2005 | -3.2105e-10 | 26.0905
-6.8212e-12 | 26.6045 | -2.7012e-10 | 26.4945
-9.5497e-12 | 27.0115 | -3.2196e-10 | 26.8995
-1.0004e-11 | 27.4156 | -3.5175e-10 | 27.3056
-6.5938e-12 | 27.8226 | -3.2151e-10 | 27.7096
-8.6402e-12 | 28.2316 | -3.09e-10 | 28.1156
-6.5938e-12 | 28.6366 | -2.8649e-10 | 28.5206
-5.6843e-12 | 29.0437 | -2.9786e-10 | 28.9246
-8.4128e-12 | 29.4467 | -3.0059e-10 | 29.3287
-8.1855e-12 | 29.8527 | -3.0627e-10 | 29.7327
-6.8212e-12 | 30.2587 | -3.1673e-10 | 30.1367
-9.5497e-12 | 30.6648 | -3.4015e-10 | 30.5418
-5.9117e-12 | 31.0688 | -3.1173e-10 | 30.9468
-7.7307e-12 | 31.4728 | -2.9809e-10 | 31.3528
-1.0459e-11 | 31.8778 | -3.422e-10 | 31.7568
-4.7748e-12 | 32.2839 | -3.0741e-10 | 32.1628
-6.1391e-12 | 32.6939 | -3.2605e-10 | 32.5699
-8.1855e-12 | 33.0999 | -3.0104e-10 | 32.9729
-5.6843e-12 | 33.5069 | -3.1332e-10 | 33.3789
-6.8212e-12 | 33.9129 | -3.3265e-10 | 33.7839
-4.7748e-12 | 34.318 | -2.9581e-10 | 34.188
-5.6843e-12 | 34.724 | -3.1719e-10 | 34.594
-8.1855e-12 | 35.132 | -2.9354e-10 | 34.998
-2.2737e-12 | 35.537 | -3.2992e-10 | 35.404
-4.0927e-12 | 35.9451 | -3.4083e-10 | 35.8081
-6.1391e-12 | 36.3511 | -3.1855e-10 | 36.2131
-8.6402e-12 | 36.7551 | -3.2742e-10 | 36.6171
-4.7748e-12 | 37.1611 | -3.1173e-10 | 37.0231
-7.9581e-12 | 37.5662 | -3.1082e-10 | 37.4271
-4.3201e-12 | 37.9712 | -3.2333e-10 | 37.8312
-7.276e-12 | 38.3752 | -3.6403e-10 | 38.2352
-5.9117e-12 | 38.7802 | -3.3219e-10 | 38.6402
-7.0486e-12 | 39.1872 | -3.1855e-10 | 39.0452
-4.5475e-12 | 39.5933 | -3.1446e-10 | 39.4493
-6.1391e-12 | 39.9973 | -3.4288e-10 | 39.8533
-7.276e-12 | 40.4033 | -3.0786e-10 | 40.2573
-5.9117e-12 | 40.8083 | -3.2082e-10 | 40.6633
-4.7748e-12 | 41.2134 | -2.9854e-10 | 41.0673
-6.1391e-12 | 41.6174 | -3.0764e-10 | 41.4714
-5.457e-12 | 42.0244 | -3.2742e-10 | 41.8784
-7.7307e-12 | 42.4314 | -3.2173e-10 | 42.2844
-5.9117e-12 | 42.8364 | -3.4038e-10 | 42.6894
-6.8212e-12 | 43.2415 | -3.2924e-10 | 43.0935
-4.5475e-12 | 43.6465 | -3.0923e-10 | 43.4975
-6.1391e-12 | 44.0535 | -3.1719e-10 | 43.9015
-4.0927e-12 | 44.4575 | -2.9468e-10 | 44.3065
-5.457e-12 | 44.8636 | -3.2105e-10 | 44.7116
-4.7748e-12 | 45.2676 | -3.422e-10 | 45.1176
-9.0949e-12 | 45.6726 | -2.949e-10 | 45.5236
-7.9581e-12 | 46.0776 | -3.1037e-10 | 45.9276
-5.457e-12 | 46.4827 | -2.8763e-10 | 46.3336
-5.457e-12 | 46.8857 | -3.2924e-10 | 46.7387
-5.457e-12 | 47.2937 | -3.5288e-10 | 47.1427
-4.7748e-12 | 47.6997 | -3.1332e-10 | 47.5467
-8.4128e-12 | 48.1048 | -2.9786e-10 | 47.9507
-5.6843e-12 | 48.5088 | -3.2696e-10 | 48.3568
-5.2296e-12 | 48.9148 | -3.2014e-10 | 48.7628
-3.8654e-12 | 49.3198 | -3.447e-10 | 49.1688
-8.4128e-12 | 49.7338 | -3.1423e-10 | 49.5738
-5.457e-12 | 50.1399 | -3.2287e-10 | 49.9789
-7.276e-12 | 50.5469 | -3.1173e-10 | 50.3829
-4.0927e-12 | 50.9539 | -3.4515e-10 | 50.7879
-1.3188e-11 | 51.3589 | -3.281e-10 | 51.1929
-7.7307e-12 | 51.765 | -3.1764e-10 | 51.5969
-3.1832e-12 | 52.17 | -3.3879e-10 | 52.003
-7.0486e-12 | 52.574 | -3.1719e-10 | 52.408
-7.7307e-12 | 52.979 | -3.2742e-10 | 52.814
-7.276e-12 | 53.388 | -2.9672e-10 | 53.219
-6.1391e-12 | 53.7931 | -3.4834e-10 | 53.6251
-6.1391e-12 | 54.1981 | -2.9763e-10 | 54.0301
-4.7748e-12 | 54.6041 | -2.7626e-10 | 54.4341
-8.1855e-12 | 55.0082 | -3.0263e-10 | 54.8391
-2.9559e-12 | 55.4142 | -2.5148e-10 | 55.2432
-7.276e-12 | 55.8192 | -2.9218e-10 | 55.6482
-6.8212e-12 | 56.2242 | -3.3106e-10 | 56.0532
-6.5938e-12 | 56.6292 | -3.3583e-10 | 56.4572
-4.5475e-12 | 57.0323 | -3.2014e-10 | 56.8623
-5.6843e-12 | 57.4373 | -2.9809e-10 | 57.2673
-7.276e-12 | 57.8433 | -3.2856e-10 | 57.6713
-4.3201e-12 | 58.2493 | -3.2856e-10 | 58.0773
-5.457e-12 | 58.6553 | -3.09e-10 | 58.4823
-3.8654e-12 | 59.0594 | -3.115e-10 | 58.8884
-9.0949e-12 | 59.4644 | -3.497e-10 | 59.2924
-7.0486e-12 | 59.8704 | -3.1946e-10 | 59.6984
-5.457e-12 | 60.2745 | -3.5175e-10 | 60.1034
-5.457e-12 | 60.6795 | -3.0172e-10 | 60.5095
-6.1391e-12 | 61.0845 | -3.206e-10 | 60.9135
-8.1855e-12 | 61.4885 | -3.0514e-10 | 61.3185
-1.0004e-11 | 61.8925 | -3.0059e-10 | 61.7235
-7.0486e-12 | 62.2996 | -3.3651e-10 | 62.1285
-6.5938e-12 | 62.7036 | -3.015e-10 | 62.5346
-5.457e-12 | 63.1106 | -3.1787e-10 | 62.9396
-4.0927e-12 | 63.5166 | -3.3651e-10 | 63.3436
-7.7307e-12 | 63.9267 | -3.1014e-10 | 63.7486
-5.6843e-12 | 64.3327 | -2.8695e-10 | 64.1527
-5.9117e-12 | 64.7367 | -3.04e-10 | 64.5567
-6.8212e-12 | 65.1437 | -3.2674e-10 | 64.9607
-6.8212e-12 | 65.5477 | -3.5834e-10 | 65.3657
-5.9117e-12 | 65.9528 | -3.3901e-10 | 65.7708
-4.0927e-12 | 66.3558 | -3.4879e-10 | 66.1758
-6.1391e-12 | 66.7608 | -2.9422e-10 | 66.5798
-6.8212e-12 | 67.1648 | -3.0445e-10 | 66.9858
-7.9581e-12 | 67.5699 | -3.14e-10 | 67.3909
-5.457e-12 | 67.9749 | -3.2946e-10 | 67.7949
-4.0927e-12 | 68.3799 | -2.8649e-10 | 68.1999
-6.5938e-12 | 68.7839 | -3.0764e-10 | 68.6039
-7.276e-12 | 69.189 | -3.3992e-10 | 69.0079
-4.0927e-12 | 69.595 | -3.2605e-10 | 69.412
-5.6843e-12 | 69.999 | -2.8126e-10 | 69.816
-6.5938e-12 | 70.404 | -2.8194e-10 | 70.221
-5.6843e-12 | 70.8101 | -3.1309e-10 | 70.626
-7.276e-12 | 71.2151 | -3.1969e-10 | 71.0301
-5.457e-12 | 71.6201 | -3.6175e-10 | 71.4351
-4.7748e-12 | 72.0261 | -2.6603e-10 | 71.8391
-7.7307e-12 | 72.4311 | -3.0627e-10 | 72.2441
-5.9117e-12 | 72.8352 | -2.9786e-10 | 72.6482
-5.457e-12 | 73.2412 | -3.14e-10 | 73.0532
-8.1855e-12 | 73.6462 | -2.7353e-10 | 73.4572
-5.9117e-12 | 74.0512 | -2.8899e-10 | 73.8632
-6.5938e-12 | 74.4553 | -2.8353e-10 | 74.2682
-7.0486e-12 | 74.8613 | -3.0218e-10 | 74.6733
-4.7748e-12 | 75.2663 | -3.5402e-10 | 75.0763
-6.8212e-12 | 75.6703 | -3.14e-10 | 75.4823
-5.6843e-12 | 76.0754 | -2.7558e-10 | 75.8873
-4.5475e-12 | 76.4814 | -3.1127e-10 | 76.2934
-5.2296e-12 | 76.8874 | -3.1832e-10 | 76.6984
-7.276e-12 | 77.2914 | -3.6835e-10 | 77.1034
-5.457e-12 | 77.6974 | -3.4902e-10 | 77.5084
-7.9581e-12 | 78.1015 | -2.9922e-10 | 77.9145
-5.457e-12 | 78.5065 | -3.4993e-10 | 78.3195
-6.5938e-12 | 78.9115 | -3.2151e-10 | 78.7245
-4.3201e-12 | 79.3165 | -3.3197e-10 | 79.1285
-5.457e-12 | 79.7226 | -2.7467e-10 | 79.5326
-8.1855e-12 | 80.1286 | -3.1309e-10 | 79.9376
-5.6843e-12 | 80.5346 | -2.9968e-10 | 80.3426
-5.6843e-12 | 80.9386 | -2.9422e-10 | 80.7486
-5.6843e-12 | 81.3427 | -3.1287e-10 | 81.1536
-6.1391e-12 | 81.7467 | -3.0536e-10 | 81.5587
-6.1391e-12 | 82.1517 | -3.2537e-10 | 81.9627
-6.8212e-12 | 82.5577 | -3.522e-10 | 82.3677
-6.8212e-12 | 82.9627 | -3.1582e-10 | 82.7727
-6.1391e-12 | 83.3688 | -3.1832e-10 | 83.1768
-4.7748e-12 | 83.7748 | -2.949e-10 | 83.5818
-5.9117e-12 | 84.1798 | -3.4834e-10 | 83.9868

## 680uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-8.3928e-12 | 2.1697e-12 | -4.5082e-10 | 3.4871e-11
GEMs On |  | GEMs On
-8.6402e-12 | 0.31202 | -4.4679e-10 | 0.30902
-7.7307e-12 | 0.99606 | -4.3087e-10 | 0.99206
-7.0486e-12 | 1.4031 | -5.2341e-10 | 1.3971
-1.0004e-11 | 1.8081 | -4.5316e-10 | 1.8021
-7.7307e-12 | 2.2141 | -4.7703e-10 | 2.2061
-7.5033e-12 | 2.6181 | -5.0795e-10 | 2.6111
-9.3223e-12 | 3.0242 | -4.4156e-10 | 3.0162
-7.0486e-12 | 3.4282 | -3.8108e-10 | 3.4202
-3.8654e-12 | 3.8322 | -4.8067e-10 | 3.8252
-6.5938e-12 | 4.2382 | -4.4861e-10 | 4.2282
-7.7307e-12 | 4.6423 | -4.9181e-10 | 4.6343
-6.8212e-12 | 5.0463 | -4.2382e-10 | 5.0383
-8.6402e-12 | 5.4513 | -4.8362e-10 | 5.4433
-5.2296e-12 | 5.8563 | -4.3929e-10 | 5.8493
-7.7307e-12 | 6.2614 | -4.336e-10 | 6.2544
-1.4097e-11 | 6.6664 | -4.8544e-10 | 6.6604
-8.1855e-12 | 7.0724 | -3.7494e-10 | 7.0644
-5.2296e-12 | 7.4774 | -3.8472e-10 | 7.4704
-8.6402e-12 | 7.8835 | -4.1405e-10 | 7.8825
-6.8212e-12 | 8.2885 | -4.3838e-10 | 8.2865
-8.1855e-12 | 8.6935 | -4.4088e-10 | 8.6905
-5.6843e-12 | 9.0975 | -4.5202e-10 | 9.0945
-9.0949e-12 | 9.5025 | -4.3019e-10 | 9.5005
-8.1855e-12 | 9.9086 | -4.9545e-10 | 9.9066
-8.8676e-12 | 10.3136 | -4.1427e-10 | 10.3096
-7.5033e-12 | 10.7186 | -4.527e-10 | 10.7146
-1.4325e-11 | 11.1236 | -4.7771e-10 | 11.1196
-7.9581e-12 | 11.5287 | -4.2769e-10 | 11.5237
-7.0486e-12 | 11.9337 | -5.0659e-10 | 11.9287
-5.2296e-12 | 12.3387 | -4.02e-10 | 12.3337
-5.457e-12 | 12.7437 | -4.7976e-10 | 12.7387
-1.2733e-11 | 13.1487 | -4.2769e-10 | 13.1427
-5.2296e-12 | 13.5538 | -4.1018e-10 | 13.5458
-6.1391e-12 | 13.9578 | -4.261e-10 | 13.9508
-7.7307e-12 | 14.3628 | -4.2223e-10 | 14.3558
-7.0486e-12 | 14.7668 | -4.361e-10 | 14.7598
-6.8212e-12 | 15.1729 | -4.3792e-10 | 15.1649
-7.0486e-12 | 15.5769 | -4.5975e-10 | 15.5699
-9.0949e-12 | 15.9809 | -4.3769e-10 | 15.9749
-7.0486e-12 | 16.3859 | -4.6066e-10 | 16.3819
-6.3665e-12 | 16.791 | -4.7203e-10 | 16.787
-7.9581e-12 | 17.196 | -4.7294e-10 | 17.191
-7.9581e-12 | 17.601 | -5.0431e-10 | 17.595
-6.5938e-12 | 18.007 | -4.3929e-10 | 17.999
-7.5033e-12 | 18.4121 | -5.075e-10 | 18.4031
-6.8212e-12 | 18.8161 | -4.1928e-10 | 18.8071
-6.8212e-12 | 19.2211 | -4.0018e-10 | 19.2141
-7.5033e-12 | 19.6271 | -4.1609e-10 | 19.6191
-7.7307e-12 | 20.0312 | -4.6839e-10 | 20.0241
-1.1596e-11 | 20.4352 | -5.0068e-10 | 20.4292
-8.8676e-12 | 20.8402 | -4.5156e-10 | 20.8332
-5.6843e-12 | 21.2452 | -5.2819e-10 | 21.2382
-1.1369e-11 | 21.6502 | -3.9358e-10 | 21.6432
-1.1596e-11 | 22.0563 | -4.2087e-10 | 22.0493
-9.0949e-12 | 22.4623 | -5.1728e-10 | 22.4553
-6.5938e-12 | 22.8673 | -4.0973e-10 | 22.8603
-6.8212e-12 | 23.2713 | -4.2951e-10 | 23.2663
-7.0486e-12 | 23.6754 | -4.411e-10 | 23.6714
-5.6843e-12 | 24.0804 | -4.8135e-10 | 24.0764
-7.0486e-12 | 24.4854 | -4.4633e-10 | 24.4834
-6.1391e-12 | 24.8904 | -4.1609e-10 | 24.8884
-7.7307e-12 | 25.2954 | -4.195e-10 | 25.2924
-5.6843e-12 | 25.7005 | -5.1159e-10 | 25.6975
-7.5033e-12 | 26.1055 | -4.4247e-10 | 26.1035
-5.457e-12 | 26.5085 | -4.7885e-10 | 26.5075
-6.8212e-12 | 26.9125 | -4.4884e-10 | 26.9115
-1.1823e-11 | 27.3176 | -4.9022e-10 | 27.3186
-1.2733e-11 | 27.7216 | -4.286e-10 | 27.7226
-7.0486e-12 | 28.1266 | -4.7498e-10 | 28.1266
-8.1855e-12 | 28.5306 | -4.8863e-10 | 28.5316
-5.6843e-12 | 28.9357 | -4.8885e-10 | 28.9377
-1.5916e-11 | 29.3417 | -4.5634e-10 | 29.3417
-5.457e-12 | 29.7467 | -4.1541e-10 | 29.7457
-8.8676e-12 | 30.1507 | -5.2614e-10 | 30.1497
-8.8676e-12 | 30.5557 | -4.3406e-10 | 30.5547
-7.5033e-12 | 30.9598 | -4.0768e-10 | 30.9608
-7.9581e-12 | 31.3648 | -4.5452e-10 | 31.3648
-8.1855e-12 | 31.7708 | -4.9522e-10 | 31.7698
-7.5033e-12 | 32.1768 | -4.6089e-10 | 32.1748
-6.5938e-12 | 32.5809 | -4.8112e-10 | 32.5779
-6.5938e-12 | 32.9849 | -4.1723e-10 | 32.9829
-1.2506e-11 | 33.3899 | -4.2633e-10 | 33.3879
-7.5033e-12 | 33.7949 | -4.1314e-10 | 33.7939
-4.5475e-12 | 34.201 | -4.7339e-10 | 34.2
-7.5033e-12 | 34.605 | -4.3133e-10 | 34.604
-7.9581e-12 | 35.01 | -4.5111e-10 | 35.009
-9.5497e-12 | 35.415 | -4.6839e-10 | 35.415
-7.7307e-12 | 35.8201 | -4.1859e-10 | 35.8201
-6.1391e-12 | 36.2251 | -4.5429e-10 | 36.2261
-1.0232e-11 | 36.6291 | -5.1227e-10 | 36.6311
-8.6402e-12 | 37.0341 | -4.6157e-10 | 37.0361
-7.0486e-12 | 37.4401 | -4.1086e-10 | 37.4411
-1.0004e-11 | 37.8442 | -4.3451e-10 | 37.8472
-7.0486e-12 | 38.2502 | -4.9135e-10 | 38.2522
-1.0914e-11 | 38.6562 | -5.2728e-10 | 38.6572
-7.7307e-12 | 39.0612 | -4.6202e-10 | 39.0612
-1.0687e-11 | 39.4663 | -4.6748e-10 | 39.4653
-1.3415e-11 | 39.8713 | -4.5202e-10 | 39.8703
-8.8676e-12 | 40.2763 | -4.095e-10 | 40.2753
-6.3665e-12 | 40.6813 | -4.9295e-10 | 40.6793
-5.2296e-12 | 41.0853 | -4.4088e-10 | 41.0834
-7.7307e-12 | 41.4914 | -4.2041e-10 | 41.4874
-7.9581e-12 | 41.8954 | -4.2928e-10 | 41.8924
-6.5938e-12 | 42.3004 | -5.0522e-10 | 42.2964
-5.6843e-12 | 42.7054 | -4.3747e-10 | 42.7004
-1.0459e-11 | 43.1105 | -4.502e-10 | 43.1055
-9.3223e-12 | 43.5165 | -4.2724e-10 | 43.5115
-8.8676e-12 | 43.9205 | -5.1978e-10 | 43.9155
-6.1391e-12 | 44.3255 | -4.1314e-10 | 44.3215
-8.6402e-12 | 44.7316 | -4.1291e-10 | 44.7266
-7.7307e-12 | 45.1366 | -3.9677e-10 | 45.1306
-7.5033e-12 | 45.5426 | -4.4338e-10 | 45.5366
-5.0022e-12 | 45.9466 | -4.4088e-10 | 45.9406
-1.0004e-11 | 46.3516 | -4.6521e-10 | 46.3466
-8.6402e-12 | 46.7567 | -4.1337e-10 | 46.7507
-7.0486e-12 | 47.1617 | -5.0841e-10 | 47.1557
-1.1823e-11 | 47.5647 | -3.7812e-10 | 47.5617
-6.1391e-12 | 47.9707 | -4.1132e-10 | 47.9657
-8.8676e-12 | 48.3758 | -4.7203e-10 | 48.3708
-1.4097e-11 | 48.7818 | -4.2996e-10 | 48.7758
-9.5497e-12 | 49.1868 | -4.1314e-10 | 49.1818
-8.1855e-12 | 49.5908 | -3.9654e-10 | 49.5858
-1.0687e-11 | 49.9949 | -4.8703e-10 | 49.9919
-1.296e-11 | 50.3999 | -4.4406e-10 | 50.3969
-8.1855e-12 | 50.8059 | -4.502e-10 | 50.8009
-7.9581e-12 | 51.2109 | -4.5793e-10 | 51.2049
-9.3223e-12 | 51.6159 | -4.6816e-10 | 51.609
-9.3223e-12 | 52.02 | -4.9567e-10 | 52.014
-9.5497e-12 | 52.426 | -4.4611e-10 | 52.418
-6.5938e-12 | 52.832 | -3.6857e-10 | 52.823
-1.1596e-11 | 53.2381 | -4.0927e-10 | 53.228
-9.5497e-12 | 53.6431 | -4.7066e-10 | 53.6331
-6.3665e-12 | 54.0471 | -4.411e-10 | 54.0371
-9.0949e-12 | 54.4521 | -4.552e-10 | 54.4411
-7.7307e-12 | 54.8571 | -4.236e-10 | 54.8461
-7.7307e-12 | 55.2622 | -5.2387e-10 | 55.2512
-9.3223e-12 | 55.6682 | -4.1041e-10 | 55.6552
-1.0232e-11 | 56.0742 | -4.2337e-10 | 56.0602
-1.0004e-11 | 56.4792 | -4.452e-10 | 56.4642
-9.0949e-12 | 56.8843 | -4.4338e-10 | 56.8683
-1.2733e-11 | 57.2893 | -4.1291e-10 | 57.2733
-1.0232e-11 | 57.6943 | -4.618e-10 | 57.6773
-7.0486e-12 | 58.0993 | -4.8863e-10 | 58.0813
-1.1823e-11 | 58.5053 | -4.2201e-10 | 58.4843
-7.0486e-12 | 58.9094 | -4.9135e-10 | 58.8894
-6.3665e-12 | 59.3144 | -4.9909e-10 | 59.2934
-1.0459e-11 | 59.7184 | -4.2178e-10 | 59.6984
-8.6402e-12 | 60.1244 | -4.5156e-10 | 60.1034
-6.8212e-12 | 60.5285 | -4.1223e-10 | 60.5085
-8.6402e-12 | 60.9335 | -4.3747e-10 | 60.9125
-7.0486e-12 | 61.3385 | -4.9181e-10 | 61.3195
-7.7307e-12 | 61.7425 | -4.2337e-10 | 61.7245
-6.5938e-12 | 62.1476 | -4.7498e-10 | 62.1285
-7.9581e-12 | 62.5526 | -5.0068e-10 | 62.5336
-8.1855e-12 | 62.9566 | -4.7498e-10 | 62.9376
-9.3223e-12 | 63.3616 | -4.5748e-10 | 63.3426
-7.7307e-12 | 63.7666 | -4.4224e-10 | 63.7467
-8.8676e-12 | 64.1717 | -4.4338e-10 | 64.1507
-7.5033e-12 | 64.5767 | -5.3274e-10 | 64.5567
-1.0004e-11 | 64.9817 | -4.3474e-10 | 64.9607
-7.9581e-12 | 65.3867 | -4.8476e-10 | 65.3657
-6.1391e-12 | 65.7918 | -4.3383e-10 | 65.7698
-9.3223e-12 | 66.1958 | -4.718e-10 | 66.1738
-1.1823e-11 | 66.5998 | -4.4793e-10 | 66.5788
-7.9581e-12 | 67.0058 | -4.7089e-10 | 66.9838
-9.0949e-12 | 67.4109 | -4.8021e-10 | 67.3879
-8.1855e-12 | 67.8159 | -3.9518e-10 | 67.7939
-8.6402e-12 | 68.2209 | -4.809e-10 | 68.1979
-7.7307e-12 | 68.6259 | -4.7226e-10 | 68.6019
-6.8212e-12 | 69.0299 | -4.6521e-10 | 69.0079
-8.8676e-12 | 69.433 | -4.793e-10 | 69.413
-9.0949e-12 | 69.837 | -3.8835e-10 | 69.817
-9.0949e-12 | 70.242 | -4.3337e-10 | 70.222
-9.3223e-12 | 70.647 | -4.8385e-10 | 70.626
-9.0949e-12 | 71.0531 | -5.2864e-10 | 71.0301
-1.2051e-11 | 71.4571 | -4.4679e-10 | 71.4341
-5.6843e-12 | 71.8631 | -4.0109e-10 | 71.8391
-6.3665e-12 | 72.2681 | -4.4679e-10 | 72.2441
-9.3223e-12 | 72.6722 | -4.4997e-10 | 72.6492
-1.5916e-11 | 73.0762 | -4.9408e-10 | 73.0542
-7.0486e-12 | 73.4822 | -4.6452e-10 | 73.4582
-8.6402e-12 | 73.8862 | -4.4247e-10 | 73.8642
-1.6826e-11 | 74.2922 | -3.9404e-10 | 74.2692
-8.8676e-12 | 74.6983 | -4.5839e-10 | 74.6743
-8.1855e-12 | 75.1033 | -4.3792e-10 | 75.0783
-6.8212e-12 | 75.5073 | -4.2951e-10 | 75.4823
-8.6402e-12 | 75.9113 | -4.2314e-10 | 75.8863
-7.0486e-12 | 76.3164 | -4.5679e-10 | 76.2924
-1.0004e-11 | 76.7234 | -4.668e-10 | 76.6964
-9.0949e-12 | 77.1284 | -3.8517e-10 | 77.1024
-7.9581e-12 | 77.5344 | -4.8703e-10 | 77.5074
-7.7307e-12 | 77.9395 | -4.1564e-10 | 77.9125
-9.5497e-12 | 78.3445 | -4.2814e-10 | 78.3175
-7.7307e-12 | 78.7495 | -4.4383e-10 | 78.7225
-7.9581e-12 | 79.1535 | -4.934e-10 | 79.1285
-1.0687e-11 | 79.5596 | -4.8453e-10 | 79.5336
-8.8676e-12 | 79.9636 | -4.411e-10 | 79.9396
-9.5497e-12 | 80.3676 | -4.5338e-10 | 80.3446
-5.457e-12 | 80.7716 | -4.2178e-10 | 80.7506
-1.1369e-11 | 81.1776 | -4.5952e-10 | 81.1546
-7.7307e-12 | 81.5817 | -4.386e-10 | 81.5587
-1.296e-11 | 81.9867 | -4.1064e-10 | 81.9627
-6.8212e-12 | 82.3917 | -4.5134e-10 | 82.3687
-1.0459e-11 | 82.7977 | -4.8453e-10 | 82.7737
|  | -4.6407e-10 | 83.1788
|  | -4.3883e-10 | 83.5838

## 690uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.0173e-11 | 3.0022e-12 | -6.48e-10 | 5.0319e-11
GEMs On |  | GEMs On
-7.5033e-12 | 0.31102 | -6.7348e-10 | 0.31202
-5.2296e-12 | 0.99606 | -6.9008e-10 | 0.99506
-9.0949e-12 | 1.4011 | -7.1327e-10 | 1.4011
-5.9117e-12 | 1.8071 | -6.4597e-10 | 1.8061
-1.1141e-11 | 2.2121 | -6.3596e-10 | 2.2111
-9.5497e-12 | 2.6191 | -5.9617e-10 | 2.6162
-9.5497e-12 | 3.0242 | -6.3892e-10 | 3.0212
-6.5938e-12 | 3.4292 | -5.9981e-10 | 3.4272
-6.5938e-12 | 3.8402 | -6.3778e-10 | 3.8312
-6.5938e-12 | 4.2462 | -6.869e-10 | 4.2362
-8.1855e-12 | 4.6573 | -6.2732e-10 | 4.6403
-1.3415e-11 | 5.0613 | -6.8212e-10 | 5.0443
-7.0486e-12 | 5.4663 | -7.4328e-10 | 5.4503
-6.5938e-12 | 5.8713 | -7.0281e-10 | 5.8553
-1.728e-11 | 6.2764 | -5.9458e-10 | 6.2594
-7.276e-12 | 6.6824 | -5.9504e-10 | 6.6644
-8.1855e-12 | 7.0884 | -6.2596e-10 | 7.0684
-9.3223e-12 | 7.4924 | -7.1259e-10 | 7.4724
-8.1855e-12 | 7.8985 | -6.4847e-10 | 7.8765
-6.8212e-12 | 8.3035 | -5.8094e-10 | 8.2815
-2.1146e-11 | 8.7085 | -6.1254e-10 | 8.6865
-9.0949e-12 | 9.1135 | -6.6711e-10 | 9.0925
-9.3223e-12 | 9.5185 | -6.0004e-10 | 9.4975
-1.4325e-11 | 9.9236 | -6.2869e-10 | 9.9026
-8.6402e-12 | 10.3296 | -5.914e-10 | 10.3076
-9.7771e-12 | 10.7376 | -6.2164e-10 | 10.7136
-9.0949e-12 | 11.1426 | -6.321e-10 | 11.1186
-6.8212e-12 | 11.5477 | -6.5143e-10 | 11.5237
-1.4325e-11 | 11.9527 | -6.5916e-10 | 11.9277
-8.1855e-12 | 12.3577 | -5.823e-10 | 12.3327
-1.2051e-11 | 12.7627 | -6.3301e-10 | 12.7387
-8.6402e-12 | 13.1668 | -6.1391e-10 | 13.1418
-7.0486e-12 | 13.5718 | -6.4119e-10 | 13.5468
-7.276e-12 | 13.9778 | -7.1441e-10 | 13.9518
-8.1855e-12 | 14.3828 | -6.6711e-10 | 14.3578
-7.0486e-12 | 14.7868 | -6.537e-10 | 14.7618
-7.0486e-12 | 15.1919 | -6.2596e-10 | 15.1659
-7.276e-12 | 15.5969 | -6.3937e-10 | 15.5699
-1.0004e-11 | 16.0019 | -5.5275e-10 | 15.9759
-9.3223e-12 | 16.4069 | -6.4097e-10 | 16.3809
-1.0004e-11 | 16.812 | -5.9481e-10 | 16.786
-8.6402e-12 | 17.217 | -6.4279e-10 | 17.191
-9.7771e-12 | 17.623 | -7.0509e-10 | 17.594
-6.8212e-12 | 18.029 | -6.3505e-10 | 18
-7.276e-12 | 18.434 | -6.446e-10 | 18.4041
-1.0459e-11 | 18.8401 | -6.2914e-10 | 18.8081
-8.1855e-12 | 19.2451 | -7.0168e-10 | 19.2131
-1.2051e-11 | 19.6531 | -6.9122e-10 | 19.6171
-1.6144e-11 | 20.0591 | -7.6579e-10 | 20.0221
-9.3223e-12 | 20.4652 | -7.1054e-10 | 20.4262
-8.1855e-12 | 20.8702 | -6.4597e-10 | 20.8292
-1.1369e-11 | 21.2752 | -7.1168e-10 | 21.2332
-9.0949e-12 | 21.6812 | -6.0481e-10 | 21.6392
-1.2506e-11 | 22.0873 | -6.6029e-10 | 22.0433
-9.0949e-12 | 22.4923 | -6.8258e-10 | 22.4473
-5.2296e-12 | 22.8963 | -5.9845e-10 | 22.8513
-1.0459e-11 | 23.3003 | -6.8781e-10 | 23.2583
-8.4128e-12 | 23.7054 | -6.6939e-10 | 23.6623
-1.296e-11 | 24.1104 | -6.3847e-10 | 24.0664
-7.276e-12 | 24.5154 | -6.2914e-10 | 24.4704
-9.3223e-12 | 24.9204 | -6.4483e-10 | 24.8744
-7.0486e-12 | 25.3255 | -5.5729e-10 | 25.2784
-8.6402e-12 | 25.7325 | -6.8576e-10 | 25.6825
-1.0687e-11 | 26.1385 | -6.2482e-10 | 26.0865
-7.276e-12 | 26.5435 | -6.8735e-10 | 26.4915
-9.3223e-12 | 26.9495 | -6.5415e-10 | 26.8965
-9.0949e-12 | 27.3536 | -6.1914e-10 | 27.3016
-8.1855e-12 | 27.7586 | -6.9963e-10 | 27.7056
-6.8212e-12 | 28.1626 | -6.0254e-10 | 28.1096
-8.6402e-12 | 28.5666 | -5.3979e-10 | 28.5146
-7.276e-12 | 28.9707 | -7.542e-10 | 28.9176
-9.3223e-12 | 29.3757 | -6.9417e-10 | 29.3227
-8.4128e-12 | 29.7807 | -6.296e-10 | 29.7267
-7.9581e-12 | 30.1847 | -7.2555e-10 | 30.1337
-1.0459e-11 | 30.5888 | -5.8731e-10 | 30.5377
-7.9581e-12 | 30.9938 | -5.6912e-10 | 30.9428
-1.0459e-11 | 31.3988 | -5.5297e-10 | 31.3478
-9.0949e-12 | 31.8048 | -7.1259e-10 | 31.7528
-1.1141e-11 | 32.2098 | -6.1436e-10 | 32.1578
-8.1855e-12 | 32.6159 | -6.6711e-10 | 32.5629
-1.296e-11 | 33.0209 | -6.7416e-10 | 32.9679
-1.0004e-11 | 33.4279 | -5.6593e-10 | 33.3729
-1.1823e-11 | 33.8319 | -6.296e-10 | 33.7779
-1.1141e-11 | 34.237 | -6.537e-10 | 34.183
-7.276e-12 | 34.641 | -7.0395e-10 | 34.587
-7.276e-12 | 35.047 | -5.3865e-10 | 34.991
-9.0949e-12 | 35.452 | -6.8599e-10 | 35.396
-8.1855e-12 | 35.858 | -7.4078e-10 | 35.8
-8.6402e-12 | 36.2631 | -6.4006e-10 | 36.2041
-8.1855e-12 | 36.6691 | -6.1641e-10 | 36.6101
-9.5497e-12 | 37.0751 | -6.5393e-10 | 37.0151
-7.5033e-12 | 37.4791 | -6.3596e-10 | 37.4201
-1.296e-11 | 37.8852 | -5.889e-10 | 37.8262
-1.0687e-11 | 38.2922 | -6.3756e-10 | 38.2302
-8.6402e-12 | 38.6972 | -6.0891e-10 | 38.6362
-1.4552e-11 | 39.1032 | -7.451e-10 | 39.0402
-9.0949e-12 | 39.5093 | -7.2964e-10 | 39.4453
-9.7771e-12 | 39.9143 | -7.1395e-10 | 39.8503
-1.1823e-11 | 40.3173 | -6.6552e-10 | 40.2553
-7.0486e-12 | 40.7233 | -6.6916e-10 | 40.6593
-9.3223e-12 | 41.1283 | -5.7798e-10 | 41.0654
-8.1855e-12 | 41.5354 | -6.2141e-10 | 41.4694
-1.9327e-11 | 41.9404 | -6.9258e-10 | 41.8754
-8.4128e-12 | 42.3464 | -6.0163e-10 | 42.2794
-8.4128e-12 | 42.7514 | -5.889e-10 | 42.6834
-7.9581e-12 | 43.1565 | -5.6934e-10 | 43.0895
-1.0459e-11 | 43.5625 | -6.3164e-10 | 43.4935
-1.3415e-11 | 43.9685 | -6.2801e-10 | 43.8975
-1.1141e-11 | 44.3745 | -8.1195e-10 | 44.3015
-8.4128e-12 | 44.7806 | -6.4324e-10 | 44.7076
-7.0486e-12 | 45.1856 | -7.2555e-10 | 45.1206
-1.2051e-11 | 45.5896 | -6.7075e-10 | 45.5266
-8.4128e-12 | 45.9946 | -7.0418e-10 | 45.9326
-1.0687e-11 | 46.3997 | -5.6025e-10 | 46.3366
-1.0914e-11 | 46.8047 | -6.5097e-10 | 46.7407
-1.3415e-11 | 47.2107 | -7.6056e-10 | 47.1447
-1.1369e-11 | 47.6157 | -5.4411e-10 | 47.5497
-8.6402e-12 | 48.0207 | -7.16e-10 | 47.9537
-8.1855e-12 | 48.4248 | -6.5165e-10 | 48.3598
-2.1146e-11 | 48.8288 | -7.3828e-10 | 48.7648
-8.6402e-12 | 49.2338 | -6.5415e-10 | 49.2058
-1.296e-11 | 49.6398 | -6.8462e-10 | 49.6348
-8.6402e-12 | 50.0449 | -6.5097e-10 | 50.0399
-1.2278e-11 | 50.4509 | -5.9117e-10 | 50.4449
-8.6402e-12 | 50.8569 | -6.7598e-10 | 50.8499
-9.3223e-12 | 51.2629 | -5.8935e-10 | 51.2549
-1.0459e-11 | 51.669 | -7.0122e-10 | 51.66
-8.6402e-12 | 52.075 | -7.3965e-10 | 52.064
-1.0004e-11 | 52.48 | -6.3892e-10 | 52.468
-9.7771e-12 | 52.885 | -5.889e-10 | 52.872
-7.0486e-12 | 53.291 | -6.5529e-10 | 53.2771
-1.1141e-11 | 53.6961 | -6.3051e-10 | 53.6831
-8.6402e-12 | 54.1011 | -7.1395e-10 | 54.0891
-1.6144e-11 | 54.5061 | -7.1418e-10 | 54.4941
-8.1855e-12 | 54.9111 | -6.271e-10 | 54.9001
-1.1823e-11 | 55.3162 | -6.1709e-10 | 55.3052
-9.7771e-12 | 55.7232 | -6.6279e-10 | 55.7092
-9.5497e-12 | 56.1292 | -6.7303e-10 | 56.1152
-9.0949e-12 | 56.5352 | -6.0027e-10 | 56.5202
-7.5033e-12 | 56.9403 | -6.1323e-10 | 56.9253
-7.5033e-12 | 57.3463 | -6.2141e-10 | 57.3293
-1.3188e-11 | 57.7503 | -6.3164e-10 | 57.7333
-1.3188e-11 | 58.1553 | -5.5525e-10 | 58.1383
-7.0486e-12 | 58.5613 | -6.0481e-10 | 58.5433
-1.0914e-11 | 58.9664 | -6.4483e-10 | 58.9544
-9.0949e-12 | 59.3724 | -5.6139e-10 | 59.3594
-7.276e-12 | 59.7774 | -6.3073e-10 | 59.7654
-7.5033e-12 | 60.1824 | -6.6848e-10 | 60.1714
-1.296e-11 | 60.5865 | -6.4369e-10 | 60.5765
-2.2055e-11 | 60.9925 | -6.6984e-10 | 60.9805
-1.2506e-11 | 61.3975 | -7.0622e-10 | 61.3845
-9.7771e-12 | 61.8025 | -6.1209e-10 | 61.7895
-1.4779e-11 | 62.2086 | -6.0822e-10 | 62.1936
-1.1369e-11 | 62.6146 | -6.321e-10 | 62.5986
-8.1855e-12 | 63.0206 | -5.7548e-10 | 63.0046
-1.6144e-11 | 63.4256 | -7.0759e-10 | 63.4086
-9.7771e-12 | 63.8306 | -6.1596e-10 | 63.8147
-1.1823e-11 | 64.2357 | -6.5756e-10 | 64.2187
-7.9581e-12 | 64.6407 | -6.8712e-10 | 64.6247
-1.5689e-11 | 65.0457 | -6.7962e-10 | 65.0297
-1.3188e-11 | 65.4507 | -7.0531e-10 | 65.4337
-9.5497e-12 | 65.8568 | -6.3324e-10 | 65.8388
-7.5033e-12 | 66.2618 | -6.462e-10 | 66.2458
-1.2506e-11 | 66.6678 | -7.3464e-10 | 66.6498
-9.5497e-12 | 67.0728 | -6.6802e-10 | 67.0538
-1.0914e-11 | 67.4779 | -6.1027e-10 | 67.4599
-7.9581e-12 | 67.8839 | -6.18e-10 | 67.8669
-1.6371e-11 | 68.2889 | -6.5097e-10 | 68.2719
-1.1823e-11 | 68.6939 | -6.0095e-10 | 68.6759
-1.6371e-11 | 69.1 | -6.3164e-10 | 69.082
-8.6402e-12 | 69.505 | -6.0004e-10 | 69.485
-2.2283e-11 | 69.908 | -6.4801e-10 | 69.891
-1.5461e-11 | 70.314 | -6.5438e-10 | 70.298
-1.2278e-11 | 70.719 | -6.7962e-10 | 70.701
-9.0949e-12 | 71.1261 | -6.6393e-10 | 71.1051
-1.0459e-11 | 71.5311 | -6.7462e-10 | 71.5141
-9.5497e-12 | 71.9371 | -7.1009e-10 | 71.9191
-8.4128e-12 | 72.3411 | -5.9003e-10 | 72.3241
-9.3223e-12 | 72.7462 | -6.6984e-10 | 72.7292
-1.0914e-11 | 73.1512 | -5.5093e-10 | 73.1342
-1.0459e-11 | 73.5562 | -5.5661e-10 | 73.5382
-1.0687e-11 | 73.9622 | -6.4892e-10 | 73.9432
-6.8212e-12 | 74.3672 | -6.421e-10 | 74.3473
-1.296e-11 | 74.7713 | -6.1664e-10 | 74.7523
-1.0459e-11 | 75.1763 | -7.5056e-10 | 75.1573
-1.1369e-11 | 75.5823 | -5.9367e-10 | 75.5623
-1.2278e-11 | 75.9874 | -6.1436e-10 | 75.9673
-2.16e-11 | 76.3924 | -6.6211e-10 | 76.3714
-1.0687e-11 | 76.7984 | -7.0122e-10 | 76.7784
-9.3223e-12 | 77.2044 | -6.7712e-10 | 77.1824
-9.0949e-12 | 77.6094 | -6.8553e-10 | 77.5884
-1.0459e-11 | 78.0135 | -5.9003e-10 | 77.9945
-1.1141e-11 | 78.4185 | -6.4847e-10 | 78.3995
-1.1369e-11 | 78.8235 | -6.7985e-10 | 78.8035
-8.6402e-12 | 79.2285 | -6.4892e-10 | 79.2075
-1.0459e-11 | 79.6355 | -5.773e-10 | 79.6115
-9.7771e-12 | 80.0406 | -6.4347e-10 | 80.0166
-1.2051e-11 | 80.4446 | -7.0577e-10 | 80.4206
-9.7771e-12 | 80.8506 | -6.6439e-10 | 80.8276
-9.3223e-12 | 81.2546 | -6.4915e-10 | 81.2317
-1.1369e-11 | 81.6607 | -6.8189e-10 | 81.6357
-1.0004e-11 | 82.0657 | -6.5302e-10 | 82.0417
-9.7771e-12 | 82.4707 | -7.3442e-10 | 82.4467
-1.1823e-11 | 82.8767 | -5.9322e-10 | 82.8517
-1.0004e-11 | 83.2828 | -6.2437e-10 | 83.2578
-1.3188e-11 | 83.6888
-1.0459e-11 | 84.0948
-9.0949e-12 | 84.4998
-1.2051e-11 | 84.9049
-1.4325e-11 | 85.3109

## 700uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.7362e-11 | 5.6878e-12 | -9.1142e-10 | 6.4454e-11
GEMs On |  | GEMs On
-8.6402e-12 | 1.5441 | -9.7361e-10 | 0.31402
-1.2733e-11 | 2.3161 | -9.1723e-10 | 1.0001
-1.0459e-11 | 2.7222 | -8.0695e-10 | 1.4051
-1.0459e-11 | 3.1292 | -1.0123e-09 | 1.8101
-8.1855e-12 | 3.5342 | -8.6857e-10 | 2.2171
-1.0004e-11 | 3.9392 | -7.7853e-10 | 2.6212
-1.2051e-11 | 4.3452 | -7.9331e-10 | 3.0252
-1.3642e-11 | 4.7493 | -8.7539e-10 | 3.4302
-9.7771e-12 | 5.1543 | -9.0563e-10 | 3.8342
-3.5016e-11 | 5.5593 | -8.647e-10 | 4.2382
-1.4552e-11 | 5.9643 | -8.7971e-10 | 4.6433
-1.1823e-11 | 6.3704 | -8.9449e-10 | 5.0483
-1.1369e-11 | 6.7754 | -8.622e-10 | 5.4553
-1.5007e-11 | 7.1814 | -9.4769e-10 | 5.8593
-1.1369e-11 | 7.5874 | -1.0339e-09 | 6.2644
-1.2051e-11 | 7.9925 | -8.7743e-10 | 6.6684
-1.1369e-11 | 8.3975 | -8.4788e-10 | 7.0714
-9.7771e-12 | 8.8035 | -8.647e-10 | 7.4764
-1.2278e-11 | 9.2095 | -8.3651e-10 | 7.8805
-8.8676e-12 | 9.6156 | -8.8539e-10 | 8.2845
-2.3419e-11 | 10.0206 | -8.9062e-10 | 8.6895
-1.1141e-11 | 10.4256 | -8.074e-10 | 9.0935
-1.0914e-11 | 10.8316 | -9.4724e-10 | 9.4975
-1.0004e-11 | 11.2366 | -9.918e-10 | 9.9026
-8.6402e-12 | 11.6417 | -9.2837e-10 | 10.3066
-1.6371e-11 | 12.0467 | -8.1059e-10 | 10.7116
-1.2506e-11 | 12.4527 | -9.1336e-10 | 11.1146
-8.6402e-12 | 12.8587 | -9.4406e-10 | 11.5197
-9.5497e-12 | 13.2648 | -9.2996e-10 | 11.9247
-1.4779e-11 | 13.6688 | -9.0199e-10 | 12.3297
-1.4779e-11 | 14.0728 | -9.2291e-10 | 12.7337
-1.1823e-11 | 14.4788 | -8.7812e-10 | 13.1387
-9.5497e-12 | 14.8849 | -9.4337e-10 | 13.5438
-1.5916e-11 | 15.2899 | -7.6375e-10 | 13.9488
-1.0687e-11 | 15.6969 | -1.0089e-09 | 14.3528
-1.2278e-11 | 16.1009 | -9.4019e-10 | 14.7578
-7.5033e-12 | 16.5069 | -9.2405e-10 | 15.1619
-1.5007e-11 | 16.911 | -9.1177e-10 | 15.5659
-1.1823e-11 | 17.316 | -7.683e-10 | 15.9699
-1.1141e-11 | 17.721 | -8.3583e-10 | 16.3739
-1.2051e-11 | 18.126 | -1.0448e-09 | 16.779
-1.4779e-11 | 18.5301 | -8.929e-10 | 17.184
-1.3415e-11 | 18.9351 | -9.5497e-10 | 17.589
-2.251e-11 | 19.3401 | -9.4178e-10 | 17.994
-1.2733e-11 | 19.7451 | -8.0536e-10 | 18.3981
-1.3415e-11 | 20.1502 | -9.9203e-10 | 18.8041
-1.3642e-11 | 20.5562 | -8.9017e-10 | 19.2081
-1.0687e-11 | 20.9612 | -8.7334e-10 | 19.6141
-1.1369e-11 | 21.3662 | -9.7384e-10 | 20.0191
-1.2733e-11 | 21.7722 | -9.7975e-10 | 20.4242
-1.8417e-11 | 22.1773 | -9.7111e-10 | 20.8302
-1.3415e-11 | 22.5823 | -9.8885e-10 | 21.2342
-1.2051e-11 | 22.9863 | -9.0836e-10 | 21.6392
-2.7512e-11 | 23.3913 | -8.1059e-10 | 22.0443
-1.5916e-11 | 23.7974 | -8.9585e-10 | 22.4483
-1.819e-11 | 24.2024 | -8.7539e-10 | 22.8523
-1.2278e-11 | 24.6084 | -8.6129e-10 | 23.2583
-1.1141e-11 | 25.0144 | -9.5747e-10 | 23.6634
-2.5921e-11 | 25.4204 | -8.6607e-10 | 24.0704
-1.728e-11 | 25.8255 | -7.1714e-10 | 24.4764
-1.9099e-11 | 26.2325 | -7.8808e-10 | 24.8814
-1.0914e-11 | 26.6385 | -8.4424e-10 | 25.2855
-1.3188e-11 | 27.0446 | -7.9194e-10 | 25.6905
-1.2733e-11 | 27.4496 | -9.1904e-10 | 26.0965
-2.4102e-11 | 27.8556 | -8.8312e-10 | 26.5005
-1.0914e-11 | 28.2606 | -9.1745e-10 | 26.9045
-1.4779e-11 | 28.6666 | -1.0157e-09 | 27.3086
-1.4325e-11 | 29.0727 | -9.2086e-10 | 27.7126
-1.3642e-11 | 29.4777 | -9.6247e-10 | 28.1176
-1.6144e-11 | 29.8827 | -9.4201e-10 | 28.5226
-1.387e-11 | 30.2887 | -9.9317e-10 | 28.9266
-1.5007e-11 | 30.6938 | -1.002e-09 | 29.3317
-1.3188e-11 | 31.0998 | -9.8953e-10 | 29.7347
-1.4779e-11 | 31.5048 | -9.1495e-10 | 30.1397
-2.4556e-11 | 31.9098 | -8.9585e-10 | 30.5448
-1.5007e-11 | 32.3169 | -9.0995e-10 | 30.9488
-1.4552e-11 | 32.7229 | -8.5311e-10 | 31.3548
-1.0914e-11 | 33.1289 | -8.3196e-10 | 31.7588
-1.4779e-11 | 33.5339 | -9.1086e-10 | 32.1638
-1.2733e-11 | 33.9409 | -8.2468e-10 | 32.5699
-1.4552e-11 | 34.346 | -8.738e-10 | 32.9749
-1.5234e-11 | 34.751 | -1.0355e-09 | 33.3779
-2.16e-11 | 35.157 | -9.6452e-10 | 33.7819
-1.2506e-11 | 35.561 | -9.0404e-10 | 34.186
-2.3419e-11 | 35.9671 | -8.1127e-10 | 34.591
-3.638e-11 | 36.3711 | -9.2359e-10 | 34.995
-1.3188e-11 | 36.7761 | -8.9426e-10 | 35.399
-1.7508e-11 | 37.1821 | -7.9922e-10 | 35.8021
-1.9554e-11 | 37.5881 | -9.6816e-10 | 36.2061
-1.819e-11 | 37.9932 | -9.8748e-10 | 36.6111
-1.5689e-11 | 38.3992 | -9.3792e-10 | 37.0151
-1.6826e-11 | 38.8042 | -8.4924e-10 | 37.4201
-2.3647e-11 | 39.2102 | -8.9972e-10 | 37.8252
-1.3188e-11 | 39.6173 | -8.8789e-10 | 38.2292
-2.1373e-11 | 40.0233 | -8.888e-10 | 38.6352
-1.5234e-11 | 40.4283 | -9.2155e-10 | 39.0402
-1.8645e-11 | 40.8323 | -9.8657e-10 | 39.4453
-1.387e-11 | 41.2384 | -9.9249e-10 | 39.8503
-2.2737e-11 | 41.6444 | -9.2245e-10 | 40.2543
-2.0236e-11 | 42.0504 | -8.9653e-10 | 40.6603
-1.3642e-11 | 42.4554 | -8.2036e-10 | 41.0643
-1.4552e-11 | 42.8605 | -9.1063e-10 | 41.4694
-1.2733e-11 | 43.2665 | -9.1632e-10 | 41.8744
-1.3188e-11 | 43.6735 | -1.0004e-09 | 42.2784
-1.7508e-11 | 44.0785 | -8.9449e-10 | 42.6834
-1.7508e-11 | 44.4835 | -9.1904e-10 | 43.0875
-2.1146e-11 | 44.8886 | -8.6311e-10 | 43.4925
-1.3188e-11 | 45.2946 | -8.6902e-10 | 43.8975
-1.4325e-11 | 45.7006 | -8.5402e-10 | 44.3025
-1.5689e-11 | 46.1066 | -9.6043e-10 | 44.7066
-2.8422e-11 | 46.5117 | -9.4019e-10 | 45.1116
-1.5007e-11 | 46.9157 | -1.0482e-09 | 45.5166
-1.819e-11 | 47.3207 | -9.2609e-10 | 45.9206
-1.5007e-11 | 47.7267 | -1.0266e-09 | 46.3246
-1.8417e-11 | 48.1307 | -8.9176e-10 | 46.7307
-2.2055e-11 | 48.5358 | -8.506e-10 | 47.1347
-3.5243e-11 | 48.9408 | -9.1768e-10 | 47.5387
-1.3415e-11 | 49.3458 | -9.6338e-10 | 47.9447
-1.7053e-11 | 49.7529 | -8.3264e-10 | 48.3478
-1.6371e-11 | 50.1579 | -8.8835e-10 | 48.7528
-2.2283e-11 | 50.5639 | -9.1654e-10 | 49.1578
-1.5234e-11 | 50.9699 | -9.2746e-10 | 49.5628
-3.2742e-11 | 51.3759 | -9.5588e-10 | 49.9669
-1.6826e-11 | 51.782 | -1.0127e-09 | 50.3709
-2.6148e-11 | 52.187 | -9.2086e-10 | 50.7759
-1.7508e-11 | 52.593 | -9.3405e-10 | 51.1799
-1.387e-11 | 52.999 | -1.0168e-09 | 51.5849
-2.2965e-11 | 53.404 | -8.5174e-10 | 51.989
-1.6826e-11 | 53.8101 | -9.0631e-10 | 52.394
-3.2742e-11 | 54.2171 | -9.0859e-10 | 52.799
-1.4552e-11 | 54.6211 | -8.6561e-10 | 53.205
-1.6144e-11 | 55.0281 | -8.9062e-10 | 53.6101
-1.6371e-11 | 55.4322 | -9.4428e-10 | 54.0161
-2.683e-11 | 55.8382 | -9.4542e-10 | 54.4211
-2.0918e-11 | 56.2442 | -8.6607e-10 | 54.8241
-2.5238e-11 | 56.6492 | -1.0127e-09 | 55.2292
-1.5689e-11 | 57.0553 | -9.3564e-10 | 55.6342
-2.5238e-11 | 57.4603 | -9.7293e-10 | 56.0372
-1.6826e-11 | 57.8653 | -9.6179e-10 | 56.4422
-1.9782e-11 | 58.2693 | -8.5151e-10 | 56.8462
-3.5243e-11 | 58.6744 | -9.6975e-10 | 57.2503
-2.8649e-11 | 59.0804 | -9.1427e-10 | 57.6553
-1.5234e-11 | 59.4854 | -1.03e-09 | 58.0613
-2.1146e-11 | 59.8894 | -8.5606e-10 | 58.4653
-1.7508e-11 | 60.2955 | -8.3764e-10 | 58.8684
-1.3642e-11 | 60.7015 | -9.0631e-10 | 59.2744
-2.0918e-11 | 61.1055 | -8.8744e-10 | 59.6784
-1.5234e-11 | 61.5105 | -9.2427e-10 | 60.0834
-1.728e-11 | 61.9155 | -9.3246e-10 | 60.4895
-1.4779e-11 | 62.3216 | -9.7157e-10 | 60.8945
-1.8645e-11 | 62.7286 | -9.2655e-10 | 61.2995
-1.728e-11 | 63.1336 | -9.2018e-10 | 61.7045
-1.5916e-11 | 63.5396 | -9.2336e-10 | 62.1095
-3.1832e-11 | 63.9447 | -8.9085e-10 | 62.5146
-1.6144e-11 | 64.3497 | -9.9635e-10 | 62.9186
-1.5916e-11 | 64.7547 | -9.9067e-10 | 63.3246
-2.2965e-11 | 65.1587 | -9.2837e-10 | 63.7286
-2.16e-11 | 65.5667 | -8.3241e-10 | 64.1337
-1.7053e-11 | 65.9718 | -9.22e-10 | 64.5377
-1.7053e-11 | 66.3778 | -8.7289e-10 | 64.9437
-1.5916e-11 | 66.7828 | -9.054e-10 | 65.3487
-2.8422e-11 | 67.1888 | -9.7748e-10 | 65.7528
-1.7508e-11 | 67.5939 | -8.7221e-10 | 66.1568
-1.728e-11 | 67.9989 | -9.0677e-10 | 66.5618
-1.8645e-11 | 68.4039 | -9.3291e-10 | 66.9658
-1.5916e-11 | 68.8089 | -8.522e-10 | 67.3718
-2.9104e-11 | 69.214 | -1.0175e-09 | 67.7759
-1.8872e-11 | 69.62 | -1.0086e-09 | 68.1799
-1.8872e-11 | 70.024 | -1.0266e-09 | 68.5849
-1.6144e-11 | 70.429 | -8.5856e-10 | 68.989
-1.9099e-11 | 70.835 | -1.0541e-09 | 69.393
-1.7053e-11 | 71.2401 | -8.5242e-10 | 69.797
-1.9782e-11 | 71.6441 | -8.8835e-10 | 70.202
-2.0918e-11 | 72.0501 | -9.2928e-10 | 70.606
-1.8417e-11 | 72.4561 | -9.2109e-10 | 71.0101
-1.728e-11 | 72.8622 | -9.0677e-10 | 71.4141
-1.8645e-11 | 73.2682 | -1.0305e-09 | 71.8201
-2.8876e-11 | 73.6722 | -9.7975e-10 | 72.2241
-3.115e-11 | 74.0772 | -9.4838e-10 | 72.6282
-1.9554e-11 | 74.4823 | -8.7994e-10 | 73.0332
-2.5466e-11 | 74.8883 | -1.0789e-09 | 73.4372
-1.9554e-11 | 75.2933 | -8.697e-10 | 73.8412
-1.8645e-11 | 75.6993 | -9.2541e-10 | 74.2463
-1.9099e-11 | 76.1034 | -8.5424e-10 | 74.6513
-1.819e-11 | 76.5084 | -9.3291e-10 | 75.0553
-1.7508e-11 | 76.9134 | -8.6311e-10 | 75.4593
-1.7508e-11 | 77.3184 | -9.843e-10 | 75.8643
-1.9554e-11 | 77.7245 | -8.5925e-10 | 76.2674
-2.1373e-11 | 78.1295 | -9.029e-10 | 76.6714
-1.9782e-11 | 78.5355 | -9.0358e-10 | 77.0774
-1.7508e-11 | 78.9405 | -7.9149e-10 | 77.4824
-2.2055e-11 | 79.3455 | -9.0699e-10 | 77.8864
-2.3647e-11 | 79.7506 | -8.7971e-10 | 78.2915
-2.3874e-11 | 80.1556 | -8.1309e-10 | 78.6965
-2.0691e-11 | 80.5606 | -7.8126e-10 | 79.1025
-2.0236e-11 | 80.9656 | -7.8353e-10 | 79.5075
-1.8645e-11 | 81.3717 | -9.2427e-10 | 79.9126
-1.7735e-11 | 81.7777 | -8.7084e-10 | 80.3176
-1.8417e-11 | 82.1827 | -9.0608e-10 | 80.7226
-2.4556e-11 | 82.5877 | -8.9949e-10 | 81.1276
-1.9554e-11 | 82.9928 | -8.7675e-10 | 81.5307
-1.9782e-11 | 83.3988 | -9.4337e-10 | 81.9347
-2.4784e-11 | 83.8048 | -9.3314e-10 | 82.3397
-2.0009e-11 | 84.2108 | -8.9085e-10 | 82.7437
-2.4102e-11 | 84.6158 | -9.7884e-10 | 83.1498
|  | -9.8021e-10 | 83.5548

## 710uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 720uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 730uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 740uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 750uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 760uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 770uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 780uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 790uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 800uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 810uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On