## Data Summary
QC5 Effective gain
Username: | Christos Roskas | QC5 Template | V2 |  | Environment |  |  |  |  | For Measuring Rate |  |  |  | For Measuring Current
Amplifier Settings |  | Date | 2018-04-05 |  | Vmon | Imon | Time | Pressure | Temp | Source Off |  | Source On |  | Source Off |  | Source On
[10 empty] | Counts | Err | Counts | Err | Current | Err | Current | Err
Make & Model: | ORTEC 474 |  |  |  | (V) | (uA) | (HH:MM) | (mbar) | (Deg C) | (N) | (N) | (N) | (N) | (A) | (A) | (A) | (A)
Coarse Gain: | 4x |  |  | Ar/CO2:70/30 Range | 3501 | 700 | 10:07 | 1005 | 20.6 | 381 | 19.5192 | 97314 | 311.9519 | -1.8069e-11 | 4.9387e-12 | -7.4341e-10 | 6.0247e-11
Fine Gain: | 4.5 |  |  |  | 3451.4 | 690 | 10:17 | 1006 | 20.5 | 294 | 17.1464 | 93632 | 305.9935 | -1.1172e-11 | 3.6573e-12 | -5.1993e-10 | 4.018e-11
Diff Time (ns): | 100 |  |  |  | 3401.4 | 680 | 10:27 | 1006 | 20.5 | 206 | 14.3527 | 85190 | 291.8733 | -9.7287e-12 | 2.5061e-12 | -3.669e-10 | 2.7818e-11
Int Time (ns): | 100 |  |  |  | 3351.6 | 670 | 10:37 | 1005 | 20.5 | 161 | 12.6886 | 63953 | 252.8893 | -6.9893e-12 | 2.0441e-12 | -2.6041e-10 | 2.0503e-11
Detector Settings |  |  |  |  | 3301.8 | 660 | 10:47 | 1005 | 20.5 | 131 | 11.4455 | 40957 | 202.3784 | -6.6092e-12 | 2.9071e-12 | -1.8263e-10 | 1.3303e-11
[5 empty] | 3252 | 650 | 10:57 | 1005 | 20.4 | 97 | 9.8489 | 28110 | 167.6604 | -4.9719e-12 | 1.7634e-12 | -1.2918e-10 | 9.0658e-12
Serial Number: | GE1/1-X-L-GHENT-0009 |  |  |  | 3202 | 640 | 11:07 | 1005 | 20.3 | 56 | 7.4833 | 11584 | 107.629 | -4.0873e-12 | 1.8483e-12 | -9.1954e-11 | 7.3798e-12
Position (ieta,iphi): | 4.2 |  |  |  | 3152.2 | 630 | 11:17 | 1005 | 20.3 | 20 | 4.4721 | 1832 | 42.8019 | -3.5018e-12 | 2.0328e-12 | -6.6796e-11 | 5.5259e-12
Gas (X/Y/Z): | Ar/CO2 |  |  |  | 3102.2 | 620 | 11:27 | 1005 | 20.3 | 12 | 3.4641 | 42 | 6.4807 | -2.9044e-12 | 2.0197e-12 | -4.8062e-11 | 3.8576e-12
Gas Frac (%X/%Y/%Z): | 70/30 |  |  |  | 3052.4 | 610 | 11:37 | 1005 | 20.2 | 7 | 2.6458 | 14 | 3.7417 | -2.4385e-12 | 1.7434e-12 | -3.4505e-11 | 2.6931e-12
Flow Rate (L/hr): | 5 |  |  |  | 3002.4 | 600 | 11:47 | 1005 | 20.2 | 7 | 2.6458 | 14 | 3.7417 | -1.8698e-12 | 1.6547e-12 | -2.5157e-11 | 2.5908e-12
HV Circuit R_Equiv (MOhm): | 5 |  |  |  | 2952.6 | 590 | 11:57 | 1005 | 20.2 | 6 | 2.4495 | 12 | 3.4641 | -2.5011e-12 | 1.2956e-26 | -1.844e-11 | 2.172e-12
Total HV Divider R (MOhm): | 4.7 |  |  |  | 2902.6 | 580 | 12:07 | 1004 | 20.1 | 5 | 2.2361 | 13 | 3.6056 | -1.2836e-12 | 1.7336e-12 | -1.3419e-11 | 2.0412e-12
R_Drift (MOhm): | Enter value |  |  |  | 2852.6 | 570 | 12:17 | 1004 | 20.1 | 5 | 2.2361 | 11 | 3.3166 | -1.0836e-12 | 1.6134e-12 | -1.0229e-11 | 1.8255e-12
R_GEM1 (MOhm): | Enter value |  |  |  | 2802.8 | 560 | 12:27 | 1004 | 20.1 | 6 | 2.4495 | 11 | 3.3166 | -1.1951e-12 | 1.7341e-12 | -7.9972e-12 | 1.8117e-12
R_Transfer1 (MOhm): | Enter value |  |  |  | 2752.8 | 550 | 12:37 | 1004 | 20.1 | 4 | 2 | 9 | 3 | -4.262e-13 | 1.8029e-12 | -5.2307e-12 | 1.5475e-12
R_GEM2 (MOhm): | Enter value
R_Transfer2 (MOhm): | Enter value |  |  |  |  | Fe55 | Cd109 | Cu X-Ray
R_GEM3 (MOhm): | Enter value |  |  | N_primary | 346 | 196 | 322 | 346 | Electron Charge |  | 1.602e-19
R_Induction (MOhm): | Enter value |  |  | Sigma N_primary | 2.9 | 1.8 | 2.8 | 2.9
Discriminator Settings |  |  |  | After T/P Corrections
Make & Model: | ORTEC 935 |  |  | Imon Equiv | V_mon | V_Drift | Rate | Rate Err | Current | Current Err | Gain | Gain Err
Threshold (mV): | -65 |  |  | (uA) | (V) | (V) | (Hz) | (Hz) | (A) | (A) | (A.U.) | (A.U.)
Walk Adjust (mV): | -0.001 |  |  | 700 | 3501 | 3290 | 1615.55 | 5.2094 | -7.2535e-10 | 6.0449e-11 | 8100.012 | 678.9453
Width (ns): | Enter value |  |  | 690 | 3451.4 | 3243 | 1555.6333 | 5.1079 | -5.0876e-10 | 4.0346e-11 | 5681.3541 | 453.4287
Picoammeter Settings |  |  |  | 680 | 3401.4 | 3196 | 1416.4 | 4.8704 | -3.5717e-10 | 2.7931e-11 | 3988.5967 | 313.9559
|  |  |  | 670 | 3351.6 | 3149 | 1063.2 | 4.2201 | -2.5342e-10 | 2.0605e-11 | 2829.9489 | 231.4975
Make & Model: | Keithley 6487 |  |  | 660 | 3301.8 | 3102 | 680.4333 | 3.3784 | -1.7602e-10 | 1.3617e-11 | 1965.6517 | 153.0884
Triax Red: | Signal |  |  | 650 | 3252 | 3055 | 466.8833 | 2.7992 | -1.2421e-10 | 9.2357e-12 | 1387.0455 | 103.8853
Triax Black: | ground |  |  | 640 | 3202 | 3008 | 192.1333 | 1.7981 | -8.7866e-11 | 7.6077e-12 | 981.2131 | 85.4118
Triax Green: | ground |  |  | 630 | 3152.2 | 2961 | 30.2 | 0.71725 | -6.3294e-11 | 5.8879e-12 | 706.8128 | 66.0569
Scalar Settings |  |  |  | 620 | 3102.2 | 2914 | 0.5 | 0.12247 | -4.5158e-11 | 4.3543e-12 | 504.2821 | 48.8356
|  |  |  | 610 | 3052.4 | 2867 | 0.11667 | 0.076376 | -3.2067e-11 | 3.2082e-12 | 358.0908 | 35.9702
|  |  |  | 600 | 3002.4 | 2820 | 0.11667 | 0.076376 | -2.3287e-11 | 3.0741e-12 | 260.0491 | 34.4081
Make & Model: | ORTEC 974 |  |  | 590 | 2952.6 | 2773 | 0.1 | 0.070711 | -1.5939e-11 | 2.172e-12 | 177.9915 | 24.3072
Acquisition Time (s): | 60 |  |  | 580 | 2902.6 | 2726 | 0.13333 | 0.070711 | -1.2136e-11 | 2.678e-12 | 135.522 | 29.9304
Source Settings |  |  |  | 570 | 2852.6 | 2679 | 0.1 | 0.066667 | -9.145e-12 | 2.4363e-12 | 102.1228 | 27.2224
|  |  |  | 560 | 2802.8 | 2632 | 0.083333 | 0.068718 | -6.8022e-12 | 2.5078e-12 | 75.9608 | 28.0137
Source: | Ag (silver) |  |  | 550 | 2752.8 | 2585 | 0.083333 | 0.060093 | -4.8045e-12 | 2.376e-12 | 53.6522 | 26.5375
HV (kV): | 40
Current (uA): | 5 |  |  | Summary |  |  | Correction parameters (P5) |  |  |  | For Ana. Framwork
Activity (MBq): | Enter value |  |  | Avrg temperature (C) | 293.4367 |  | P0 | 964.4 |  |  | 700,8100.01195955509,678.945299455347
Filter Status: | 3Cu tape layer + 5mm Al filter |  |  | Err temperature (C) | 0.15434 |  | T0 | 297.1 |  |  | 690,5681.35411580299,453.428662154527
Collimator Status: | removed |  |  | Avrg pressure (mbar) | 1004.875 | [5 empty] | 680,3988.59665349291,313.955867126686
|  |  |  | Err pressure (mbar) | 0.59948 | [5 empty] | 670,2829.94890386553,231.497513722651
|  |  |  | Expo_fit_1_param_1 | 5.1077e-07 | [5 empty] | 660,1965.65174113041,153.088412880724
|  |  |  | Expo_fit_1_param_2 | 0.033464 | [5 empty] | 650,1387.04546145074,103.885346048968
640,981.213137183948,85.4118373987007
630,706.812770704626,66.0569280467476
620,504.282054693647,48.8355956721013
610,358.090824123477,35.9702127359071
600,260.049143587326,34.408102434407
590,177.991481354473,24.3071911596919
580,135.522021299634,29.930365682234
570,102.122819182117,27.2223684247922
560,75.9607794041838,28.0136641954242
550,53.6521791363769,26.5374538151927

## 550uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-4.262e-13 | 1.8029e-12 | -5.2307e-12 | 1.5475e-12
GEMs On |  | GEMs On
-6.8212e-13 | 0.31202 | -3.8654e-12 | 0.30902
-2.0464e-12 | 0.99606 | -4.7748e-12 | 0.99306
4.5475e-13 | 1.4021 | -5.9117e-12 | 1.3981
4.0927e-12 | 1.8111 | -2.9559e-12 | 1.8041
1.5916e-12 | 2.2161 | -5.457e-12 | 2.2081
-3.1832e-12 | 2.6231 | -5.457e-12 | 2.6132
4.5475e-13 | 3.0432 | -8.1855e-12 | 3.0182
1.1369e-12 | 3.4872 | -5.457e-12 | 3.4222
-2.7285e-12 | 3.8942 | -6.5938e-12 | 3.8272
9.0949e-13 | 4.2982 | -5.2296e-12 | 4.2332
2.9559e-12 | 4.7033 | -5.2296e-12 | 4.6373
-1.819e-12 | 5.1083 | -6.5938e-12 | 5.0423
-9.0949e-13 | 5.5143 | -3.4106e-12 | 5.4513
-1.5916e-12 | 5.9183 | -2.7285e-12 | 5.8553
-9.0949e-13 | 6.3244 | -4.5475e-12 | 6.2604
4.5475e-13 | 6.7294 | -4.5475e-12 | 6.6654
-6.8212e-13 | 7.1354 | -4.7748e-12 | 7.0734
-2.2737e-13 | 7.5864 | -4.3201e-12 | 7.4804
2.5011e-12 | 7.9995 | -3.4106e-12 | 7.8875
6.8212e-13 | 8.4055 | -2.5011e-12 | 8.2965
-4.5475e-13 | 8.8235 | -4.5475e-12 | 8.7105
6.8212e-13 | 9.2285 | -5.6843e-12 | 9.2195
2.9559e-12 | 9.6336 | -6.5938e-12 | 9.6256
-2.0464e-12 | 10.0566 | -3.1832e-12 | 10.0966
-2.2737e-13 | 10.4616 | -6.8212e-12 | 10.5086
1.5916e-12 | 11.1506 | -3.8654e-12 | 10.9216
-4.5475e-13 | 11.5707 | -7.276e-12 | 11.3256
-3.1832e-12 | 11.9747 | -4.3201e-12 | 11.7307
-3.4106e-12 | 12.3907 | -7.7307e-12 | 12.1407
2.2737e-13 | 12.7967 | -5.6843e-12 | 12.5457
-2.2737e-13 | 13.2027 | -4.5475e-12 | 12.9587
-4.5475e-13 | 13.6068 | -5.457e-12 | 13.3648
2.9559e-12 | 14.0188 | -5.6843e-12 | 13.7698
4.5475e-13 | 14.4308 | -5.457e-12 | 14.1748
-2.0464e-12 | 14.8359 | -4.7748e-12 | 14.6428
-4.5475e-13 | 15.2409 | -4.5475e-12 | 15.0559
-9.0949e-13 | 15.6469 | -3.4106e-12 | 15.4619
2.0464e-12 | 16.0509 | -5.6843e-12 | 15.9519
-2.0464e-12 | 16.4559 | -4.5475e-12 | 16.3579
2.2737e-13 | 16.86 | -6.8212e-12 | 16.763
2.2737e-13 | 17.263 | -2.9559e-12 | 17.168
4.5475e-13 | 17.667 | -5.6843e-12 | 17.604
-2.2737e-12 | 18.071 | -6.3665e-12 | 18.009
6.8212e-13 | 18.4771 | -4.5475e-12 | 18.414
2.2737e-13 | 18.8801 | -4.0927e-12 | 18.8191
-1.5916e-12 | 19.2851 | -5.457e-12 | 19.2231
9.0949e-13 | 19.6911 | -4.0927e-12 | 19.6271
-2.9559e-12 | 20.0952 | -5.2296e-12 | 20.0331
-9.0949e-13 | 20.4992 | -6.5938e-12 | 20.4392
1.1369e-12 | 20.9042 | -5.2296e-12 | 20.8512
-3.4106e-12 | 21.3092 | -5.9117e-12 | 21.4722
2.9559e-12 | 21.7142 | -5.6843e-12 | 22.0363
6.8212e-13 | 22.1193 | -2.7285e-12 | 22.6543
1.819e-12 | 22.5233 | -5.9117e-12 | 23.0603
-9.0949e-13 | 22.9283 | -2.9559e-12 | 23.5853
-3.1832e-12 | 23.3313 | -7.7307e-12 | 24.2484
6.8212e-13 | 23.7364 | -6.3665e-12 | 24.7564
-1.5916e-12 | 24.1414 | -5.9117e-12 | 25.1644
9.0949e-13 | 24.5484 | -2.7285e-12 | 25.5915
-3.638e-12 | 24.9514 | -5.9117e-12 | 26.1985
2.2737e-13 | 25.3555 | -3.4106e-12 | 26.7065
1.1369e-12 | 25.7605 | -4.3201e-12 | 27.3266
1.1369e-12 | 26.1665 | -4.5475e-12 | 27.8166
1.5916e-12 | 26.5715 | -4.3201e-12 | 28.2396
-9.0949e-13 | 26.9775 | -4.3201e-12 | 28.6446
-6.8212e-13 | 27.3816 | -5.6843e-12 | 29.0507
1.819e-12 | 27.7846 | -6.3665e-12 | 29.4717
-2.2737e-13 | 28.1896 | -5.457e-12 | 29.8777
-1.819e-12 | 28.5936 | -6.3665e-12 | 30.2847
4.5475e-13 | 28.9977 | -7.276e-12 | 30.6918
6.8212e-13 | 29.4017 | -5.2296e-12 | 31.0958
-1.819e-12 | 29.8067 | -7.0486e-12 | 31.5198
-2.7285e-12 | 30.2117 | -3.4106e-12 | 32.0698
-4.5475e-13 | 30.6178 | -4.5475e-12 | 32.9309
-1.3642e-12 | 31.0218 | -6.8212e-12 | 33.3359
2.7285e-12 | 31.4258 | -4.7748e-12 | 34.13
-2.0464e-12 | 31.8308 | -2.5011e-12 | 34.546
-6.8212e-13 | 32.2368 | -3.8654e-12 | 35.003
-1.819e-12 | 32.6409 | -4.0927e-12 | 35.425
2.5011e-12 | 33.0459 | -6.8212e-12 | 35.83
-9.0949e-13 | 33.4499 | -1.0914e-11 | 36.2351
4.5475e-13 | 33.8549 | -7.0486e-12 | 36.6411
-6.8212e-13 | 34.259 | -8.1855e-12 | 37.0461
-1.819e-12 | 34.663 | -4.3201e-12 | 37.4581
-2.2737e-13 | 35.068 | -5.2296e-12 | 37.8642
-1.5916e-12 | 35.472 | -4.0927e-12 | 38.2702
-1.819e-12 | 35.877 | -7.276e-12 | 38.6782
-9.0949e-13 | 36.2811 | -8.4128e-12 | 39.0852
4.5475e-13 | 36.6861 | -7.9581e-12 | 39.4903
1.819e-12 | 37.0891 | -7.7307e-12 | 39.8943
-2.9559e-12 | 37.4941 | -3.8654e-12 | 40.2983
-1.819e-12 | 37.8992 | -5.457e-12 | 40.7033
4.5475e-13 | 38.3042 | -6.5938e-12 | 41.1093
2.2737e-13 | 38.7072 | -3.4106e-12 | 41.5134
-4.5475e-12 | 39.1112 | -4.5475e-12 | 41.9194
-2.2737e-13 | 39.5173 | -4.5475e-12 | 42.3294
1.5916e-12 | 39.9213 | -4.3201e-12 | 42.7344
1.819e-12 | 40.3273 | -3.1832e-12 | 43.1395
-3.638e-12 | 40.7323 | -4.5475e-12 | 43.5435
6.8212e-13 | 41.1364 | -5.457e-12 | 43.9525
-4.5475e-13 | 41.5414 | -3.4106e-12 | 44.3585
-1.819e-12 | 41.9464 | -5.9117e-12 | 44.7836
2.9559e-12 | 42.3524 | -5.2296e-12 | 45.1896
-1.819e-12 | 42.7555 | -2.7285e-12 | 45.5956
-3.638e-12 | 43.1605 | -6.3665e-12 | 46.0026
-3.4106e-12 | 43.5665 | -5.2296e-12 | 46.4177
-2.9559e-12 | 43.9725 | -6.3665e-12 | 46.8247
-4.5475e-13 | 44.3765 | -8.1855e-12 | 47.2307
-2.0464e-12 | 44.7816 | -2.9559e-12 | 47.6347
9.0949e-13 | 45.1876 | -5.6843e-12 | 48.0388
-2.7285e-12 | 45.5936 | -5.457e-12 | 48.4548
6.8212e-13 | 46.0146 | -2.9559e-12 | 48.8618
-1.3642e-12 | 46.4196 | -4.3201e-12 | 49.2698
9.0949e-13 | 46.8247 | -5.457e-12 | 49.6928
9.0949e-13 | 47.2287 | -5.457e-12 | 50.0979
-6.8212e-13 | 47.6337 | -6.3665e-12 | 50.5029
1.1369e-12 | 48.0498 | -7.276e-12 | 50.9089
1.5916e-12 | 48.4638 | -4.5475e-12 | 51.3139
-2.0464e-12 | 48.8728 | -7.0486e-12 | 51.719
-2.2737e-12 | 49.2768 | -6.3665e-12 | 52.123
2.5011e-12 | 49.6818 | -3.8654e-12 | 52.53
-2.9559e-12 | 50.0879 | -5.457e-12 | 53.316
-1.5916e-12 | 50.4929 | -6.3665e-12 | 53.7241
2.5011e-12 | 50.8979 | -3.4106e-12 | 54.1291
-4.5475e-13 | 51.3019 | -5.6843e-12 | 54.5371
-2.0464e-12 | 51.706 | -3.4106e-12 | 55.2372
9.0949e-13 | 52.111 | -7.0486e-12 | 55.8442
-1.5916e-12 | 52.515 | -5.6843e-12 | 56.2492
1.819e-12 | 52.92 | -6.3665e-12 | 56.6532
4.5475e-13 | 53.325 | -4.7748e-12 | 57.0583
1.5916e-12 | 53.7291 | -4.7748e-12 | 57.4623
-1.5916e-12 | 54.1331 | -3.8654e-12 | 57.8683
6.8212e-13 | 54.5391 | -5.2296e-12 | 58.2733
-2.9559e-12 | 54.9451 | -7.0486e-12 | 58.6784
2.0464e-12 | 55.3512 | -5.9117e-12 | 59.0844
4.5475e-13 | 55.7562 | -5.9117e-12 | 59.4904
-4.5475e-13 | 56.1612 | -7.0486e-12 | 59.8954
9.0949e-13 | 56.5642 | -5.457e-12 | 60.2995
6.8212e-13 | 56.9693 | -6.3665e-12 | 60.7045
-6.8212e-13 | 57.3733 | -6.5938e-12 | 61.1105
-1.3642e-12 | 57.7763 | -3.8654e-12 | 61.5155
-4.0927e-12 | 58.1813 | -6.3665e-12 | 61.9205
-6.8212e-13 | 58.5864 | -5.9117e-12 | 62.3266
-4.5475e-13 | 58.9924 | -5.9117e-12 | 62.7326
-2.0464e-12 | 59.3984 | -2.7285e-12 | 63.1366
-1.5916e-12 | 59.8034 | -4.0927e-12 | 63.5426
-2.9559e-12 | 60.2074 | -4.7748e-12 | 63.9487
3.4106e-12 | 60.6125 | -5.6843e-12 | 64.3547
4.5475e-13 | 61.0175 | -5.9117e-12 | 64.7597
-9.0949e-13 | 61.4215 | -9.0949e-12 | 65.1647
-3.4106e-12 | 61.8255 | -4.0927e-12 | 65.5708
9.0949e-13 | 62.2316 | -6.8212e-12 | 65.9748
-2.2737e-13 | 62.6346 | -3.8654e-12 | 66.3788
-4.5475e-13 | 63.0386 | -4.7748e-12 | 66.7848
-6.8212e-13 | 63.4416 | -3.8654e-12 | 67.1898
-2.9559e-12 | 63.8456 | -2.9559e-12 | 67.5949
-2.2737e-12 | 64.2537 | -5.457e-12 | 68.0009
-2.9559e-12 | 64.6647 | -4.3201e-12 | 68.4059
2.2737e-13 | 65.0697 | -2.0464e-12 | 68.8109
6.8212e-13 | 65.4727 | -4.3201e-12 | 69.215
2.5011e-12 | 65.8758 | -4.7748e-12 | 69.621
-2.7285e-12 | 66.2798 | -2.9559e-12 | 70.026
-9.0949e-13 | 66.6838 | -4.5475e-12 | 70.431
-3.1832e-12 | 67.0878 | -4.5475e-12 | 70.835
-2.2737e-13 | 67.4919 | -5.457e-12 | 71.2401
-4.5475e-13 | 67.8949 | -6.3665e-12 | 71.6461
2.2737e-13 | 68.2999 | -6.8212e-12 | 72.0511
-2.2737e-13 | 68.7039 | -1.5916e-12 | 72.4561
-9.0949e-13 | 69.1069 | -7.276e-12 | 72.8612
-1.3642e-12 | 69.521 | -7.0486e-12 | 73.2672
-9.0949e-13 | 69.926 | -7.276e-12 | 73.6712
-9.0949e-13 | 70.33 | -3.8654e-12 | 74.0772
-1.5916e-12 | 70.7861 | -2.5011e-12 | 74.4843
4.5475e-13 | 71.1901 | -3.4106e-12 | 74.8893
2.0464e-12 | 71.5941 | -5.457e-12 | 75.2943
2.5011e-12 | 71.9991 | -5.6843e-12 | 75.6993
-2.7285e-12 | 72.4051 | -6.3665e-12 | 76.1043
-1.5916e-12 | 72.8082 | -5.6843e-12 | 76.5084
2.0464e-12 | 73.2132 | -4.7748e-12 | 76.9134
-2.0464e-12 | 73.6182 | -4.5475e-12 | 77.3194
-4.5475e-13 | 74.0232 | -7.9581e-12 | 77.7245
-2.2737e-12 | 74.4283 | -1.819e-12 | 78.1285
-6.8212e-13 | 74.8333 | -6.5938e-12 | 78.5335
1.819e-12 | 75.2383 | -4.0927e-12 | 78.9385
-9.0949e-13 | 75.6423 | -8.4128e-12 | 79.3435
-3.1832e-12 | 76.0464 | -8.6402e-12 | 79.7486
-1.5916e-12 | 76.4514 | -3.4106e-12 | 80.1536
-2.9559e-12 | 76.8564 | -5.457e-12 | 80.5586
-4.5475e-13 | 77.2604 | -4.3201e-12 | 80.9626
-4.5475e-13 | 77.6684 | -5.9117e-12 | 81.3676
-3.4106e-12 | 78.0725 | -4.7748e-12 | 81.7737
-2.2737e-13 | 78.4785 | -6.3665e-12 | 82.2447
1.819e-12 | 78.8825 | -3.4106e-12 | 82.6497
-2.9559e-12 | 79.2885 | -2.7285e-12 | 83.0558
-2.2737e-12 | 79.6986 | -6.3665e-12 | 83.4598
1.1369e-12 | 80.1576 | -7.0486e-12 | 83.8648
-1.819e-12 | 80.5656 | -2.9559e-12 | 84.2928
-2.2737e-13 | 80.9716 | -4.7748e-12 | 84.7318
-1.5916e-12 | 81.3766 | -5.2296e-12 | 85.2149
4.5475e-13 | 81.7837 | -6.8212e-12 | 85.6199
-9.0949e-13 | 82.1947 | -4.5475e-12 | 86.0249
1.5916e-12 | 82.6437 | -4.0927e-12 | 86.4519
-1.5916e-12 | 83.0487 | -4.7748e-12 | 86.857
2.0464e-12 | 83.4788 | -5.9117e-12 | 87.263
2.2737e-13 | 83.9108 | -2.7285e-12 | 87.668
-4.5475e-13 | 84.3148 | -7.0486e-12 | 88.075
1.5916e-12 | 84.7198 | -6.8212e-12 | 88.4801
-2.0464e-12 | 85.1519 | -5.457e-12 | 88.8841
-2.2737e-12 | 85.5619 | -6.3665e-12 | 89.2891
1.5916e-12 | 85.9679
2.9559e-12 | 86.3739
2.2737e-13 | 86.779
1.1369e-12 | 87.221
-1.819e-12 | 87.638
-3.1832e-12 | 88.044
2.5011e-12 | 88.4501
2.2737e-13 | 88.8541
-2.9559e-12 | 89.2591
-9.0949e-13 | 89.6641
2.5011e-12 | 90.0691
-5.457e-12 | 90.4742
-2.2737e-13 | 90.9532
1.819e-12 | 91.3602
9.0949e-13 | 91.9363
1.1369e-12 | 92.3403
2.5011e-12 | 92.7503
2.2737e-13 | 93.1563
-6.8212e-13 | 93.5603
-3.1832e-12 | 93.9894
1.1369e-12 | 94.3954
9.0949e-13 | 94.8044

## 560uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.1951e-12 | 1.7341e-12 | -7.9972e-12 | 1.8117e-12
GEMs On |  | GEMs On
6.8212e-13 | 0.30902 | -7.7307e-12 | 0.31402
-2.7285e-12 | 0.99306 | -6.8212e-12 | 0.99806
-4.5475e-13 | 1.3981 | -9.5497e-12 | 1.4051
-2.2737e-12 | 1.8021 | -1.1596e-11 | 1.8101
6.8212e-13 | 2.2071 | -7.7307e-12 | 2.2141
-6.8212e-13 | 2.6101 | -9.5497e-12 | 2.6191
-2.5011e-12 | 3.0162 | -8.6402e-12 | 3.0242
-2.2737e-13 | 3.4212 | -1.0232e-11 | 3.4302
-2.9559e-12 | 3.8282 | -3.8654e-12 | 3.8352
2.2737e-13 | 4.2372 | -9.3223e-12 | 4.2402
-2.2737e-12 | 4.6413 | -9.3223e-12 | 4.6453
-1.5916e-12 | 5.0473 | -5.457e-12 | 5.0543
-2.5011e-12 | 5.4533 | -9.3223e-12 | 5.4593
-2.7285e-12 | 5.8613 | -1.0232e-11 | 5.8653
-5.0022e-12 | 6.2674 | -8.1855e-12 | 6.2704
-2.9559e-12 | 6.6734 | -5.6843e-12 | 6.6754
-4.0927e-12 | 7.0794 | -7.5033e-12 | 7.0814
-2.2737e-12 | 7.4854 | -8.1855e-12 | 7.4964
6.8212e-13 | 7.8905 | -1.0687e-11 | 7.9075
2.0464e-12 | 8.2945 | -6.8212e-12 | 8.3145
-1.3642e-12 | 8.6985 | -8.8676e-12 | 8.7195
1.1369e-12 | 9.1035 | -9.3223e-12 | 9.1265
1.3642e-12 | 9.5085 | -8.6402e-12 | 9.5315
-2.2737e-13 | 9.9136 | -1.1369e-11 | 9.9396
-4.5475e-12 | 10.3196 | -9.5497e-12 | 10.3466
-2.9559e-12 | 10.7256 | -5.6843e-12 | 10.7526
-1.5916e-12 | 11.1406 | -4.0927e-12 | 11.1586
-1.819e-12 | 11.5457 | -6.3665e-12 | 11.5647
-3.8654e-12 | 11.9497 | -5.6843e-12 | 11.9707
9.0949e-13 | 12.3557 | -5.2296e-12 | 12.3767
-2.7285e-12 | 12.7607 | -7.0486e-12 | 12.7837
-2.9559e-12 | 13.1657 | -7.7307e-12 | 13.2008
-1.3642e-12 | 13.5758 | -7.5033e-12 | 13.6048
9.0949e-13 | 13.9808 | -7.9581e-12 | 14.0088
-2.2737e-13 | 14.4078 | -7.9581e-12 | 14.4168
2.2737e-13 | 14.8118 | -8.1855e-12 | 14.8229
-4.5475e-13 | 15.2189 | -5.6843e-12 | 15.2289
-4.5475e-13 | 15.6229 | -7.0486e-12 | 15.6369
2.2737e-13 | 16.0389 | -5.0022e-12 | 16.0439
-1.3642e-12 | 16.4469 | -9.5497e-12 | 16.4489
2.2737e-13 | 16.859 | -8.1855e-12 | 16.946
-4.5475e-13 | 17.265 | -6.3665e-12 | 17.352
-1.3642e-12 | 17.671 | -1.0232e-11 | 17.758
1.1369e-12 | 18.075 | -6.8212e-12 | 18.164
-1.5916e-12 | 18.4801 | -8.1855e-12 | 18.7191
-2.5011e-12 | 18.8841 | -1.0459e-11 | 19.1301
9.0949e-13 | 19.2961 | -1.2506e-11 | 19.5711
-6.8212e-13 | 19.7041 | -8.6402e-12 | 20.1001
-2.9559e-12 | 20.1141 | -1.0232e-11 | 20.5052
-1.5916e-12 | 20.5202 | -8.6402e-12 | 20.9112
-4.3201e-12 | 20.9262 | -8.6402e-12 | 21.3142
-2.7285e-12 | 21.3342 | -8.6402e-12 | 21.7772
2.7285e-12 | 21.7392 | -1.0232e-11 | 22.1863
6.8212e-13 | 22.1443 | -4.5475e-12 | 22.6113
-4.5475e-13 | 22.5493 | -9.0949e-12 | 23.0153
-4.5475e-13 | 22.9543 | -7.5033e-12 | 23.4213
2.7285e-12 | 23.3593 | -7.0486e-12 | 23.8274
6.8212e-13 | 23.7654 | -8.6402e-12 | 24.2324
-1.3642e-12 | 24.1704 | -8.8676e-12 | 24.6394
1.1369e-12 | 24.5764 | -7.5033e-12 | 25.0444
-2.2737e-12 | 24.9814 | -6.8212e-12 | 25.4495
-1.819e-12 | 25.3875 | -1.0687e-11 | 25.8535
-1.1369e-12 | 25.7915 | -1.1596e-11 | 26.2585
1.1369e-12 | 26.1985 | -6.8212e-12 | 26.6655
-1.1369e-12 | 26.6045 | -7.9581e-12 | 27.0715
2.0464e-12 | 27.0085 | -8.8676e-12 | 27.4786
1.3642e-12 | 27.4146 | -8.8676e-12 | 27.8846
-3.638e-12 | 27.8206 | -6.3665e-12 | 28.2886
-2.2737e-13 | 28.2276 | -5.6843e-12 | 28.6936
-5.0022e-12 | 28.6306 | -5.2296e-12 | 29.0997
6.8212e-13 | 29.0357 | -4.5475e-12 | 29.5047
-2.2737e-13 | 29.4557 | -5.6843e-12 | 29.9097
-1.5916e-12 | 29.8677 | -6.1391e-12 | 30.3157
-1.819e-12 | 30.2727 | -5.0022e-12 | 30.7308
0 | 30.6777 | -6.8212e-12 | 31.1358
9.0949e-13 | 31.0958 | -5.6843e-12 | 31.5418
-2.9559e-12 | 31.5048 | -1.0232e-11 | 31.9498
-3.8654e-12 | 31.9088 | -6.1391e-12 | 32.3558
9.0949e-13 | 32.3148 | -7.5033e-12 | 32.7599
-3.638e-12 | 32.7199 | -1.0004e-11 | 33.1649
-6.8212e-13 | 33.1249 | -8.1855e-12 | 33.5709
-1.819e-12 | 33.5289 | -6.8212e-12 | 33.9759
-4.5475e-12 | 33.9339 | -8.6402e-12 | 34.381
-1.3642e-12 | 34.34 | -9.3223e-12 | 34.786
-2.2737e-12 | 34.749 | -1.1369e-11 | 35.191
0 | 35.155 | -9.0949e-12 | 35.597
-4.5475e-13 | 35.557 | -1.0687e-11 | 36.0021
0 | 35.9611 | -8.8676e-12 | 36.4071
-2.2737e-12 | 36.3651 | -1.2051e-11 | 36.8121
-3.1832e-12 | 36.7711 | -7.5033e-12 | 37.2181
-1.3642e-12 | 37.1831 | -8.1855e-12 | 37.6251
6.8212e-13 | 37.5891 | -7.7307e-12 | 38.0292
-2.7285e-12 | 38.0052 | -7.9581e-12 | 38.4352
0 | 38.4142 | -8.8676e-12 | 38.8402
-1.819e-12 | 38.8182 | -6.3665e-12 | 39.2452
-2.5011e-12 | 39.2322 | -8.1855e-12 | 39.6513
-6.8212e-13 | 39.6373 | -9.3223e-12 | 40.0563
-2.5011e-12 | 40.0433 | -5.6843e-12 | 40.4623
-4.5475e-12 | 40.4693 | -5.0022e-12 | 40.8673
6.8212e-13 | 40.8743 | -1.0232e-11 | 41.2724
6.8212e-13 | 41.2784 | -9.3223e-12 | 41.6784
-1.5916e-12 | 41.6834 | -1.0459e-11 | 42.0834
2.0464e-12 | 42.0884 | -7.0486e-12 | 42.4894
9.0949e-13 | 42.4934 | -8.1855e-12 | 42.8944
6.8212e-13 | 42.8974 | -6.3665e-12 | 43.2995
-4.5475e-13 | 43.3025 | -7.5033e-12 | 43.7055
-1.819e-12 | 43.7075 | -9.0949e-12 | 44.1115
-2.2737e-13 | 44.1125 | -5.6843e-12 | 44.5166
-4.5475e-13 | 44.5185 | -7.0486e-12 | 44.9216
-2.2737e-12 | 44.9226 | -7.0486e-12 | 45.3276
-1.3642e-12 | 45.3396 | -1.0914e-11 | 45.7326
-6.8212e-13 | 45.7476 | -7.9581e-12 | 46.1396
0 | 46.1516 | -8.1855e-12 | 46.5437
-2.9559e-12 | 46.5567 | -7.7307e-12 | 46.9477
-3.1832e-12 | 46.9637 | -6.5938e-12 | 47.3517
-4.5475e-13 | 47.3677 | -5.0022e-12 | 47.7567
-2.5011e-12 | 47.7727 | -1.0459e-11 | 48.1617
-2.9559e-12 | 48.1848 | -1.0004e-11 | 48.5678
-3.1832e-12 | 48.6028 | -5.6843e-12 | 48.9728
2.0464e-12 | 49.0078 | -8.1855e-12 | 49.3798
1.5916e-12 | 49.4128 | -7.7307e-12 | 49.7848
-1.5916e-12 | 49.8178 | -1.0459e-11 | 50.1899
9.0949e-13 | 50.2269 | -7.0486e-12 | 50.5939
-2.2737e-13 | 50.6299 | -7.7307e-12 | 50.9989
-2.5011e-12 | 51.0339 | -6.8212e-12 | 51.4029
1.1369e-12 | 51.4379 | -6.8212e-12 | 51.808
-2.2737e-12 | 51.842 | -6.8212e-12 | 52.213
-5.0022e-12 | 52.246 | -7.9581e-12 | 52.617
2.2737e-13 | 52.65 | -8.6402e-12 | 53.021
-2.5011e-12 | 53.054 | -8.8676e-12 | 53.4271
-2.9559e-12 | 53.4591 | -6.5938e-12 | 53.8311
-2.7285e-12 | 53.8641 | -6.1391e-12 | 54.2361
9.0949e-13 | 54.2681 | -7.7307e-12 | 54.6391
9.0949e-13 | 54.6721 | -9.0949e-12 | 55.0442
-4.5475e-13 | 55.0761 | -9.5497e-12 | 55.4492
-1.5916e-12 | 55.4802 | -6.8212e-12 | 55.8542
-1.5916e-12 | 55.8872 | -8.8676e-12 | 56.2612
-1.819e-12 | 56.2912 | -7.9581e-12 | 56.6692
-3.1832e-12 | 56.6982 | -7.7307e-12 | 57.0753
-2.9559e-12 | 57.1023 | -7.7307e-12 | 57.4793
-2.2737e-13 | 57.5063 | -8.6402e-12 | 57.8843
0 | 57.9233 | -1.0687e-11 | 58.2893
-3.638e-12 | 58.3273 | -7.7307e-12 | 58.7014
2.2737e-13 | 58.7324 | -7.5033e-12 | 59.1104
2.2737e-13 | 59.1474 | -4.0927e-12 | 59.5154
-1.819e-12 | 59.5564 | -7.0486e-12 | 59.9204
0 | 59.9664 | -5.2296e-12 | 60.3254
-4.5475e-13 | 60.3704 | -8.8676e-12 | 60.7305
-4.5475e-13 | 60.7745 | -7.9581e-12 | 61.1355
-4.0927e-12 | 61.1785 | -9.3223e-12 | 61.5415
0 | 61.5835 | -7.9581e-12 | 61.9475
-2.2737e-12 | 61.9875 | -6.3665e-12 | 62.3526
-3.8654e-12 | 62.3936 | -1.1596e-11 | 62.7576
-1.819e-12 | 62.7986 | -6.5938e-12 | 63.1636
-2.7285e-12 | 63.2036 | -1.2506e-11 | 63.5686
-1.819e-12 | 63.6076 | -6.8212e-12 | 63.9747
6.8212e-13 | 64.0137 | -1.0687e-11 | 64.3797
-1.5916e-12 | 64.4197 | -7.5033e-12 | 64.7857
0 | 64.8237 | -6.3665e-12 | 65.1887
0 | 65.2267 | -8.8676e-12 | 65.5948
6.8212e-13 | 65.6287 | -6.8212e-12 | 65.9988
-1.3642e-12 | 66.0328 | -6.3665e-12 | 66.4038
-1.1369e-12 | 66.4398 | -7.0486e-12 | 66.8078
2.0464e-12 | 66.8448 | -6.1391e-12 | 67.2128
-1.819e-12 | 67.2489 | -9.0949e-12 | 67.6179
-1.819e-12 | 67.6549 | -7.0486e-12 | 68.0229
9.0949e-13 | 68.0589 | -1.0004e-11 | 68.4279
2.2737e-13 | 68.4619 | -9.0949e-12 | 68.8339
-1.819e-12 | 68.8659 | -7.9581e-12 | 69.239
-3.638e-12 | 69.272 | -8.8676e-12 | 69.645
9.0949e-13 | 69.676 | -5.0022e-12 | 70.05
3.1832e-12 | 70.082 | -9.3223e-12 | 70.454
-2.5011e-12 | 70.487 | -5.2296e-12 | 70.859
-4.5475e-12 | 70.8931 | -7.9581e-12 | 71.2631
-1.819e-12 | 71.2981 | -9.3223e-12 | 71.6671
-2.2737e-13 | 71.7021 | -5.2296e-12 | 72.0721
-2.7285e-12 | 72.1071 | -9.0949e-12 | 72.4771
-1.1369e-12 | 72.5122 | -6.8212e-12 | 72.8822
1.3642e-12 | 72.9172 | -7.0486e-12 | 73.2862
-1.3642e-12 | 73.3212 | -1.0004e-11 | 73.6912
-1.3642e-12 | 73.7262 | -5.457e-12 | 74.0962
-2.2737e-13 | 74.1312 | -1.0687e-11 | 74.5013
-1.819e-12 | 74.5353 | -9.0949e-12 | 74.9073
-1.5916e-12 | 74.9403 | -8.1855e-12 | 75.3133
-1.1369e-12 | 75.3443 | -1.0914e-11 | 75.7193
-3.8654e-12 | 75.7493 | -7.9581e-12 | 76.1253
-3.8654e-12 | 76.1554 | -6.1391e-12 | 76.5324
-2.9559e-12 | 76.5604 | -2.9559e-12 | 76.9364
-4.3201e-12 | 76.9694 | -8.1855e-12 | 77.3414
1.3642e-12 | 77.3744 | -6.8212e-12 | 77.7455
-6.8212e-13 | 77.7785 | -7.5033e-12 | 78.1505
-1.819e-12 | 78.1825 | -8.8676e-12 | 78.5555
-1.5916e-12 | 78.5885 | -1.0914e-11 | 78.9605
-6.8212e-13 | 78.9945 | -6.3665e-12 | 79.3675
-3.1832e-12 | 79.3985 | -9.5497e-12 | 79.7726
-3.1832e-12 | 79.8026 | -1.0459e-11 | 80.1776
-2.7285e-12 | 80.2076 | -7.0486e-12 | 80.5826
0 | 80.6136 | -7.7307e-12 | 80.9886
-6.8212e-13 | 81.0156 | -8.6402e-12 | 81.3937
-2.9559e-12 | 81.4207 | -9.0949e-12 | 81.7987
-2.5011e-12 | 81.8257 | -1.0459e-11 | 82.2037
-2.2737e-13 | 82.2537 | -7.5033e-12 | 82.6307
-1.3642e-12 | 82.6577 | -1.0004e-11 | 83.0357
-1.5916e-12 | 83.0627 | -9.5497e-12 | 83.4418
-2.2737e-12 | 83.4678 | -1.0232e-11 | 83.8458
9.0949e-13 | 83.8728 | -5.6843e-12 | 84.2518
-2.9559e-12 | 84.2778 | -7.9581e-12 | 84.6558
-1.3642e-12 | 84.6818 | -5.6843e-12 | 85.0619
1.5916e-12 | 85.0869 | -7.7307e-12 | 85.4669
-3.638e-12 | 85.4929
-1.819e-12 | 85.8979

## 570uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.0836e-12 | 1.6134e-12 | -1.0229e-11 | 1.8255e-12
GEMs On |  | GEMs On
-1.3642e-12 | 0.31202 | -1.0914e-11 | 0.31102
-1.3642e-12 | 0.99606 | -1.1596e-11 | 0.99606
-1.5916e-12 | 1.4011 | -1.1369e-11 | 1.4021
6.8212e-13 | 1.8061 | -3.1832e-12 | 1.8071
-1.3642e-12 | 2.2101 | -1.0232e-11 | 2.2121
9.0949e-13 | 2.6141 | -7.276e-12 | 2.6172
2.0464e-12 | 3.0202 | -9.7771e-12 | 3.0242
9.0949e-13 | 3.4252 | -1.1369e-11 | 3.4292
1.1369e-12 | 3.8282 | -9.7771e-12 | 3.8342
9.0949e-13 | 4.2352 | -1.0914e-11 | 4.2392
0 | 4.6383 | -9.5497e-12 | 4.6443
-1.1369e-12 | 5.0423 | -1.0004e-11 | 5.0493
-4.5475e-13 | 5.4463 | -1.2278e-11 | 5.4553
-5.0022e-12 | 5.8503 | -1.5007e-11 | 5.8603
-2.2737e-13 | 6.2554 | -1.2278e-11 | 6.2674
-1.3642e-12 | 6.6594 | -8.1855e-12 | 6.6714
0 | 7.0664 | -1.0459e-11 | 7.0764
2.2737e-13 | 7.4684 | -9.5497e-12 | 7.4824
-2.5011e-12 | 7.8724 | -8.4128e-12 | 7.8875
1.1369e-12 | 8.2775 | -1.2278e-11 | 8.2915
-4.0927e-12 | 8.6815 | -1.296e-11 | 8.6975
-6.8212e-13 | 9.0865 | -8.1855e-12 | 9.1035
2.2737e-13 | 9.4915 | -1.1369e-11 | 9.5085
-2.0464e-12 | 9.8956 | -1.0459e-11 | 9.9146
-1.819e-12 | 10.3006 | -1.0232e-11 | 10.3206
1.1369e-12 | 10.7056 | -9.7771e-12 | 10.7266
-4.5475e-13 | 11.1106 | -8.1855e-12 | 11.1316
-1.819e-12 | 11.5147 | -1.0914e-11 | 11.5367
-2.9559e-12 | 11.9197 | -9.7771e-12 | 11.9407
0 | 12.3247 | -1.2278e-11 | 12.3457
-1.1369e-12 | 12.7277 | -7.5033e-12 | 12.7517
-2.9559e-12 | 13.1327 | -1.2278e-11 | 13.1557
1.1369e-12 | 13.5388 | -1.2733e-11 | 13.5608
-2.9559e-12 | 13.9428 | -8.6402e-12 | 13.9648
-4.5475e-13 | 14.3478 | -1.3642e-11 | 14.3698
0 | 14.7518 | -1.1369e-11 | 14.7749
-1.5916e-12 | 15.1549 | -7.5033e-12 | 15.1819
2.2737e-13 | 15.5599 | -1.0004e-11 | 15.5869
1.3642e-12 | 15.9629 | -1.0459e-11 | 15.9919
-6.8212e-13 | 16.3679 | -9.7771e-12 | 16.3989
6.8212e-13 | 16.77 | -8.6402e-12 | 16.803
9.0949e-13 | 17.175 | -9.7771e-12 | 17.209
-2.7285e-12 | 17.58 | -1.2051e-11 | 17.613
-3.638e-12 | 17.983 | -1.2051e-11 | 18.018
-2.0464e-12 | 18.388 | -1.0232e-11 | 18.4241
-1.819e-12 | 18.7921 | -9.0949e-12 | 18.8291
1.3642e-12 | 19.1991 | -8.1855e-12 | 19.2341
-6.8212e-13 | 19.6041 | -1.0459e-11 | 19.6381
-1.3642e-12 | 20.0071 | -9.7771e-12 | 20.0441
-2.2737e-13 | 20.4122 | -1.2278e-11 | 20.4502
1.1369e-12 | 20.8172 | -9.7771e-12 | 20.8552
-1.819e-12 | 21.2212 | -1.0914e-11 | 21.2602
-2.5011e-12 | 21.6262 | -9.7771e-12 | 21.6642
9.0949e-13 | 22.0303 | -9.7771e-12 | 22.0683
2.0464e-12 | 22.4343 | -8.8676e-12 | 22.4743
-2.5011e-12 | 22.8393 | -1.0459e-11 | 22.8793
0 | 23.2443 | -1.2506e-11 | 23.2853
-2.7285e-12 | 23.6473 | -6.1391e-12 | 23.6924
-2.2737e-13 | 24.0514 | -9.5497e-12 | 24.0974
-2.0464e-12 | 24.4564 | -1.0232e-11 | 24.5024
-4.5475e-13 | 24.8604 | -1.0232e-11 | 24.9084
-4.5475e-13 | 25.2644 | -8.8676e-12 | 25.3134
-4.5475e-12 | 25.6695 | -1.0459e-11 | 25.7175
-2.5011e-12 | 26.0745 | -1.0914e-11 | 26.1235
-6.8212e-13 | 26.4795 | -1.0914e-11 | 26.5285
-1.819e-12 | 26.8845 | -9.7771e-12 | 26.9335
-6.8212e-13 | 27.2896 | -1.0004e-11 | 27.3386
-2.7285e-12 | 27.6936 | -1.1369e-11 | 27.7436
-6.8212e-13 | 28.0986 | -1.1141e-11 | 28.1496
-1.5916e-12 | 28.5006 | -1.2733e-11 | 28.5566
-6.8212e-13 | 28.9077 | -9.0949e-12 | 28.9617
-4.5475e-13 | 29.3117 | -8.1855e-12 | 29.3677
-2.5011e-12 | 29.7167 | -1.3415e-11 | 29.7727
-2.9559e-12 | 30.1207 | -1.2051e-11 | 30.1777
-2.0464e-12 | 30.5257 | -1.1141e-11 | 30.5828
-2.7285e-12 | 30.9308 | -8.4128e-12 | 30.9888
-2.7285e-12 | 31.3358 | -1.0232e-11 | 31.3948
-1.1369e-12 | 31.7408 | -7.7307e-12 | 31.7998
-2.0464e-12 | 32.1458 | -1.1596e-11 | 32.2038
-2.5011e-12 | 32.5499 | -7.7307e-12 | 32.6089
-2.0464e-12 | 32.9549 | -9.5497e-12 | 33.0149
-2.9559e-12 | 33.3589 | -1.2506e-11 | 33.4199
-6.8212e-13 | 33.7639 | -1.0004e-11 | 33.8259
2.2737e-12 | 34.1669 | -1.1141e-11 | 34.232
-1.1369e-12 | 34.572 | -1.2051e-11 | 34.637
-2.9559e-12 | 34.978 | -1.0914e-11 | 35.042
-1.819e-12 | 35.383 | -9.7771e-12 | 35.447
-1.819e-12 | 35.788 | -1.0004e-11 | 35.852
3.1832e-12 | 36.1931 | -1.296e-11 | 36.2581
-6.8212e-13 | 36.5991 | -1.3415e-11 | 36.6631
1.1369e-12 | 37.0041 | -1.1596e-11 | 37.0691
-2.0464e-12 | 37.4081 | -1.0232e-11 | 37.4751
6.8212e-13 | 37.8132 | -1.2278e-11 | 37.8812
-3.8654e-12 | 38.2172 | -9.0949e-12 | 38.2882
-1.5916e-12 | 38.6222 | -1.0459e-11 | 38.6932
-1.5916e-12 | 39.0282 | -1.387e-11 | 39.0982
-4.5475e-13 | 39.4323 | -1.2733e-11 | 39.5043
2.2737e-13 | 39.8383 | -1.0232e-11 | 39.9083
1.3642e-12 | 40.2433 | -1.2278e-11 | 40.3143
0 | 40.6483 | -7.7307e-12 | 40.7203
-1.5916e-12 | 41.0504 | -9.0949e-12 | 41.1253
-3.638e-12 | 41.4544 | -9.7771e-12 | 41.5294
-1.1369e-12 | 41.8584 | -8.6402e-12 | 41.9344
-2.7285e-12 | 42.2624 | -9.7771e-12 | 42.3394
-2.0464e-12 | 42.6684 | -8.6402e-12 | 42.7444
0 | 43.0735 | -8.6402e-12 | 43.1475
-3.638e-12 | 43.4765 | -9.7771e-12 | 43.5525
2.5011e-12 | 43.8825 | -1.3415e-11 | 43.9575
0 | 44.2875 | -7.7307e-12 | 44.3635
6.8212e-13 | 44.6906 | -1.3642e-11 | 44.7686
-2.7285e-12 | 45.0936 | -8.8676e-12 | 45.1746
2.2737e-13 | 45.4976 | -8.6402e-12 | 45.5806
-6.8212e-13 | 45.9016 | -1.1369e-11 | 45.9846
-1.819e-12 | 46.3047 | -9.5497e-12 | 46.3907
-1.5916e-12 | 46.7097 | -7.276e-12 | 46.7987
-1.3642e-12 | 47.1157 | -9.7771e-12 | 47.2027
-4.3201e-12 | 47.5197 | -9.0949e-12 | 47.6067
-1.5916e-12 | 47.9247 | -8.6402e-12 | 48.0108
-1.819e-12 | 48.3298 | -8.4128e-12 | 48.4158
-2.5011e-12 | 48.7338 | -1.2051e-11 | 48.8218
-6.8212e-13 | 49.1388 | -9.7771e-12 | 49.2268
0 | 49.5448 | -9.7771e-12 | 49.6318
-1.819e-12 | 49.9469 | -8.4128e-12 | 50.0369
-3.1832e-12 | 50.3519 | -1.0459e-11 | 50.4429
-2.2737e-13 | 50.7569 | -1.0914e-11 | 50.8489
-2.2737e-13 | 51.1609 | -7.0486e-12 | 51.2549
-1.819e-12 | 51.5639 | -1.0914e-11 | 51.6589
-1.5916e-12 | 51.969 | -1.1141e-11 | 52.064
-2.5011e-12 | 52.373 | -1.2051e-11 | 52.47
-3.1832e-12 | 52.778 | -8.4128e-12 | 52.876
-3.1832e-12 | 53.184 | -6.5938e-12 | 53.2811
0 | 53.5891 | -8.6402e-12 | 53.6861
-5.457e-12 | 53.9911 | -1.2051e-11 | 54.0911
-4.5475e-13 | 54.3961 | -8.4128e-12 | 54.4961
-1.1369e-12 | 54.8001 | -1.1369e-11 | 54.9011
-2.7285e-12 | 55.2052 | -1.1141e-11 | 55.3052
-2.0464e-12 | 55.6092 | -1.2278e-11 | 55.7102
0 | 56.0142 | -6.5938e-12 | 56.1152
-2.0464e-12 | 56.4182 | -8.8676e-12 | 56.5202
-4.5475e-13 | 56.8222 | -1.0914e-11 | 56.9253
2.2737e-13 | 57.2253 | -1.0914e-11 | 57.3313
-1.5916e-12 | 57.6303 | -8.8676e-12 | 57.7373
1.819e-12 | 58.0353 | -9.7771e-12 | 58.1443
-4.5475e-13 | 58.4403 | -1.296e-11 | 58.5493
-2.7285e-12 | 58.8464 | -1.0914e-11 | 58.9544
-1.819e-12 | 59.2514 | -1.1596e-11 | 59.3604
-1.819e-12 | 59.6574 | -9.0949e-12 | 59.7664
2.2737e-13 | 60.0614 | -1.0914e-11 | 60.1714
-2.5011e-12 | 60.4655 | -9.0949e-12 | 60.5765
2.0464e-12 | 60.8695 | -1.2051e-11 | 60.9815
-1.5916e-12 | 61.2745 | -1.1596e-11 | 61.3885
1.1369e-12 | 61.6785 | -1.0004e-11 | 61.7945
-2.9559e-12 | 62.0836 | -6.5938e-12 | 62.2006
-2.2737e-13 | 62.4886 | -7.7307e-12 | 62.6066
1.3642e-12 | 62.8936 | -9.7771e-12 | 63.0116
-3.638e-12 | 63.2986 | -9.5497e-12 | 63.4166
-2.9559e-12 | 63.7036 | -1.0914e-11 | 63.8207
-2.5011e-12 | 64.1087 | -1.0232e-11 | 64.2267
-1.5916e-12 | 64.5137 | -1.1596e-11 | 64.6327
-1.5916e-12 | 64.9187 | -7.276e-12 | 65.0387
-6.8212e-13 | 65.3237 | -1.0232e-11 | 65.4437
-2.0464e-12 | 65.7268 | -1.0914e-11 | 65.8488
6.8212e-13 | 66.1308 | -1.296e-11 | 66.2548
2.9559e-12 | 66.5348 | -1.1141e-11 | 66.6608
-1.3642e-12 | 66.9408 | -9.7771e-12 | 67.0678
-1.1369e-12 | 67.3448 | -8.8676e-12 | 67.4739
9.0949e-13 | 67.7509 | -1.0459e-11 | 67.8789
-1.819e-12 | 68.1549 | -1.5461e-11 | 68.2839
-3.1832e-12 | 68.5609 | -7.276e-12 | 68.6889
-2.9559e-12 | 68.9669 | -1.1596e-11 | 69.094
2.0464e-12 | 69.372 | -1.4779e-11 | 69.501
-2.2737e-13 | 69.778 | -1.1596e-11 | 69.906
-3.1832e-12 | 70.18 | -8.8676e-12 | 70.311
1.819e-12 | 70.586 | -9.0949e-12 | 70.716
-1.3642e-12 | 70.9901 | -9.7771e-12 | 71.1211
2.2737e-13 | 71.3951 | -8.1855e-12 | 71.5261
-4.0927e-12 | 71.8001 | -8.8676e-12 | 71.9311
-2.5011e-12 | 72.2051 | -1.1369e-11 | 72.3351
-2.7285e-12 | 72.6102 | -1.0232e-11 | 72.7422
-2.7285e-12 | 73.0162 | -1.0914e-11 | 73.1482
6.8212e-13 | 73.4202 | -7.7307e-12 | 73.5532
-1.3642e-12 | 73.8242 | -7.0486e-12 | 73.9592
2.2737e-13 | 74.2292 | -8.8676e-12 | 74.3642
-2.7285e-12 | 74.6343 | -9.5497e-12 | 74.7703
-1.1369e-12 | 75.0403 | -7.5033e-12 | 75.1753
-2.2737e-13 | 75.4453 | -1.0232e-11 | 75.5813
6.8212e-13 | 75.8493 | -8.8676e-12 | 75.9853
-2.2737e-13 | 76.2534 | -1.2278e-11 | 76.3904
2.2737e-13 | 76.6584 | -1.2278e-11 | 76.7964
-2.2737e-13 | 77.0624 | -7.276e-12 | 77.2024
0 | 77.4664 | -1.0232e-11 | 77.6064
0 | 77.8695 | -1.2733e-11 | 78.0115
-1.5916e-12 | 78.2715 | -1.0459e-11 | 78.4175
-6.8212e-13 | 78.6765 | -1.0232e-11 | 78.8225
-3.1832e-12 | 79.0795 | -9.5497e-12 | 79.2275
2.0464e-12 | 79.4845 | -1.0004e-11 | 79.6335
6.8212e-13 | 79.8896 | -8.8676e-12 | 80.0386
-2.5011e-12 | 80.2926 | -1.0459e-11 | 80.4446
-3.8654e-12 | 80.7016 | -9.5497e-12 | 80.8496
0 | 81.1076 | -1.0914e-11 | 81.2557
-2.2737e-13 | 81.5117 | -1.1369e-11 | 81.6607
-4.5475e-13 | 81.9407 | -1.3415e-11 | 82.0887
-2.0464e-12 | 82.3477 | -6.5938e-12 | 82.4967
-1.3642e-12 | 82.7527 | -1.296e-11 | 82.9027
-1.1369e-12 | 83.1598 | -1.0459e-11 | 83.3088
-3.638e-12 | 83.5658 | -1.1596e-11 | 83.7128
-4.5475e-13 | 83.9708 | -1.2051e-11 | 84.1188
-1.3642e-12 | 84.3748
-1.5916e-12 | 84.7789

## 580uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.2836e-12 | 1.7336e-12 | -1.3419e-11 | 2.0412e-12
GEMs On |  | GEMs On
1.3642e-12 | 0.31202 | -1.296e-11 | 0.31002
-4.5475e-13 | 0.99606 | -1.4325e-11 | 0.99606
-2.5011e-12 | 1.4021 | -1.4779e-11 | 1.4001
2.2737e-12 | 1.8101 | -1.0459e-11 | 1.8071
2.7285e-12 | 2.2161 | -1.2051e-11 | 2.2131
0 | 2.6212 | -1.3415e-11 | 2.6181
2.9559e-12 | 3.0272 | -1.4552e-11 | 3.0232
-1.819e-12 | 3.4322 | -1.0914e-11 | 3.4292
-2.0464e-12 | 3.8382 | -1.1596e-11 | 3.8352
2.0464e-12 | 4.2442 | -1.2278e-11 | 4.2392
4.5475e-13 | 4.6533 | -1.4097e-11 | 4.6463
-4.0927e-12 | 5.0573 | -1.4325e-11 | 5.0523
-2.7285e-12 | 5.4613 | -1.3188e-11 | 5.4593
-3.4106e-12 | 5.8683 | -1.4552e-11 | 5.8653
9.0949e-13 | 6.2744 | -1.2051e-11 | 6.2714
2.7285e-12 | 6.6774 | -1.2278e-11 | 6.6774
2.0464e-12 | 7.0834 | -1.5689e-11 | 7.0814
-9.0949e-13 | 7.4884 | -1.4097e-11 | 7.4874
1.3642e-12 | 7.8945 | -1.2733e-11 | 7.8925
-2.7285e-12 | 8.3015 | -1.5689e-11 | 8.2985
-2.0464e-12 | 8.7105 | -1.4325e-11 | 8.7035
-2.2737e-13 | 9.1165 | -1.3642e-11 | 9.1095
-1.819e-12 | 9.5215 | -1.5007e-11 | 9.5145
-2.2737e-13 | 9.9266 | -1.5916e-11 | 9.9206
-9.0949e-13 | 10.3326 | -1.4552e-11 | 10.3266
6.8212e-13 | 10.7446 | -1.5007e-11 | 10.7316
-3.4106e-12 | 11.1496 | -1.5461e-11 | 11.1376
-6.8212e-13 | 11.5617 | -1.3642e-11 | 11.5427
-1.3642e-12 | 11.9667 | -1.3415e-11 | 11.9487
2.9559e-12 | 12.3707 | -1.728e-11 | 12.3537
-4.0927e-12 | 12.7757 | -1.2733e-11 | 12.7587
-1.819e-12 | 13.1877 | -1.3415e-11 | 13.1647
6.8212e-13 | 13.5938 | -1.4779e-11 | 13.5698
-2.0464e-12 | 14.0018 | -1.2051e-11 | 13.9768
1.1369e-12 | 14.4078 | -1.5916e-11 | 14.3818
1.3642e-12 | 14.8218 | -1.0459e-11 | 14.7858
-2.5011e-12 | 15.2289 | -1.3415e-11 | 15.1919
-2.2737e-13 | 15.6399 | -1.3188e-11 | 15.5989
-6.8212e-13 | 16.0469 | -1.4097e-11 | 16.0039
-2.2737e-13 | 16.4509 | -1.3642e-11 | 16.4099
-4.5475e-13 | 16.855 | -1.0232e-11 | 16.815
-9.0949e-13 | 17.27 | -1.2733e-11 | 17.221
-3.8654e-12 | 17.674 | -1.296e-11 | 17.626
-2.0464e-12 | 18.08 | -1.3188e-11 | 18.031
0 | 18.4861 | -1.0687e-11 | 18.437
-2.9559e-12 | 18.8891 | -1.1596e-11 | 18.8431
6.8212e-13 | 19.2951 | -1.1596e-11 | 19.2481
-9.0949e-13 | 19.6991 | -1.2733e-11 | 19.6541
-2.5011e-12 | 20.1042 | -1.3642e-11 | 20.0602
-9.0949e-13 | 20.5102 | -1.5916e-11 | 20.4672
6.8212e-13 | 20.9142 | -1.2051e-11 | 20.8722
-2.0464e-12 | 21.3182 | -1.2278e-11 | 21.2812
-1.5916e-12 | 21.7242 | -1.3642e-11 | 21.6872
-1.819e-12 | 22.1293 | -1.4097e-11 | 22.0943
1.1369e-12 | 22.5343 | -1.2278e-11 | 22.4983
-3.4106e-12 | 22.9403 | -1.5461e-11 | 22.9043
0 | 23.3463 | -1.2278e-11 | 23.3103
4.5475e-13 | 23.7494 | -1.3188e-11 | 23.7164
-2.5011e-12 | 24.1534 | -1.6144e-11 | 24.1214
-3.1832e-12 | 24.5594 | -1.4779e-11 | 24.5264
6.8212e-13 | 24.9774 | -1.5461e-11 | 24.9324
-3.8654e-12 | 25.3834 | -1.3188e-11 | 25.3375
-3.1832e-12 | 25.7885 | -1.2278e-11 | 25.7425
-4.0927e-12 | 26.1945 | -1.296e-11 | 26.1495
-3.1832e-12 | 26.6125 | -8.6402e-12 | 26.5545
-2.7285e-12 | 27.0266 | -1.3642e-11 | 26.9615
-2.5011e-12 | 27.4306 | -1.4552e-11 | 27.3676
-2.2737e-13 | 27.8366 | -1.4325e-11 | 27.7736
-3.4106e-12 | 28.2406 | -1.296e-11 | 28.1796
-4.5475e-12 | 28.6466 | -1.2733e-11 | 28.5856
-1.819e-12 | 29.0527 | -1.0914e-11 | 28.9917
-3.8654e-12 | 29.4597 | -1.4097e-11 | 29.3967
-1.819e-12 | 29.8667 | -1.4325e-11 | 29.8017
-3.4106e-12 | 30.3087 | -1.4552e-11 | 30.2077
-2.2737e-13 | 30.7158 | -1.5916e-11 | 30.6137
-3.1832e-12 | 31.1198 | -1.3188e-11 | 31.0188
9.0949e-13 | 31.5258 | -1.6598e-11 | 31.4238
-9.0949e-13 | 31.9288 | -1.6826e-11 | 31.8298
0 | 32.3528 | -1.1596e-11 | 32.2358
-3.4106e-12 | 32.7719 | -1.3642e-11 | 32.6419
9.0949e-13 | 33.1929 | -1.5007e-11 | 33.0469
-4.5475e-13 | 33.6069 | -1.2051e-11 | 33.4539
6.8212e-13 | 34.0119 | -1.0687e-11 | 33.8579
-1.5916e-12 | 34.416 | -1.0687e-11 | 34.264
-2.2737e-13 | 34.821 | -1.6598e-11 | 34.67
-4.0927e-12 | 35.223 | -1.0914e-11 | 35.076
-1.819e-12 | 35.63 | -1.4779e-11 | 35.482
4.5475e-13 | 36.0341 | -1.1141e-11 | 35.888
-3.1832e-12 | 36.4391 | -8.4128e-12 | 36.2931
-2.0464e-12 | 36.8441 | -1.1596e-11 | 36.6981
1.819e-12 | 37.2491 | -1.1596e-11 | 37.1051
4.5475e-13 | 37.6542 | -1.1823e-11 | 37.5112
-3.4106e-12 | 38.0592 | -1.2733e-11 | 37.9152
-5.0022e-12 | 38.4642 | -1.4325e-11 | 38.3212
-4.5475e-12 | 38.8682 | -1.4552e-11 | 38.7252
-3.4106e-12 | 39.2732 | -1.296e-11 | 39.1312
-6.8212e-13 | 39.6793 | -1.4097e-11 | 39.5363
-4.3201e-12 | 40.0833 | -1.4097e-11 | 39.9413
-9.0949e-13 | 40.4893 | -1.4325e-11 | 40.3453
-2.0464e-12 | 40.8943 | -1.6144e-11 | 40.7513
4.5475e-13 | 41.2994 | -1.296e-11 | 41.1574
-4.0927e-12 | 41.7474 | -9.3223e-12 | 41.5624
-2.0464e-12 | 42.1644 | -1.5007e-11 | 41.9664
0 | 42.5704 | -7.7307e-12 | 42.3714
-3.4106e-12 | 42.9735 | -1.4779e-11 | 42.7764
-3.8654e-12 | 43.3785 | -1.7508e-11 | 43.1835
-2.0464e-12 | 43.7845 | -1.0914e-11 | 43.5885
0 | 44.1905 | -1.2278e-11 | 43.9935
-1.819e-12 | 44.6185 | -1.4097e-11 | 44.3995
-1.819e-12 | 45.0336 | -1.3188e-11 | 44.8036
-2.9559e-12 | 45.4406 | -1.4097e-11 | 45.2096
-1.5916e-12 | 45.8456 | -1.4097e-11 | 45.6156
-9.0949e-13 | 46.2507 | -1.4097e-11 | 46.0216
-2.5011e-12 | 46.6557 | -1.4325e-11 | 46.4277
4.5475e-13 | 47.0597 | -1.4779e-11 | 46.8317
-2.7285e-12 | 47.4637 | -1.0914e-11 | 47.2387
-2.9559e-12 | 47.8687 | -1.2733e-11 | 47.6437
-2.0464e-12 | 48.3208 | -1.1141e-11 | 48.0508
6.8212e-13 | 48.7388 | -1.9327e-11 | 48.4568
-1.3642e-12 | 49.1428 | -1.728e-11 | 48.8628
-2.2737e-13 | 49.5478 | -1.5916e-11 | 49.2688
-2.9559e-12 | 49.9759 | -1.4097e-11 | 49.6748
-2.7285e-12 | 50.4009 | -1.2278e-11 | 50.0789
0 | 50.8099 | -1.0914e-11 | 50.4849
-1.3642e-12 | 51.2119 | -1.1141e-11 | 50.8899
-4.5475e-13 | 51.618 | -1.5689e-11 | 51.2949
-3.4106e-12 | 52.024 | -1.3415e-11 | 51.701
-3.4106e-12 | 52.43 | -1.0914e-11 | 52.107
4.5475e-13 | 52.835 | -1.1596e-11 | 52.513
-9.0949e-13 | 53.24 | -1.5461e-11 | 52.919
1.819e-12 | 53.6451 | -1.6826e-11 | 53.324
-1.819e-12 | 54.0841 | -1.2278e-11 | 53.7281
-1.3642e-12 | 54.5231 | -1.2051e-11 | 54.1321
-1.5916e-12 | 54.9281 | -1.2733e-11 | 54.5381
-2.9559e-12 | 55.3332 | -1.1823e-11 | 54.9431
-1.3642e-12 | 55.7382 | -1.0687e-11 | 55.3482
-4.3201e-12 | 56.1432 | -1.3188e-11 | 55.7532
1.3642e-12 | 56.5822 | -1.4097e-11 | 56.1572
-2.7285e-12 | 57.0123 | -1.4779e-11 | 56.5632
-3.1832e-12 | 57.4173 | -1.0914e-11 | 56.9683
0 | 57.8223 | -1.2278e-11 | 57.3723
-2.5011e-12 | 58.2283 | -1.4097e-11 | 57.7783
4.5475e-13 | 58.6474 | -1.4552e-11 | 58.1833
-1.5916e-12 | 59.0574 | -1.1141e-11 | 58.5883
2.9559e-12 | 59.4694 | -1.1823e-11 | 58.9924
-2.2737e-13 | 59.8774 | -1.6144e-11 | 59.3994
-2.0464e-12 | 60.2884 | -1.2278e-11 | 59.8044
-2.5011e-12 | 60.6945 | -1.5007e-11 | 60.2094
-2.7285e-12 | 61.0985 | -1.7508e-11 | 60.6155
-2.0464e-12 | 61.5045 | -8.6402e-12 | 61.0205
-2.9559e-12 | 61.9095 | -1.3188e-11 | 61.4255
-2.5011e-12 | 62.3156 | -1.1596e-11 | 61.8295
-1.5916e-12 | 62.7206 | -1.4552e-11 | 62.2366
-2.2737e-13 | 63.1256 | -9.5497e-12 | 62.6426
1.1369e-12 | 63.5296 | -1.3415e-11 | 63.0466
-4.5475e-13 | 63.9347 | -1.2733e-11 | 63.4526
-4.5475e-12 | 64.3387 | -1.3415e-11 | 63.8586
-6.8212e-13 | 64.7547 | -9.7771e-12 | 64.2637
-2.9559e-12 | 65.1617 | -1.5689e-11 | 64.6687
6.8212e-13 | 65.5667 | -1.7053e-11 | 65.0747
-3.8654e-12 | 65.9718 | -1.2733e-11 | 65.4797
-2.5011e-12 | 66.3768 | -1.2051e-11 | 65.8858
1.3642e-12 | 66.7898 | -1.4779e-11 | 66.2918
-1.5916e-12 | 67.1988 | -1.5916e-11 | 66.6958
-3.1832e-12 | 67.6029 | -1.3642e-11 | 67.1028
-6.8212e-13 | 68.0079 | -1.4325e-11 | 67.5079
1.819e-12 | 68.4129 | -1.5916e-11 | 67.9139
-6.8212e-13 | 68.8169 | -1.728e-11 | 68.3199
6.8212e-13 | 69.222 | -1.3188e-11 | 68.7269
1.3642e-12 | 69.626 | -1.296e-11 | 69.1329
-2.0464e-12 | 70.03 | -1.4325e-11 | 69.538
-2.0464e-12 | 70.435 | -1.5689e-11 | 69.944
-6.8212e-13 | 70.8401 | -1.5007e-11 | 70.352
-2.9559e-12 | 71.2451 | -1.6144e-11 | 70.7561
9.0949e-13 | 71.6511 | -1.4779e-11 | 71.1611
-2.2737e-13 | 72.0541 | -1.296e-11 | 71.5661
-4.5475e-13 | 72.4571 | -1.3642e-11 | 71.9711
-1.3642e-12 | 72.8612 | -1.2733e-11 | 72.3781
-6.8212e-13 | 73.2662 | -1.3188e-11 | 72.7832
-4.5475e-13 | 73.6702 | -1.0459e-11 | 73.1892
-2.0464e-12 | 74.0752 | -1.4097e-11 | 73.5942
-2.0464e-12 | 74.4793 | -1.4097e-11 | 74.0012
-1.819e-12 | 74.8853 | -1.6598e-11 | 74.4073
-2.2737e-13 | 75.2913 | -1.6826e-11 | 74.8123
-1.3642e-12 | 75.6963 | -1.5689e-11 | 75.2153
-2.2737e-13 | 76.1013 | -1.5007e-11 | 75.6213
-9.0949e-13 | 76.5034 | -1.3642e-11 | 76.0273
-2.0464e-12 | 76.9074 | -1.6826e-11 | 76.4324
-2.0464e-12 | 77.3114 | -1.5461e-11 | 76.8374
-2.7285e-12 | 77.7174 | -1.6144e-11 | 77.2424
-6.8212e-13 | 78.1225 | -1.3188e-11 | 77.6484
-2.2737e-13 | 78.5265 | -1.3642e-11 | 78.0545
-1.3642e-12 | 78.9305 | -1.1141e-11 | 78.4605
-3.4106e-12 | 79.3365 | -1.5007e-11 | 78.8665
-1.3642e-12 | 79.7406 | -1.2051e-11 | 79.2725
-9.0949e-13 | 80.1466 | -1.2051e-11 | 79.6766
4.5475e-13 | 80.5496 | -1.1141e-11 | 80.0826
-1.3642e-12 | 80.9526 | -1.5007e-11 | 80.4876
-2.0464e-12 | 81.3586 | -1.2278e-11 | 80.8956
1.1369e-12 | 81.7637 | -1.1823e-11 | 81.3007
-6.8212e-13 | 82.1697 | -6.3665e-12 | 81.7057
-4.0927e-12 | 82.5967 | -1.0914e-11 | 82.1327
-4.5475e-13 | 83.0018 | -1.1141e-11 | 82.5377
-4.0927e-12 | 83.4068 | -1.4325e-11 | 82.9447
-1.819e-12 | 83.8118 | -1.4552e-11 | 83.3498
-4.5475e-13 | 84.2168 | -9.0949e-12 | 83.7558
-9.0949e-13 | 84.6208 | -1.1823e-11 | 84.1608
-6.8212e-13 | 85.0249
-3.8654e-12 | 85.4279
-2.5011e-12 | 85.8329
-6.8212e-13 | 86.2389
-2.0464e-12 | 86.644
4.5475e-13 | 87.048
-9.0949e-13 | 87.452
-2.0464e-12 | 87.855
-6.8212e-13 | 88.2591
-1.5916e-12 | 88.6631
9.0949e-13 | 89.0681
-1.819e-12 | 89.4711
-2.2737e-13 | 89.8771

## 590uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-2.5011e-12 | 1.2956e-26 | -1.844e-11 | 2.172e-12
GEMs On |  | GEMs On
-2.5011e-12 | 0.30902 | -1.9327e-11 | 0.31202
-2.5011e-12 | 0.30902 | -1.9782e-11 | 0.99706
-2.5011e-12 | 0.30902 | -2.1373e-11 | 1.4041
-2.5011e-12 | 0.30902 | -1.9554e-11 | 1.8101
-2.5011e-12 | 0.30902 | -1.7963e-11 | 2.2151
-2.5011e-12 | 0.30902 | -1.9554e-11 | 2.6202
-2.5011e-12 | 0.30902 | -1.9554e-11 | 3.0272
-2.5011e-12 | 0.30902 | -2.4784e-11 | 3.4322
-2.5011e-12 | 0.30902 | -1.8872e-11 | 3.8382
-2.5011e-12 | 0.30902 | -1.8872e-11 | 4.2432
-2.5011e-12 | 0.30902 | -1.5461e-11 | 4.6513
-2.5011e-12 | 0.30902 | -1.728e-11 | 5.0583
-2.5011e-12 | 0.30902 | -1.8872e-11 | 5.4633
-2.5011e-12 | 0.30902 | -1.9554e-11 | 5.8703
-2.5011e-12 | 0.30902 | -1.9327e-11 | 6.2764
-2.5011e-12 | 0.30902 | -1.728e-11 | 6.6824
-2.5011e-12 | 0.30902 | -1.728e-11 | 7.0874
-2.5011e-12 | 0.30902 | -1.728e-11 | 7.4924
-2.5011e-12 | 0.30902 | -1.6598e-11 | 7.8975
-2.5011e-12 | 0.30902 | -2.1828e-11 | 8.3025
-2.5011e-12 | 0.30902 | -1.8417e-11 | 8.7065
-2.5011e-12 | 0.30902 | -2.0464e-11 | 9.1125
-2.5011e-12 | 0.30902 | -1.9782e-11 | 9.5175
-2.5011e-12 | 0.30902 | -2.0236e-11 | 9.9236
-2.5011e-12 | 0.30902 | -1.9327e-11 | 10.3296
-2.5011e-12 | 0.30902 | -1.9327e-11 | 10.7356
-2.5011e-12 | 0.30902 | -2.1373e-11 | 11.1406
-2.5011e-12 | 0.30902 | -1.7053e-11 | 11.5467
-2.5011e-12 | 0.30902 | -2.1828e-11 | 11.9527
-2.5011e-12 | 0.30902 | -1.819e-11 | 12.3597
-2.5011e-12 | 0.30902 | -1.9327e-11 | 12.7657
-2.5011e-12 | 0.30902 | -1.8872e-11 | 13.1698
-2.5011e-12 | 0.30902 | -1.5916e-11 | 13.5758
-2.5011e-12 | 0.30902 | -1.819e-11 | 13.9798
-2.5011e-12 | 0.30902 | -1.9782e-11 | 14.3858
-2.5011e-12 | 0.30902 | -1.6826e-11 | 14.7919
-2.5011e-12 | 0.30902 | -1.8417e-11 | 15.1969
-2.5011e-12 | 0.30902 | -1.6826e-11 | 15.6029
-2.5011e-12 | 0.30902 | -1.9327e-11 | 16.0089
-2.5011e-12 | 0.30902 | -1.5234e-11 | 16.4159
-2.5011e-12 | 0.30902 | -1.7735e-11 | 16.822
-2.5011e-12 | 0.30902 | -1.7963e-11 | 17.227
-2.5011e-12 | 0.30902 | -2.16e-11 | 17.633
-2.5011e-12 | 0.30902 | -1.5007e-11 | 18.037
-2.5011e-12 | 0.30902 | -2.0464e-11 | 18.4441
-2.5011e-12 | 0.30902 | -1.8872e-11 | 18.8501
-2.5011e-12 | 0.30902 | -1.6371e-11 | 19.2541
-2.5011e-12 | 0.30902 | -2.16e-11 | 19.6601
-2.5011e-12 | 0.30902 | -2.0918e-11 | 20.0662
-2.5011e-12 | 0.30902 | -1.9327e-11 | 20.4702
-2.5011e-12 | 0.30902 | -1.9099e-11 | 20.8762
-2.5011e-12 | 0.30902 | -1.8872e-11 | 21.2812
-2.5011e-12 | 0.30902 | -1.9099e-11 | 21.6872
-2.5011e-12 | 0.30902 | -1.8417e-11 | 22.0913
-2.5011e-12 | 0.30902 | -1.3415e-11 | 22.4953
-2.5011e-12 | 0.30902 | -1.8417e-11 | 22.9003
-2.5011e-12 | 0.30902 | -1.9327e-11 | 23.3053
-2.5011e-12 | 0.30902 | -2.0464e-11 | 23.7104
-2.5011e-12 | 0.30902 | -1.8417e-11 | 24.1154
-2.5011e-12 | 0.30902 | -2.2283e-11 | 24.5234
-2.5011e-12 | 0.30902 | -1.6371e-11 | 24.9294
-2.5011e-12 | 0.30902 | -1.6371e-11 | 25.3345
-2.5011e-12 | 0.30902 | -2.2055e-11 | 25.7395
-2.5011e-12 | 0.30902 | -2.1828e-11 | 26.1445
-2.5011e-12 | 0.30902 | -1.8872e-11 | 26.5495
-2.5011e-12 | 0.30902 | -2.0918e-11 | 26.9545
-2.5011e-12 | 0.30902 | -1.9099e-11 | 27.3606
-2.5011e-12 | 0.30902 | -1.8417e-11 | 27.7656
-2.5011e-12 | 0.30902 | -1.5916e-11 | 28.1706
-2.5011e-12 | 0.30902 | -1.7963e-11 | 28.5786
-2.5011e-12 | 0.30902 | -2.1373e-11 | 28.9837
-2.5011e-12 | 0.30902 | -1.8417e-11 | 29.3887
-2.5011e-12 | 0.30902 | -2.2283e-11 | 29.7937
-2.5011e-12 | 0.30902 | -2.0236e-11 | 30.1987
-2.5011e-12 | 0.30902 | -1.6826e-11 | 30.6058
-2.5011e-12 | 0.30902 | -1.5461e-11 | 31.0108
-2.5011e-12 | 0.30902 | -1.7963e-11 | 31.4168
-2.5011e-12 | 0.30902 | -1.728e-11 | 31.8218
-2.5011e-12 | 0.30902 | -1.5916e-11 | 32.2278
-2.5011e-12 | 0.30902 | -2.3419e-11 | 32.6329
-2.5011e-12 | 0.30902 | -1.6826e-11 | 33.0379
-2.5011e-12 | 0.30902 | -1.7963e-11 | 33.4429
-2.5011e-12 | 0.30902 | -1.8417e-11 | 33.8469
-2.5011e-12 | 0.30902 | -1.8872e-11 | 34.253
-2.5011e-12 | 0.30902 | -1.6598e-11 | 34.659
-2.5011e-12 | 0.30902 | -1.9554e-11 | 35.066
-2.5011e-12 | 0.30902 | -1.5916e-11 | 35.471
-2.5011e-12 | 0.30902 | -1.7053e-11 | 35.877
-2.5011e-12 | 0.30902 | -1.728e-11 | 36.2831
-2.5011e-12 | 0.30902 | -1.6826e-11 | 36.6881
-2.5011e-12 | 0.30902 | -1.9327e-11 | 37.0951
-2.5011e-12 | 0.30902 | -1.2506e-11 | 37.5001
-2.5011e-12 | 0.30902 | -1.7053e-11 | 37.9062
-2.5011e-12 | 0.30902 | -1.7963e-11 | 38.3122
-2.5011e-12 | 0.30902 | -2.1828e-11 | 38.7192
-2.5011e-12 | 0.30902 | -1.387e-11 | 39.1262
-2.5011e-12 | 0.30902 | -1.728e-11 | 39.5323
-2.5011e-12 | 0.30902 | -2.0918e-11 | 39.9383
-2.5011e-12 | 0.30902 | -1.9327e-11 | 40.3443
-2.5011e-12 | 0.30902 | -1.7053e-11 | 40.7493
-2.5011e-12 | 0.30902 | -1.9554e-11 | 41.1544
-2.5011e-12 | 0.30902 | -1.8872e-11 | 41.5604
-2.5011e-12 | 0.30902 | -1.296e-11 | 41.9664
-2.5011e-12 | 0.30902 | -2.0691e-11 | 42.3724
-2.5011e-12 | 0.30902 | -2.0236e-11 | 42.7775
-2.5011e-12 | 0.30902 | -1.5916e-11 | 43.1825
-2.5011e-12 | 0.30902 | -1.6598e-11 | 43.5875
-2.5011e-12 | 0.30902 | -2.3647e-11 | 43.9925
-2.5011e-12 | 0.30902 | -1.5689e-11 | 44.3985
-2.5011e-12 | 0.30902 | -1.5916e-11 | 44.8056
-2.5011e-12 | 0.30902 | -1.6371e-11 | 45.2116
-2.5011e-12 | 0.30902 | -1.8872e-11 | 45.6166
-2.5011e-12 | 0.30902 | -1.7963e-11 | 46.0226
-2.5011e-12 | 0.30902 | -2.1373e-11 | 46.4287
-2.5011e-12 | 0.30902 | -1.8872e-11 | 46.8337
-2.5011e-12 | 0.30902 | -2.0236e-11 | 47.2397
-2.5011e-12 | 0.30902 | -1.9099e-11 | 47.6477
-2.5011e-12 | 0.30902 | -2.0236e-11 | 48.0528
-2.5011e-12 | 0.30902 | -1.7053e-11 | 48.4598
-2.5011e-12 | 0.30902 | -1.9554e-11 | 48.8658
-2.5011e-12 | 0.30902 | -2.16e-11 | 49.2718
-2.5011e-12 | 0.30902 | -1.8417e-11 | 49.6768
-2.5011e-12 | 0.30902 | -1.6598e-11 | 50.0819
-2.5011e-12 | 0.30902 | -2.2965e-11 | 50.4869
-2.5011e-12 | 0.30902 | -1.9099e-11 | 50.8929
-2.5011e-12 | 0.30902 | -1.4325e-11 | 51.2969
-2.5011e-12 | 0.30902 | -1.3415e-11 | 51.702
-2.5011e-12 | 0.30902 | -1.819e-11 | 52.107
-2.5011e-12 | 0.30902 | -1.7053e-11 | 52.513
-2.5011e-12 | 0.30902 | -1.5689e-11 | 52.92
-2.5011e-12 | 0.30902 | -1.728e-11 | 53.327
-2.5011e-12 | 0.30902 | -2.1373e-11 | 53.7321
-2.5011e-12 | 0.30902 | -1.5461e-11 | 54.1371
-2.5011e-12 | 0.30902 | -2.0236e-11 | 54.5441
-2.5011e-12 | 0.30902 | -1.8417e-11 | 54.9511
-2.5011e-12 | 0.30902 | -1.6826e-11 | 55.3572
-2.5011e-12 | 0.30902 | -1.819e-11 | 55.7622
-2.5011e-12 | 0.30902 | -2.1373e-11 | 56.1682
-2.5011e-12 | 0.30902 | -1.8872e-11 | 56.5732
-2.5011e-12 | 0.30902 | -1.5461e-11 | 56.9783
-2.5011e-12 | 0.30902 | -1.9782e-11 | 57.3833
-2.5011e-12 | 0.30902 | -1.7053e-11 | 57.7913
-2.5011e-12 | 0.30902 | -1.6826e-11 | 58.1973
-2.5011e-12 | 0.30902 | -2.0236e-11 | 58.6063
-2.5011e-12 | 0.30902 | -1.6371e-11 | 59.0124
-2.5011e-12 | 0.30902 | -1.819e-11 | 59.4174
-2.5011e-12 | 0.30902 | -1.6371e-11 | 59.8244
-2.5011e-12 | 0.30902 | -1.7053e-11 | 60.2285
-2.5011e-12 | 0.30902 | -1.6371e-11 | 60.6325
-2.5011e-12 | 0.30902 | -1.6371e-11 | 61.0395
-2.5011e-12 | 0.30902 | -2.0236e-11 | 61.4445
-2.5011e-12 | 0.30902 | -1.5916e-11 | 61.8505
-2.5011e-12 | 0.30902 | -1.9554e-11 | 62.2556
-2.5011e-12 | 0.30902 | -2.3419e-11 | 62.6606
-2.5011e-12 | 0.30902 | -1.728e-11 | 63.0656
-2.5011e-12 | 0.30902 | -1.7735e-11 | 63.4716
-2.5011e-12 | 0.30902 | -1.8417e-11 | 63.8787
-2.5011e-12 | 0.30902 | -1.7963e-11 | 64.2847
-2.5011e-12 | 0.30902 | -2.2055e-11 | 64.6897
-2.5011e-12 | 0.30902 | -1.5689e-11 | 65.0947
-2.5011e-12 | 0.30902 | -1.7963e-11 | 65.5007
-2.5011e-12 | 0.30902 | -1.9554e-11 | 65.9068
-2.5011e-12 | 0.30902 | -1.8417e-11 | 66.3108
-2.5011e-12 | 0.30902 | -2.16e-11 | 66.7168
-2.5011e-12 | 0.30902 | -1.7735e-11 | 67.1228
-2.5011e-12 | 0.30902 | -1.819e-11 | 67.5319
-2.5011e-12 | 0.30902 | -1.9099e-11 | 67.9369
-2.5011e-12 | 0.30902 | -1.9554e-11 | 68.3429
-2.5011e-12 | 0.30902 | -1.4325e-11 | 68.7469
-2.5011e-12 | 0.30902 | -2.0236e-11 | 69.153
-2.5011e-12 | 0.30902 | -1.5916e-11 | 69.558
-2.5011e-12 | 0.30902 | -1.6826e-11 | 69.962
-2.5011e-12 | 0.30902 | -2.0691e-11 | 70.368
-2.5011e-12 | 0.30902 | -1.7963e-11 | 70.773
-2.5011e-12 | 0.30902 | -1.819e-11 | 71.1791
-2.5011e-12 | 0.30902 | -2.16e-11 | 71.5851
-2.5011e-12 | 0.30902 | -2.0236e-11 | 71.9911
-2.5011e-12 | 0.30902 | -2.0918e-11 | 72.3971
-2.5011e-12 | 0.30902 | -2.0236e-11 | 72.8022
-2.5011e-12 | 0.30902 | -1.9554e-11 | 73.2072
-2.5011e-12 | 0.30902 | -1.4097e-11 | 73.6122
-2.5011e-12 | 0.30902 | -1.819e-11 | 74.0182
-2.5011e-12 | 0.30902 | -1.7735e-11 | 74.4243
-2.5011e-12 | 0.30902 | -1.7735e-11 | 74.8303
-2.5011e-12 | 0.30902 | -2.0464e-11 | 75.2353
-2.5011e-12 | 0.30902 | -1.5916e-11 | 75.6423
-2.5011e-12 | 0.30902 | -2.0236e-11 | 76.0473
-2.5011e-12 | 0.30902 | -1.819e-11 | 76.4524
-2.5011e-12 | 0.30902 | -2.16e-11 | 76.8574
-2.5011e-12 | 0.30902 | -1.4325e-11 | 77.2634
-2.5011e-12 | 0.30902 | -1.6826e-11 | 77.6684
-2.5011e-12 | 0.30902 | -1.9782e-11 | 78.0745
-2.5011e-12 | 0.30902 | -1.9099e-11 | 78.4815
-2.5011e-12 | 0.30902 | -2.1373e-11 | 78.8865
-2.5011e-12 | 0.30902 | -1.7963e-11 | 79.2925
-2.5011e-12 | 0.30902 | -1.6826e-11 | 79.6976
-2.5011e-12 | 0.30902 | -1.7735e-11 | 80.1046
-2.5011e-12 | 0.30902 | -1.5689e-11 | 80.5106
-2.5011e-12 | 0.30902 | -1.728e-11 | 80.9146
-2.5011e-12 | 0.30902 | -1.5461e-11 | 81.3196
-2.5011e-12 | 0.30902 | -1.728e-11 | 81.7247
-2.5011e-12 | 0.30902 | -1.819e-11 | 82.1537
-2.5011e-12 | 0.30902 | -1.5461e-11 | 82.5607
-2.5011e-12 | 0.30902 | -1.9554e-11 | 82.9677
-2.5011e-12 | 0.30902 | -2.0236e-11 | 83.3728
-2.5011e-12 | 0.30902 | -1.6371e-11 | 83.7778
|  | -2.16e-11 | 84.1838
|  | -1.7963e-11 | 84.5888
|  | -1.6826e-11 | 84.9949
|  | -1.7963e-11 | 85.3999

## 600uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.8698e-12 | 1.6547e-12 | -2.5157e-11 | 2.5908e-12
GEMs On |  | GEMs On
-1.1369e-12 | 0.47103 | -2.2283e-11 | 0.31102
0 | 1.1561 | -2.1373e-11 | 0.99706
-2.7285e-12 | 1.5581 | -2.4556e-11 | 1.4031
-5.0022e-12 | 1.9641 | -2.3192e-11 | 1.8081
-3.4106e-12 | 2.3691 | -3.0013e-11 | 2.2141
-2.0464e-12 | 2.7742 | -2.5693e-11 | 2.6202
-3.4106e-12 | 3.1792 | -2.5238e-11 | 3.0252
-1.3642e-12 | 3.5842 | -2.7285e-11 | 3.4302
-4.5475e-13 | 3.9892 | -2.5011e-11 | 3.8372
-3.4106e-12 | 4.3933 | -2.3419e-11 | 4.2412
-3.1832e-12 | 4.7993 | -2.251e-11 | 4.6463
4.5475e-13 | 5.2053 | -2.0918e-11 | 5.0533
-4.5475e-13 | 5.6103 | -2.0691e-11 | 5.4593
-4.7748e-12 | 6.0153 | -2.9559e-11 | 5.8663
-1.3642e-12 | 6.4194 | -2.7285e-11 | 6.2724
-4.3201e-12 | 6.8254 | -2.1828e-11 | 6.6784
-2.0464e-12 | 7.2304 | -2.2055e-11 | 7.0824
-2.0464e-12 | 7.6344 | -2.4784e-11 | 7.4884
2.2737e-13 | 8.0405 | -2.7285e-11 | 7.8945
4.5475e-13 | 8.4445 | -2.3419e-11 | 8.3005
-3.638e-12 | 8.8485 | -2.2283e-11 | 8.7065
-2.7285e-12 | 9.2535 | -2.6148e-11 | 9.1115
-4.7748e-12 | 9.6586 | -2.4556e-11 | 9.5175
4.5475e-13 | 10.0646 | -2.3419e-11 | 9.9236
-4.3201e-12 | 10.4686 | -2.4329e-11 | 10.3296
-2.0464e-12 | 10.8726 | -2.3192e-11 | 10.7346
-3.638e-12 | 11.2766 | -2.6148e-11 | 11.1396
-2.5011e-12 | 11.6827 | -3.1605e-11 | 11.5457
-4.5475e-13 | 12.0877 | -2.1828e-11 | 11.9517
-2.0464e-12 | 12.4927 | -3.0468e-11 | 12.3557
-2.2737e-12 | 12.8987 | -2.6603e-11 | 12.7607
-4.3201e-12 | 13.3038 | -2.8649e-11 | 13.1647
-3.638e-12 | 13.7108 | -1.9782e-11 | 13.5718
-3.1832e-12 | 14.1158 | -2.3192e-11 | 13.9778
-2.0464e-12 | 14.5218 | -2.8422e-11 | 14.3828
-1.1369e-12 | 14.9259 | -1.8417e-11 | 14.7879
-4.3201e-12 | 15.3309 | -2.5011e-11 | 15.1949
-5.2296e-12 | 15.7349 | -2.6148e-11 | 15.5999
-3.8654e-12 | 16.1399 | -2.4784e-11 | 16.0059
-3.1832e-12 | 16.545 | -2.9559e-11 | 16.4119
-3.1832e-12 | 16.949 | -2.8194e-11 | 16.817
-1.1369e-12 | 17.354 | -2.6375e-11 | 17.222
4.5475e-13 | 17.759 | -2.5011e-11 | 17.626
-2.2737e-12 | 18.163 | -2.8422e-11 | 18.032
-2.5011e-12 | 18.5671 | -2.8422e-11 | 18.4381
-2.5011e-12 | 18.9721 | -2.7057e-11 | 18.8451
-1.3642e-12 | 19.3781 | -2.3647e-11 | 19.2511
-3.1832e-12 | 19.7831 | -2.4329e-11 | 19.6581
-2.7285e-12 | 20.1882 | -2.5238e-11 | 20.0642
-4.5475e-12 | 20.5922 | -2.3419e-11 | 20.4702
0 | 20.9972 | -2.2283e-11 | 20.8742
4.5475e-13 | 21.4012 | -2.5921e-11 | 21.2812
0 | 21.8042 | -2.6375e-11 | 21.6872
-9.0949e-13 | 22.2063 | -2.2737e-11 | 22.0933
-3.4106e-12 | 22.6093 | -2.774e-11 | 22.4983
-1.3642e-12 | 23.0143 | -2.4556e-11 | 22.9053
-2.5011e-12 | 23.4193 | -2.8422e-11 | 23.3113
-3.638e-12 | 23.8244 | -2.7285e-11 | 23.7164
-2.0464e-12 | 24.2284 | -2.4329e-11 | 24.1224
-1.1369e-12 | 24.6324 | -2.7057e-11 | 24.5284
-2.2737e-12 | 25.0374 | -2.3192e-11 | 24.9334
-1.819e-12 | 25.4425 | -2.9104e-11 | 25.3394
-2.5011e-12 | 25.8475 | -2.2055e-11 | 25.7455
1.1369e-12 | 26.2535 | -2.7285e-11 | 26.1515
-1.819e-12 | 26.6565 | -2.3647e-11 | 26.5565
-4.5475e-13 | 27.0626 | -2.4329e-11 | 26.9635
-6.8212e-13 | 27.4676 | -2.3419e-11 | 27.3696
2.2737e-13 | 27.8706 | -2.7512e-11 | 27.7736
-3.638e-12 | 28.2756 | -2.4556e-11 | 28.1786
6.8212e-13 | 28.6796 | -2.0918e-11 | 28.5846
-1.3642e-12 | 29.0837 | -2.7285e-11 | 28.9907
-2.5011e-12 | 29.4897 | -2.6148e-11 | 29.3967
4.5475e-13 | 29.8957 | -2.2737e-11 | 29.8017
1.1369e-12 | 30.2997 | -2.5238e-11 | 30.2057
-3.8654e-12 | 30.7038 | -2.1828e-11 | 30.6127
-1.819e-12 | 31.1078 | -2.5011e-11 | 31.0188
-1.3642e-12 | 31.5118 | -2.6603e-11 | 31.4248
-9.0949e-13 | 31.9178 | -2.5011e-11 | 31.8318
-4.5475e-13 | 32.3218 | -2.5693e-11 | 32.2368
-1.819e-12 | 32.7259 | -2.1146e-11 | 32.6429
-2.0464e-12 | 33.1309 | -2.5693e-11 | 33.0489
-4.5475e-12 | 33.5369 | -2.1828e-11 | 33.4549
-2.5011e-12 | 33.9409 | -2.4329e-11 | 33.8589
-2.7285e-12 | 34.346 | -2.4556e-11 | 34.264
2.2737e-13 | 34.752 | -2.8876e-11 | 34.67
-4.3201e-12 | 35.156 | -2.8194e-11 | 35.076
6.8212e-13 | 35.561 | -2.3192e-11 | 35.481
-3.1832e-12 | 35.9651 | -2.4556e-11 | 35.886
6.8212e-13 | 36.3701 | -2.1373e-11 | 36.2921
4.5475e-13 | 36.7741 | -3.0923e-11 | 36.6971
-6.8212e-13 | 37.1781 | -2.5238e-11 | 37.1031
0 | 37.5832 | -3.0241e-11 | 37.5091
-2.5011e-12 | 37.9862 | -2.9104e-11 | 37.9152
-3.638e-12 | 38.3912 | -2.2055e-11 | 38.3202
-2.5011e-12 | 38.7962 | -2.4556e-11 | 38.7252
-3.4106e-12 | 39.2032 | -2.6603e-11 | 39.1302
1.3642e-12 | 39.6083 | -2.5011e-11 | 39.5353
-3.4106e-12 | 40.0133 | -2.3192e-11 | 39.9413
-9.0949e-13 | 40.4183 | -2.774e-11 | 40.3463
2.2737e-13 | 40.8223 | -2.9559e-11 | 40.7523
-1.1369e-12 | 41.2264 | -3.115e-11 | 41.1574
-2.0464e-12 | 41.6304 | -2.3874e-11 | 41.5634
-4.3201e-12 | 42.0354 | -2.6603e-11 | 41.9694
-1.819e-12 | 42.4404 | -2.3192e-11 | 42.3744
-2.0464e-12 | 42.8464 | -2.7512e-11 | 42.7805
-9.0949e-13 | 43.2515 | -2.8422e-11 | 43.1855
-3.1832e-12 | 43.6555 | -2.6148e-11 | 43.5905
-1.1369e-12 | 44.0605 | -2.2055e-11 | 43.9965
0 | 44.4665 | -2.1146e-11 | 44.4015
-4.5475e-13 | 44.8706 | -2.251e-11 | 44.8066
-2.5011e-12 | 45.2756 | -2.3419e-11 | 45.2136
-3.1832e-12 | 45.6816 | -3.1378e-11 | 45.6186
-5.2296e-12 | 46.0866 | -2.2737e-11 | 46.0236
-9.0949e-13 | 46.4917 | -2.2055e-11 | 46.4297
-5.6843e-12 | 46.8967 | -2.5921e-11 | 46.8337
-1.819e-12 | 47.3017 | -2.5011e-11 | 47.2397
1.3642e-12 | 47.7057 | -2.3419e-11 | 47.6447
-4.3201e-12 | 48.1108 | -2.5921e-11 | 48.0498
-1.819e-12 | 48.5148 | -2.5693e-11 | 48.4548
-2.5011e-12 | 48.9188 | -2.5011e-11 | 48.8608
-3.4106e-12 | 49.3238 | -2.4784e-11 | 49.2648
-4.5475e-12 | 49.7278 | -2.5011e-11 | 49.6688
-9.0949e-13 | 50.1329 | -2.3419e-11 | 50.0739
4.5475e-13 | 50.5369 | -2.2283e-11 | 50.4799
-4.7748e-12 | 50.9419 | -2.4556e-11 | 50.8859
-2.0464e-12 | 51.3479 | -2.2737e-11 | 51.2899
4.5475e-13 | 51.754 | -2.251e-11 | 51.696
4.5475e-13 | 52.158 | -1.9327e-11 | 52.102
4.5475e-13 | 52.561 | -2.3419e-11 | 52.508
-1.3642e-12 | 52.965 | -2.3874e-11 | 52.912
-2.0464e-12 | 53.37 | -2.3647e-11 | 53.3171
-1.1369e-12 | 53.7751 | -2.6375e-11 | 53.7231
-1.1369e-12 | 54.1801 | -2.7285e-11 | 54.1281
-9.0949e-13 | 54.5851 | -2.3419e-11 | 54.5351
-6.8212e-13 | 54.9901 | -3.0241e-11 | 54.9391
0 | 55.3942 | -2.4784e-11 | 55.3442
2.2737e-13 | 55.7972 | -2.8194e-11 | 55.7502
-1.3642e-12 | 56.2002 | -2.4556e-11 | 56.1572
-2.7285e-12 | 56.6042 | -2.3419e-11 | 56.5622
-2.0464e-12 | 57.0093 | -2.3874e-11 | 56.9673
-2.7285e-12 | 57.4143 | -3.0241e-11 | 57.3723
-3.638e-12 | 57.8183 | -2.5011e-11 | 57.7783
1.1369e-12 | 58.2223 | -2.774e-11 | 58.1833
2.2737e-13 | 58.6273 | -2.3192e-11 | 58.5883
-9.0949e-13 | 59.0314 | -2.6603e-11 | 58.9934
4.5475e-13 | 59.4374 | -2.251e-11 | 59.3994
0 | 59.8404 | -2.2737e-11 | 59.8054
2.2737e-13 | 60.2435 | -2.774e-11 | 60.2114
-2.7285e-12 | 60.6485 | -2.5011e-11 | 60.6175
-3.8654e-12 | 61.0535 | -2.5011e-11 | 61.0245
-4.5475e-13 | 61.4575 | -2.774e-11 | 61.4295
-9.0949e-13 | 61.8615 | -2.251e-11 | 61.8345
-1.819e-12 | 62.2656 | -2.2737e-11 | 62.2396
-9.0949e-13 | 62.6706 | -2.5921e-11 | 62.6466
6.8212e-13 | 63.0756 | -2.3419e-11 | 63.0516
-4.5475e-12 | 63.4796 | -2.6375e-11 | 63.4576
-2.5011e-12 | 63.8837 | -2.7057e-11 | 63.8636
1.1369e-12 | 64.2887 | -2.4784e-11 | 64.2687
-1.1369e-12 | 64.6937 | -2.8194e-11 | 64.6737
-2.5011e-12 | 65.0977 | -2.3192e-11 | 65.0797
-9.0949e-13 | 65.5028 | -2.7512e-11 | 65.4857
-2.7285e-12 | 65.9078 | -2.5011e-11 | 65.8908
-2.5011e-12 | 66.3128 | -2.6375e-11 | 66.2958
-2.5011e-12 | 66.7178 | -2.3419e-11 | 66.7028
-1.1369e-12 | 67.1228 | -2.5921e-11 | 67.1088
0 | 67.5269 | -2.9786e-11 | 67.5149
-4.3201e-12 | 67.9289 | -2.251e-11 | 67.9219
-6.8212e-13 | 68.3349 | -2.0918e-11 | 68.3259
-4.5475e-13 | 68.7389 | -2.8422e-11 | 68.7309
-9.0949e-13 | 69.1429 | -2.3874e-11 | 69.1369
-3.1832e-12 | 69.548 | -2.3419e-11 | 69.542
-1.1369e-12 | 69.952 | -2.6148e-11 | 69.948
-4.5475e-13 | 70.357 | -2.5238e-11 | 70.354
-6.5938e-12 | 70.761 | -2.5693e-11 | 70.7601
-2.5011e-12 | 71.1651 | -2.774e-11 | 71.1651
-4.3201e-12 | 71.5691 | -2.7512e-11 | 71.5711
-1.819e-12 | 71.9741 | -2.0691e-11 | 71.9761
-3.4106e-12 | 72.3791 | -2.4556e-11 | 72.3831
-2.5011e-12 | 72.7842 | -2.8876e-11 | 72.7872
-1.819e-12 | 73.1882 | -2.6603e-11 | 73.1932
-9.0949e-13 | 73.5932 | -2.5921e-11 | 73.5992
-2.5011e-12 | 73.9982 | -2.4784e-11 | 74.0052
-5.6843e-12 | 74.4033 | -2.5011e-11 | 74.4093
-6.8212e-13 | 74.8083 | -2.3192e-11 | 74.8153
-3.8654e-12 | 75.2123 | -2.1828e-11 | 75.2203
-2.0464e-12 | 75.6163 | -2.4784e-11 | 75.6243
-3.4106e-12 | 76.0203 | -2.6375e-11 | 76.0314
-1.3642e-12 | 76.4254 | -2.1828e-11 | 76.4374
-2.0464e-12 | 76.8304 | -2.2737e-11 | 76.8424
6.8212e-13 | 77.2364 | -2.6603e-11 | 77.2474
-1.819e-12 | 77.6404 | -2.5693e-11 | 77.6544
-1.1369e-12 | 78.0445 | -2.6148e-11 | 78.0605
4.5475e-13 | 78.4495 | -2.8422e-11 | 78.4665
-2.7285e-12 | 78.8535 | -2.4784e-11 | 78.8715
-3.638e-12 | 79.2585 | -2.5238e-11 | 79.2775
-2.0464e-12 | 79.6646 | -2.6148e-11 | 79.6816
-2.7285e-12 | 80.0696 | -2.7285e-11 | 80.0876
-9.0949e-13 | 80.4756 | -2.3647e-11 | 80.4936
-1.1369e-12 | 80.8806 | -2.6603e-11 | 80.8986
-3.638e-12 | 81.2857 | -2.9786e-11 | 81.3026
-1.819e-12 | 81.6897 | -2.6148e-11 | 81.7087
-2.0464e-12 | 82.1177 | -1.8645e-11 | 82.1357
-2.7285e-12 | 82.5237 | -2.7512e-11 | 82.5407
-2.5011e-12 | 82.9287 | -2.8422e-11 | 82.9467
-2.2737e-12 | 83.3338 | -2.3874e-11 | 83.3538
-2.5011e-12 | 83.7378 | -2.2737e-11 | 83.7588
-3.638e-12 | 84.1428
-2.2737e-12 | 84.5468
2.2737e-13 | 84.9519
-4.3201e-12 | 85.3559
-1.819e-12 | 85.7599
1.3642e-12 | 86.1649
0 | 86.57
-3.1832e-12 | 86.972
-1.1369e-12 | 87.378

## 610uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-2.4385e-12 | 1.7434e-12 | -3.4505e-11 | 2.6931e-12
GEMs On |  | GEMs On
-7.0486e-12 | 0.31402 | -3.7744e-11 | 0.31102
-4.5475e-12 | 0.99706 | -3.4561e-11 | 0.99506
-2.5011e-12 | 1.4021 | -3.2969e-11 | 1.4031
1.5916e-12 | 1.8061 | -3.5698e-11 | 1.8101
-3.638e-12 | 2.2101 | -3.5698e-11 | 2.2151
-2.0464e-12 | 2.6161 | -3.7062e-11 | 2.6202
-1.5916e-12 | 3.0202 | -3.547e-11 | 3.0252
-4.3201e-12 | 3.4252 | -3.4333e-11 | 3.4312
-2.0464e-12 | 3.8302 | -3.3424e-11 | 3.8362
-4.5475e-12 | 4.2362 | -3.7289e-11 | 4.2422
-9.0949e-13 | 4.6403 | -3.5698e-11 | 4.6483
-2.9559e-12 | 5.0443 | -3.547e-11 | 5.0543
-4.5475e-12 | 5.4493 | -3.6835e-11 | 5.4603
-6.8212e-13 | 5.8533 | -3.5698e-11 | 5.8663
-1.819e-12 | 6.2584 | -3.2514e-11 | 6.2724
-2.5011e-12 | 6.6634 | -3.7744e-11 | 6.6774
-1.5916e-12 | 7.0684 | -3.7744e-11 | 7.0814
-9.0949e-13 | 7.4734 | -2.8649e-11 | 7.4884
-4.5475e-13 | 7.8774 | -3.4333e-11 | 7.8945
-2.0464e-12 | 8.2815 | -3.5016e-11 | 8.2995
-2.5011e-12 | 8.6865 | -3.206e-11 | 8.7045
-4.5475e-12 | 9.0905 | -3.3651e-11 | 9.1105
-4.5475e-13 | 9.4955 | -3.7744e-11 | 9.5155
4.5475e-13 | 9.9006 | -3.8881e-11 | 9.9226
-6.8212e-13 | 10.3046 | -3.5925e-11 | 10.3266
-3.1832e-12 | 10.7086 | -3.547e-11 | 10.7326
-1.5916e-12 | 11.1146 | -3.4333e-11 | 11.1376
-9.0949e-13 | 11.5197 | -3.4561e-11 | 11.5437
-4.5475e-13 | 11.9337 | -3.6835e-11 | 11.9487
-4.5475e-12 | 12.3377 | -3.5698e-11 | 12.3537
-2.9559e-12 | 12.7427 | -3.6835e-11 | 12.7597
-2.9559e-12 | 13.1478 | -3.3879e-11 | 13.1657
-1.5916e-12 | 13.5538 | -3.638e-11 | 13.5708
-2.0464e-12 | 13.9768 | -2.7967e-11 | 13.9758
-1.1369e-12 | 14.3998 | -3.4333e-11 | 14.3818
-3.638e-12 | 14.8058 | -3.7289e-11 | 14.7858
-2.5011e-12 | 15.2089 | -3.3879e-11 | 15.1909
-2.2737e-13 | 15.6139 | -3.7744e-11 | 15.5969
-2.2737e-12 | 16.0189 | -3.5698e-11 | 16.0039
1.3642e-12 | 16.4239 | -3.7744e-11 | 16.4089
-4.7748e-12 | 16.829 | -3.7289e-11 | 16.816
-2.2737e-12 | 17.233 | -3.6152e-11 | 17.22
-4.0927e-12 | 17.638 | -3.2969e-11 | 17.626
-2.2737e-12 | 18.043 | -3.8881e-11 | 18.031
-2.2737e-12 | 18.4471 | -3.4561e-11 | 18.4351
-5.457e-12 | 18.8511 | -3.0923e-11 | 18.8401
-1.1369e-12 | 19.2571 | -3.5016e-11 | 19.2451
4.5475e-13 | 19.6621 | -3.6152e-11 | 19.6501
-4.0927e-12 | 20.0662 | -3.5698e-11 | 20.0552
-6.8212e-13 | 20.4722 | -3.2287e-11 | 20.4602
-4.5475e-13 | 20.8762 | -3.2287e-11 | 20.8662
-2.0464e-12 | 21.2802 | -3.3197e-11 | 21.2712
-5.9117e-12 | 21.6862 | -3.3879e-11 | 21.6762
1.819e-12 | 22.0913 | -2.7967e-11 | 22.0813
-2.0464e-12 | 22.4973 | -3.0695e-11 | 22.4873
-2.0464e-12 | 22.9013 | -3.7744e-11 | 22.8943
-2.0464e-12 | 23.3063 | -3.3424e-11 | 23.2993
-4.5475e-13 | 23.7134 | -3.2969e-11 | 23.7054
-3.638e-12 | 24.1164 | -3.5698e-11 | 24.1114
-5.457e-12 | 24.5224 | -3.9563e-11 | 24.5174
-1.819e-12 | 24.9274 | -3.4788e-11 | 24.9244
-2.5011e-12 | 25.3325 | -3.1832e-11 | 25.3304
-4.0927e-12 | 25.7375 | -2.9559e-11 | 25.7355
-3.1832e-12 | 26.1435 | -3.9563e-11 | 26.1405
-2.0464e-12 | 26.5485 | -3.6152e-11 | 26.5465
-1.819e-12 | 26.9535 | -3.0923e-11 | 26.9515
-3.1832e-12 | 27.3596 | -3.2969e-11 | 27.3576
-2.5011e-12 | 27.7646 | -3.3879e-11 | 27.7626
-4.5475e-12 | 28.1706 | -3.5925e-11 | 28.1696
-1.5916e-12 | 28.5766 | -4.07e-11 | 28.5746
-5.6843e-12 | 28.9817 | -3.4333e-11 | 28.9807
-2.2737e-12 | 29.3877 | -3.2287e-11 | 29.3857
-9.0949e-13 | 29.7927 | -2.9331e-11 | 29.7897
-3.638e-12 | 30.1967 | -3.6152e-11 | 30.1947
-3.4106e-12 | 30.6008 | -3.2514e-11 | 30.6008
-2.9559e-12 | 31.0068 | -3.3197e-11 | 31.0058
-2.2737e-12 | 31.4108 | -3.4788e-11 | 31.4128
-4.5475e-12 | 31.8148 | -3.115e-11 | 31.8178
-5.6843e-12 | 32.2198 | -3.0468e-11 | 32.2238
-4.5475e-12 | 32.6249 | -3.2969e-11 | 32.6289
-3.1832e-12 | 33.0309 | -3.4788e-11 | 33.0359
-2.5011e-12 | 33.4359 | -3.3197e-11 | 33.4409
-1.1369e-12 | 33.8399 | -2.7967e-11 | 33.8469
6.8212e-13 | 34.244 | -3.1832e-11 | 34.252
-1.5916e-12 | 34.649 | -3.4561e-11 | 34.658
-3.1832e-12 | 35.055 | -3.7289e-11 | 35.063
-4.7748e-12 | 35.46 | -3.6835e-11 | 35.47
9.0949e-13 | 35.864 | -3.3651e-11 | 35.8751
2.5011e-12 | 36.2681 | -3.3651e-11 | 36.2811
-3.4106e-12 | 36.6731 | -3.7744e-11 | 36.6871
-4.0927e-12 | 37.0771 | -3.115e-11 | 37.0921
-1.819e-12 | 37.4811 | -3.0468e-11 | 37.4971
-1.5916e-12 | 37.8872 | -3.3424e-11 | 37.9022
-5.457e-12 | 38.2922 | -3.3197e-11 | 38.3072
-2.0464e-12 | 38.6972 | -3.4333e-11 | 38.7112
2.2737e-13 | 39.1022 | -3.1378e-11 | 39.1152
-3.4106e-12 | 39.5053 | -3.4333e-11 | 39.5223
-3.4106e-12 | 39.9103 | -3.4788e-11 | 39.9273
-5.6843e-12 | 40.3143 | -3.4333e-11 | 40.3323
-2.5011e-12 | 40.7193 | -3.4788e-11 | 40.7393
-2.9559e-12 | 41.1253 | -3.2287e-11 | 41.1444
-1.5916e-12 | 41.5304 | -3.979e-11 | 41.5504
-6.8212e-13 | 41.9354 | -3.3651e-11 | 41.9554
-1.5916e-12 | 42.3394 | -3.0695e-11 | 42.3614
-2.5011e-12 | 42.7444 | -3.3651e-11 | 42.7664
-3.638e-12 | 43.1495 | -3.7289e-11 | 43.1725
1.819e-12 | 43.5545 | -3.547e-11 | 43.5775
-1.5916e-12 | 43.9605 | -3.3879e-11 | 43.9835
-2.2737e-12 | 44.3645 | -3.2514e-11 | 44.3915
-4.5475e-12 | 44.7696 | -3.3879e-11 | 44.7976
-4.3201e-12 | 45.1746 | -4.0018e-11 | 45.2026
-9.0949e-13 | 45.5816 | -3.8881e-11 | 45.6066
-3.638e-12 | 45.9856 | -3.7289e-11 | 46.0126
-2.0464e-12 | 46.3907 | -3.8199e-11 | 46.4186
-2.9559e-12 | 46.7957 | -3.3879e-11 | 46.8247
-4.3201e-12 | 47.2007 | -3.3424e-11 | 47.2297
-6.8212e-13 | 47.6057 | -3.6152e-11 | 47.6357
-1.5916e-12 | 48.0108 | -3.1378e-11 | 48.0408
-4.0927e-12 | 48.4158 | -3.5698e-11 | 48.4468
-4.5475e-12 | 48.8208 | -3.6152e-11 | 48.8528
-4.7748e-12 | 49.2258 | -3.4333e-11 | 49.2588
-5.457e-12 | 49.6318 | -3.4788e-11 | 49.6638
-1.5916e-12 | 50.0359 | -3.6152e-11 | 50.0709
-4.7748e-12 | 50.4409 | -3.3879e-11 | 50.4769
-2.5011e-12 | 50.8469 | -3.6152e-11 | 50.8829
-2.2737e-12 | 51.2519 | -3.2287e-11 | 51.2889
-6.8212e-13 | 51.6559 | -3.2514e-11 | 51.693
-4.3201e-12 | 52.059 | -3.115e-11 | 52.098
-5.0022e-12 | 52.464 | -4.1382e-11 | 52.504
-2.9559e-12 | 52.868 | -3.7289e-11 | 52.909
-5.457e-12 | 53.2741 | -3.115e-11 | 53.315
-2.2737e-12 | 53.6801 | -3.3879e-11 | 53.7201
-2.5011e-12 | 54.0861 | -3.4561e-11 | 54.1261
-2.0464e-12 | 54.4901 | -3.4788e-11 | 54.5311
-2.2737e-12 | 54.8941 | -3.4561e-11 | 54.9371
-3.638e-12 | 55.2992 | -3.3879e-11 | 55.3432
-4.3201e-12 | 55.7042 | -3.2969e-11 | 55.7482
-3.1832e-12 | 56.1092 | -3.0695e-11 | 56.1552
-1.1369e-12 | 56.5152 | -3.2969e-11 | 56.5602
-1.5916e-12 | 56.9203 | -3.4333e-11 | 56.9653
-6.8212e-13 | 57.3253 | -3.3424e-11 | 57.3723
-4.7748e-12 | 57.7293 | -3.6835e-11 | 57.7773
-1.1369e-12 | 58.1333 | -3.1832e-11 | 58.1823
-2.2737e-12 | 58.5374 | -3.7062e-11 | 58.5883
-2.5011e-12 | 58.9434 | -3.2969e-11 | 58.9944
-2.2737e-13 | 59.3484 | -3.3197e-11 | 59.3994
-6.8212e-13 | 59.7544 | -3.2287e-11 | 59.8054
-2.2737e-12 | 60.1584 | -2.9786e-11 | 60.2114
-7.276e-12 | 60.5635 | -3.3879e-11 | 60.6175
-2.5011e-12 | 60.9685 | -3.8881e-11 | 61.0215
9.0949e-13 | 61.3725 | -3.4788e-11 | 61.4265
-1.5916e-12 | 61.7775 | -3.7289e-11 | 61.8325
-2.2737e-12 | 62.1816 | -3.0695e-11 | 62.2376
-4.5475e-13 | 62.5866 | -3.7517e-11 | 62.6436
-3.1832e-12 | 62.9906 | -3.8426e-11 | 63.0506
-6.8212e-13 | 63.3956 | -3.3197e-11 | 63.4556
-2.0464e-12 | 63.7996 | -3.4561e-11 | 63.8606
-9.0949e-13 | 64.2037 | -3.206e-11 | 64.2667
-2.2737e-12 | 64.6077 | -3.6152e-11 | 64.6727
-5.457e-12 | 65.0127 | -3.1378e-11 | 65.0787
-4.3201e-12 | 65.4187 | -3.4561e-11 | 65.4838
-4.3201e-12 | 65.8238 | -3.4333e-11 | 65.8888
-1.1369e-12 | 66.2288 | -3.4788e-11 | 66.2948
-4.5475e-13 | 66.6338 | -2.7967e-11 | 66.6988
-1.5916e-12 | 67.0388 | -3.9336e-11 | 67.1038
-2.2737e-12 | 67.4439 | -3.6152e-11 | 67.5089
-3.638e-12 | 67.8479 | -3.2969e-11 | 67.9159
-6.1391e-12 | 68.2549 | -3.5016e-11 | 68.3209
-3.4106e-12 | 68.6589 | -3.3197e-11 | 68.7269
-2.2737e-12 | 69.063 | -3.3651e-11 | 69.132
1.819e-12 | 69.468 | -3.8426e-11 | 69.536
-1.819e-12 | 69.874 | -3.638e-11 | 69.942
4.5475e-13 | 70.278 | -3.2287e-11 | 70.349
-2.2737e-12 | 70.683 | -3.1832e-11 | 70.7571
-3.638e-12 | 71.0881 | -3.6835e-11 | 71.1621
-4.0927e-12 | 71.4941 | -3.0923e-11 | 71.5681
-2.0464e-12 | 71.8981 | -3.3879e-11 | 71.9731
-1.5916e-12 | 72.3031 | -3.8654e-11 | 72.3791
-3.4106e-12 | 72.7072 | -3.7062e-11 | 72.7832
-2.5011e-12 | 73.1122 | -3.2969e-11 | 73.1892
-3.1832e-12 | 73.5172 | -3.2514e-11 | 73.5942
-1.5916e-12 | 73.9222 | -3.3197e-11 | 74.0002
-4.5475e-13 | 74.3263 | -3.3879e-11 | 74.4063
-2.0464e-12 | 74.7293 | -3.0923e-11 | 74.8113
-2.5011e-12 | 75.1343 | -3.3651e-11 | 75.2163
-4.5475e-13 | 75.5403 | -3.206e-11 | 75.6213
-4.3201e-12 | 75.9443 | -3.0923e-11 | 76.0273
-1.819e-12 | 76.3494 | -3.206e-11 | 76.4324
-2.5011e-12 | 76.7534 | -3.206e-11 | 76.8394
-2.5011e-12 | 77.1594 | -4.2519e-11 | 77.2434
-1.1369e-12 | 77.5634 | -3.1832e-11 | 77.6494
-2.9559e-12 | 77.9685 | -3.5016e-11 | 78.0545
-4.3201e-12 | 78.3745 | -3.7289e-11 | 78.4595
-2.5011e-12 | 78.7795 | -2.7967e-11 | 78.8645
-1.1369e-12 | 79.1845 | -3.6835e-11 | 79.2705
-9.0949e-13 | 79.5905 | -3.0695e-11 | 79.6766
-2.5011e-12 | 79.9946 | -3.5925e-11 | 80.0816
-1.819e-12 | 80.4006 | -3.7744e-11 | 80.4866
-2.9559e-12 | 80.8056 | -3.6152e-11 | 80.8936
-3.1832e-12 | 81.2096 | -3.8881e-11 | 81.2996
-3.638e-12 | 81.6147 | -3.547e-11 | 81.7047
-2.2737e-12 | 82.0427 | -3.4561e-11 | 82.1337
-9.0949e-13 | 82.4487 | -3.3879e-11 | 82.5397
-1.819e-12 | 82.8537 | -3.4788e-11 | 82.9457
-1.819e-12 | 83.2598 | -3.5698e-11 | 83.3508
-5.0022e-12 | 83.6658 | -3.9563e-11 | 83.7558
9.0949e-13 | 84.0708 | -3.5925e-11 | 84.1608
|  | -3.4333e-11 | 84.5658

## 620uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-2.9044e-12 | 2.0197e-12 | -4.8062e-11 | 3.8576e-12
GEMs On |  | GEMs On
-5.0022e-12 | 0.31302 | -5.2069e-11 | 0.31602
-6.8212e-13 | 0.99806 | -4.9567e-11 | 1.0021
-5.457e-12 | 1.4041 | -4.3428e-11 | 1.4091
-4.3201e-12 | 1.8121 | -4.7976e-11 | 1.8151
-9.0949e-13 | 2.2181 | -4.0245e-11 | 2.2211
-1.1369e-12 | 2.6221 | -4.7976e-11 | 2.6271
-1.5916e-12 | 3.0372 | -3.979e-11 | 3.0332
-2.5011e-12 | 3.4432 | -5.2069e-11 | 3.4402
-6.3665e-12 | 3.8482 | -4.8431e-11 | 3.8462
-4.5475e-13 | 4.2542 | -4.8431e-11 | 4.2522
9.0949e-13 | 4.6593 | -4.7521e-11 | 4.6573
-2.9559e-12 | 5.0623 | -4.2291e-11 | 5.0633
-7.0486e-12 | 5.4683 | -4.9795e-11 | 5.4673
-9.0949e-13 | 5.8733 | -5.3433e-11 | 5.8723
-1.5916e-12 | 6.2774 | -5.1841e-11 | 6.2774
-3.4106e-12 | 6.6834 | -4.411e-11 | 6.6834
-2.0464e-12 | 7.0884 | -4.6157e-11 | 7.0894
0 | 7.4944 | -4.2974e-11 | 7.4944
-4.5475e-12 | 7.8965 | -5.4797e-11 | 7.8995
-1.5916e-12 | 8.3015 | -4.7976e-11 | 8.3055
-1.819e-12 | 8.7065 | -4.9567e-11 | 8.7115
-2.5011e-12 | 9.1125 | -5.1159e-11 | 9.1175
-3.1832e-12 | 9.5185 | -4.6384e-11 | 9.5225
-4.0927e-12 | 9.9246 | -5.457e-11 | 9.9326
-4.5475e-13 | 10.3316 | -5.1386e-11 | 10.3366
-5.0022e-12 | 10.7366 | -4.8203e-11 | 10.7426
-2.2737e-12 | 11.1426 | -5.5934e-11 | 11.1486
1.819e-12 | 11.5507 | -4.8658e-11 | 11.5667
-3.4106e-12 | 11.9557 | -5.0704e-11 | 11.9937
2.2737e-13 | 12.3607 | -4.7976e-11 | 12.3997
-2.9559e-12 | 12.7787 | -4.934e-11 | 12.8057
-3.1832e-12 | 13.1817 | -4.6157e-11 | 13.2118
-4.5475e-13 | 13.5878 | -5.2069e-11 | 13.6168
-3.4106e-12 | 13.9988 | -4.0245e-11 | 14.0228
-2.9559e-12 | 14.4038 | -4.8885e-11 | 14.4278
-1.819e-12 | 14.8109 | -4.1155e-11 | 14.8338
-5.9117e-12 | 15.2219 | -4.7066e-11 | 15.2389
-2.5011e-12 | 15.6299 | -4.6157e-11 | 15.6449
-9.0949e-13 | 16.0349 | -4.502e-11 | 16.0509
-3.638e-12 | 16.4489 | -5.0932e-11 | 16.4599
-4.7748e-12 | 16.856 | -4.7066e-11 | 16.865
-5.6843e-12 | 17.263 | -4.6839e-11 | 17.272
-2.2737e-12 | 17.669 | -4.5702e-11 | 17.677
-2.2737e-12 | 18.075 | -5.0932e-11 | 18.082
2.2737e-13 | 18.4791 | -4.9795e-11 | 18.4881
-2.2737e-12 | 18.8911 | -5.1841e-11 | 18.9121
-4.5475e-13 | 19.3011 | -5.3433e-11 | 19.3201
-6.8212e-13 | 19.7091 | -5.3888e-11 | 19.7291
-3.4106e-12 | 20.1132 | -4.9567e-11 | 20.1491
-6.3665e-12 | 20.5182 | -5.1841e-11 | 20.5572
-2.5011e-12 | 20.9242 | -5.0704e-11 | 20.9662
-1.819e-12 | 21.3312 | -4.5929e-11 | 21.3722
-4.5475e-12 | 21.7372 | -4.5929e-11 | 21.7972
-2.0464e-12 | 22.1423 | -5.0932e-11 | 22.2053
-4.0927e-12 | 22.5483 | -4.3656e-11 | 22.6103
-5.6843e-12 | 22.9553 | -4.4565e-11 | 23.0253
-4.5475e-13 | 23.3613 | -4.502e-11 | 23.4323
-1.1369e-12 | 23.7674 | -4.5475e-11 | 23.8364
-4.3201e-12 | 24.1724 | -4.9567e-11 | 24.2764
-3.4106e-12 | 24.5834 | -5.2296e-11 | 24.6814
-2.9559e-12 | 24.9894 | -4.8658e-11 | 25.0884
-5.0022e-12 | 25.3955 | -5.025e-11 | 25.4985
-3.4106e-12 | 25.7995 | -4.3656e-11 | 25.9045
-4.3201e-12 | 26.2085 | -4.8431e-11 | 26.3115
1.3642e-12 | 26.6145 | -4.0927e-11 | 26.7185
-4.0927e-12 | 27.0246 | -4.9795e-11 | 27.1295
-6.8212e-12 | 27.4306 | -4.9567e-11 | 27.5366
-2.9559e-12 | 27.8376 | -4.2064e-11 | 27.9426
-2.9559e-12 | 28.2606 | -4.4793e-11 | 28.3486
-4.5475e-12 | 28.6716 | -4.4565e-11 | 28.7546
-4.0927e-12 | 29.0767 | -4.1609e-11 | 29.1877
-6.8212e-12 | 29.4807 | -5.5252e-11 | 29.5957
-2.0464e-12 | 29.8877 | -4.7976e-11 | 30.0017
-1.1369e-12 | 30.2967 | -4.7976e-11 | 30.4097
-4.0927e-12 | 30.7048 | -4.1609e-11 | 30.8148
-2.5011e-12 | 31.1098 | -5.2296e-11 | 31.2338
-1.819e-12 | 31.5178 | -4.411e-11 | 31.6398
-4.5475e-12 | 31.9218 | -4.6384e-11 | 32.0458
-5.6843e-12 | 32.3278 | -4.1155e-11 | 32.4509
-3.638e-12 | 32.7329 | -5.3433e-11 | 32.8569
-6.8212e-13 | 33.1579 | -4.7976e-11 | 33.2619
-3.638e-12 | 33.5719 | -4.8885e-11 | 33.6689
-2.2737e-12 | 33.9769 | -4.8885e-11 | 34.0849
-5.9117e-12 | 34.381 | -4.7066e-11 | 34.491
-4.0927e-12 | 34.785 | -4.9795e-11 | 34.896
-4.3201e-12 | 35.19 | -5.5934e-11 | 35.301
-2.5011e-12 | 35.617 | -4.8431e-11 | 35.708
2.2737e-13 | 36.0231 | -4.4338e-11 | 36.1211
-1.819e-12 | 36.4341 | -5.025e-11 | 36.5271
-4.5475e-13 | 36.8391 | -4.8431e-11 | 36.9311
-2.0464e-12 | 37.2451 | -4.9795e-11 | 37.3461
-4.5475e-12 | 37.6742 | -4.9795e-11 | 37.7522
-3.638e-12 | 38.0782 | -5.3888e-11 | 38.1582
0 | 38.4832 | -4.7294e-11 | 38.5642
-3.4106e-12 | 38.8862 | -5.025e-11 | 38.9702
-2.0464e-12 | 39.2923 | -5.5024e-11 | 39.3772
-1.5916e-12 | 39.6973 | -4.7066e-11 | 39.7853
-2.2737e-12 | 40.1023 | -4.7976e-11 | 40.1893
-2.2737e-12 | 40.5083 | -4.7066e-11 | 40.5963
-4.0927e-12 | 40.9143 | -4.9795e-11 | 41.0014
-3.4106e-12 | 41.3194 | -4.9567e-11 | 41.4084
-3.4106e-12 | 41.7234 | -4.502e-11 | 41.8144
-1.5916e-12 | 42.1294 | -5.025e-11 | 42.2414
-2.9559e-12 | 42.5344 | -3.9563e-11 | 42.6494
2.2737e-13 | 42.9605 | -5.0932e-11 | 43.0675
-3.638e-12 | 43.3665 | -5.0932e-11 | 43.4745
-4.5475e-12 | 43.7715 | -4.7294e-11 | 43.8815
-3.4106e-12 | 44.1765 | -4.2974e-11 | 44.3035
-6.1391e-12 | 44.5906 | -4.7294e-11 | 44.7166
-1.819e-12 | 44.9976 | -4.9567e-11 | 45.1236
-1.5916e-12 | 45.4036 | -5.2296e-11 | 45.5296
-7.0486e-12 | 45.8106 | -5.2523e-11 | 45.9356
-1.1369e-12 | 46.2166 | -4.9567e-11 | 46.3417
-7.5033e-12 | 46.6227 | -4.3201e-11 | 46.7547
-3.638e-12 | 47.0357 | -4.8203e-11 | 47.1607
-3.4106e-12 | 47.5027 | -4.4338e-11 | 47.5697
-5.457e-12 | 47.9087 | -4.9567e-11 | 47.9757
-4.0927e-12 | 48.3138 | -5.2296e-11 | 48.3828
-9.0949e-13 | 48.7188 | -4.0245e-11 | 48.8148
-2.2737e-12 | 49.1228 | -4.7521e-11 | 49.2198
-3.4106e-12 | 49.5298 | -4.4565e-11 | 49.6248
-5.457e-12 | 49.9349 | -5.7071e-11 | 50.0309
-4.0927e-12 | 50.3399 | -5.1386e-11 | 50.4369
-4.5475e-12 | 50.7469 | -4.8203e-11 | 50.8429
0 | 51.1529 | -5.5707e-11 | 51.2499
-2.5011e-12 | 51.5549 | -5.0704e-11 | 51.6579
-6.8212e-12 | 51.96 | -4.7294e-11 | 52.064
-3.1832e-12 | 52.367 | -4.8885e-11 | 52.469
-4.0927e-12 | 52.785 | -5.1841e-11 | 52.875
-1.819e-12 | 53.19 | -4.6839e-11 | 53.2801
-4.0927e-12 | 53.5971 | -4.9567e-11 | 53.6911
-5.9117e-12 | 54.0021 | -5.1159e-11 | 54.0961
-6.8212e-13 | 54.4071 | -4.6384e-11 | 54.5031
-4.0927e-12 | 54.8151 | -5.8662e-11 | 54.9201
2.2737e-13 | 55.2302 | -4.7976e-11 | 55.3282
-3.1832e-12 | 55.6342 | -4.8658e-11 | 55.7342
-7.276e-12 | 56.0432 | -5.1159e-11 | 56.1392
-9.0949e-13 | 56.4482 | -4.3428e-11 | 56.5442
-3.4106e-12 | 56.8533 | -5.0704e-11 | 56.9513
-4.5475e-12 | 57.2583 | -4.7294e-11 | 57.3573
-4.0927e-12 | 57.6993 | -5.3433e-11 | 57.7633
-9.0949e-13 | 58.1043 | -4.3656e-11 | 58.1693
-2.5011e-12 | 58.5104 | -4.411e-11 | 58.5763
-5.0022e-12 | 58.9144 | -4.411e-11 | 58.9824
-2.9559e-12 | 59.3204 | -4.8885e-11 | 59.3884
-2.9559e-12 | 59.7264 | -5.025e-11 | 59.7944
-1.1369e-12 | 60.1314 | -4.2064e-11 | 60.2044
-4.0927e-12 | 60.5375 | -5.2751e-11 | 60.6115
-1.1369e-12 | 60.9445 | -4.1155e-11 | 61.0175
-1.819e-12 | 61.3505 | -4.934e-11 | 61.4225
-2.5011e-12 | 61.7585 | -4.9567e-11 | 61.8285
9.0949e-13 | 62.1636 | -4.3428e-11 | 62.2356
-7.5033e-12 | 62.5696 | -4.0245e-11 | 62.6416
-2.9559e-12 | 62.9746 | -4.6384e-11 | 63.0476
-3.4106e-12 | 63.3796 | -5.2069e-11 | 63.4526
-1.5916e-12 | 63.8006 | -4.6839e-11 | 63.8586
-3.1832e-12 | 64.2077 | -4.3428e-11 | 64.2647
-1.1369e-12 | 64.6237 | -5.2296e-11 | 64.6717
-6.8212e-13 | 65.0367 | -4.8885e-11 | 65.0787
-4.5475e-12 | 65.4417 | -4.7294e-11 | 65.4958
-3.1832e-12 | 65.8478 | -4.3656e-11 | 65.9018
-2.9559e-12 | 66.2598 | -4.8431e-11 | 66.3068
2.2737e-13 | 66.6658 | -4.9795e-11 | 66.7138
-1.1369e-12 | 67.0688 | -4.9795e-11 | 67.1228
-2.0464e-12 | 67.4739 | -4.6839e-11 | 67.5289
0 | 67.8789 | -4.3656e-11 | 67.9379
-4.5475e-12 | 68.2859 | -5.0704e-11 | 68.3439
-3.1832e-12 | 68.6919 | -4.411e-11 | 68.7519
6.8212e-13 | 69.096 | -4.7976e-11 | 69.162
-2.2737e-12 | 69.502 | -5.6161e-11 | 69.589
-3.638e-12 | 69.915 | -4.4793e-11 | 69.995
-2.2737e-12 | 70.328 | -5.0022e-11 | 70.401
-1.1369e-12 | 70.7331 | -5.0932e-11 | 70.809
-2.5011e-12 | 71.1381 | -5.0022e-11 | 71.2141
-5.9117e-12 | 71.5461 | -4.1837e-11 | 71.6241
2.2737e-13 | 71.9541 | -4.5702e-11 | 72.0321
-5.6843e-12 | 72.3641 | -4.6157e-11 | 72.4391
-1.1369e-12 | 72.7732 | -4.7976e-11 | 72.8482
-3.4106e-12 | 73.1812 | -4.7976e-11 | 73.2552
-9.0949e-13 | 73.5902 | -4.411e-11 | 73.6612
-1.819e-12 | 73.9962 | -5.2751e-11 | 74.0652
-5.9117e-12 | 74.4013 | -4.4793e-11 | 74.4713
-3.4106e-12 | 74.8083 | -4.3201e-11 | 74.8773
-3.1832e-12 | 75.2133 | -4.5929e-11 | 75.2833
-4.0927e-12 | 75.6193 | -5.3205e-11 | 75.7023
-4.7748e-12 | 76.0243 | -5.2523e-11 | 76.1094
-4.3201e-12 | 76.4304 | -5.366e-11 | 76.5254
-6.8212e-12 | 76.8354 | -5.0022e-11 | 76.9304
-4.5475e-13 | 77.2464 | -5.3205e-11 | 77.3354
-7.9581e-12 | 77.6624 | -4.6157e-11 | 77.7445
1.5916e-12 | 78.0675 | -4.9795e-11 | 78.1615
-5.6843e-12 | 78.4765 | -5.2523e-11 | 78.5665
-5.6843e-12 | 78.8815 | -5.0022e-11 | 78.9725
-3.1832e-12 | 79.2875 | -5.1386e-11 | 79.3785
-1.5916e-12 | 79.6936 | -5.1159e-11 | 79.7846
-2.5011e-12 | 80.1006 | -5.2751e-11 | 80.1926
0 | 80.5056 | -4.934e-11 | 80.6006
-2.5011e-12 | 80.9086 | -4.6157e-11 | 81.0056
-4.5475e-13 | 81.3196 | -4.7521e-11 | 81.4117
-4.5475e-12 | 81.7487 | -4.3201e-11 | 81.8177
-3.4106e-12 | 82.1537 | -5.1841e-11 | 82.2227
-5.457e-12 | 82.5807 | -4.5475e-11 | 82.6517
9.0949e-13 | 83.0507 | -4.411e-11 | 83.0567
-6.1391e-12 | 83.4748 | -3.8654e-11 | 83.4618
-3.1832e-12 | 83.8808 | -5.2069e-11 | 83.9008
4.5475e-13 | 84.2848 | -3.7289e-11 | 84.3078
0 | 84.6898 | -5.0932e-11 | 84.7138
-2.9559e-12 | 85.1029 | -4.7294e-11 | 85.1199
-4.7748e-12 | 85.5339 | -4.5475e-11 | 85.5269
-1.819e-12 | 85.9609 | -5.366e-11 | 85.9329
6.8212e-13 | 86.3659 | -5.025e-11 | 86.3419
-1.819e-12 | 86.772 | -4.6384e-11 | 86.748
-3.1832e-12 | 87.177 | -4.7066e-11 | 87.161
-2.9559e-12 | 87.584 | -4.934e-11 | 87.58
-2.9559e-12 | 87.99 | -5.0022e-11 | 87.988
-6.1391e-12 | 88.3951 | -4.8658e-11 | 88.3951
-4.7748e-12 | 88.7991 | -5.3888e-11 | 88.8011
-6.8212e-13 | 89.2041 | -4.7976e-11 | 89.2091
-2.5011e-12 | 89.6101 | -4.8658e-11 | 89.6181
-5.0022e-12 | 90.0152 | -4.3428e-11 | 90.0262
-1.819e-12 | 90.4192 | -4.1609e-11 | 90.4332
|  | -5.0704e-11 | 90.8392
|  | -4.5702e-11 | 91.2462
|  | -4.3201e-11 | 91.6522
|  | -4.4338e-11 | 92.0603
|  | -4.9567e-11 | 92.4673
|  | -4.6384e-11 | 92.8743
|  | -4.502e-11 | 93.2943
|  | -4.7976e-11 | 93.7014
|  | -5.366e-11 | 94.1074
|  | -4.7976e-11 | 94.5124
|  | -4.2974e-11 | 94.9204
|  | -4.3428e-11 | 95.3294
|  | -4.4338e-11 | 95.7375
|  | -4.8203e-11 | 96.1445
|  | -5.0022e-11 | 96.5495
|  | -4.2519e-11 | 96.9556

## 630uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-3.5018e-12 | 2.0328e-12 | -6.6796e-11 | 5.5259e-12
GEMs On |  | GEMs On
-4.3201e-12 | 0.31002 | -5.9345e-11 | 0.31202
-2.2737e-12 | 0.99506 | -7.0031e-11 | 0.99606
-3.4106e-12 | 1.4001 | -7.4806e-11 | 1.4011
2.2737e-13 | 1.8041 | -6.5484e-11 | 1.8071
-4.3201e-12 | 2.2091 | -6.1391e-11 | 2.2111
-4.7748e-12 | 2.6132 | -7.0486e-11 | 2.6161
-4.7748e-12 | 3.0182 | -6.6166e-11 | 3.0222
4.5475e-13 | 3.4222 | -6.5256e-11 | 3.4272
-1.5916e-12 | 3.8262 | -6.0254e-11 | 3.8332
-6.3665e-12 | 4.2302 | -6.0709e-11 | 4.2402
-3.1832e-12 | 4.6343 | -7.9126e-11 | 4.6463
-2.0464e-12 | 5.0393 | -6.1164e-11 | 5.0523
-4.5475e-12 | 5.4443 | -6.6166e-11 | 5.4573
-1.5916e-12 | 5.8503 | -7.3896e-11 | 5.8623
-4.7748e-12 | 6.2534 | -6.4574e-11 | 6.2684
-1.1369e-12 | 6.6584 | -6.6166e-11 | 6.6734
-3.4106e-12 | 7.0634 | -6.5484e-11 | 7.0804
-3.8654e-12 | 7.4684 | -5.8208e-11 | 7.4854
-2.2737e-12 | 7.8735 | -6.753e-11 | 7.8905
-3.4106e-12 | 8.2785 | -6.9349e-11 | 8.2975
4.5475e-13 | 8.6835 | -7.2987e-11 | 8.7035
-3.8654e-12 | 9.0875 | -8.0036e-11 | 9.1095
-5.0022e-12 | 9.4915 | -5.9799e-11 | 9.5145
-4.3201e-12 | 9.8966 | -6.4801e-11 | 9.9206
-2.5011e-12 | 10.3006 | -6.7303e-11 | 10.3256
-5.0022e-12 | 10.7056 | -6.4574e-11 | 10.7316
-6.3665e-12 | 11.1096 | -6.4801e-11 | 11.1376
-3.1832e-12 | 11.5147 | -6.1164e-11 | 11.5437
-6.1391e-12 | 11.9197 | -6.0254e-11 | 11.9487
-4.3201e-12 | 12.3247 | -6.3437e-11 | 12.3537
-3.638e-12 | 12.7307 | -7.4806e-11 | 12.7597
-1.819e-12 | 13.1348 | -6.9122e-11 | 13.1668
-1.5916e-12 | 13.5388 | -6.6393e-11 | 13.5728
-3.4106e-12 | 13.9438 | -6.662e-11 | 13.9798
-1.1369e-12 | 14.3498 | -6.5256e-11 | 14.3848
-4.5475e-13 | 14.7548 | -7.2305e-11 | 14.7898
9.0949e-13 | 15.1589 | -7.2987e-11 | 15.1969
-4.5475e-12 | 15.5629 | -6.1618e-11 | 15.6029
-3.4106e-12 | 15.9659 | -6.8894e-11 | 16.0079
-6.3665e-12 | 16.3709 | -5.7753e-11 | 16.4149
0 | 16.776 | -6.3437e-11 | 16.82
-3.1832e-12 | 17.178 | -7.185e-11 | 17.226
-3.638e-12 | 17.583 | -7.7534e-11 | 17.632
-2.0464e-12 | 17.988 | -6.2755e-11 | 18.038
-3.1832e-12 | 18.393 | -6.0709e-11 | 18.443
-1.819e-12 | 18.7971 | -6.5256e-11 | 18.8491
-6.1391e-12 | 19.2021 | -6.2755e-11 | 19.2541
-9.0949e-13 | 19.6071 | -6.1618e-11 | 19.6611
-2.0464e-12 | 20.0121 | -7.3896e-11 | 20.0672
-4.5475e-12 | 20.4172 | -6.4574e-11 | 20.4722
-5.457e-12 | 20.8232 | -6.5029e-11 | 20.8782
-3.4106e-12 | 21.2282 | -7.185e-11 | 21.2832
0 | 21.6332 | -6.5938e-11 | 21.6892
-5.457e-12 | 22.0353 | -6.9804e-11 | 22.0963
-2.5011e-12 | 22.4393 | -6.2755e-11 | 22.5013
-1.819e-12 | 22.8443 | -6.9122e-11 | 22.9103
0 | 23.2503 | -7.2987e-11 | 23.3173
-3.4106e-12 | 23.6523 | -6.1618e-11 | 23.7224
-5.9117e-12 | 24.0574 | -6.2528e-11 | 24.1284
-2.2737e-12 | 24.4624 | -6.3437e-11 | 24.5324
1.3642e-12 | 24.8674 | -6.7303e-11 | 24.9374
-9.0949e-13 | 25.2724 | -6.1164e-11 | 25.3424
-3.8654e-12 | 25.6775 | -6.5256e-11 | 25.7535
-2.5011e-12 | 26.0835 | -7.0258e-11 | 26.1595
0 | 26.4875 | -6.5256e-11 | 26.5645
-6.3665e-12 | 26.8905 | -6.9349e-11 | 26.9715
-4.3201e-12 | 27.2946 | -6.3665e-11 | 27.3796
-3.1832e-12 | 27.6996 | -6.9349e-11 | 27.8036
-5.6843e-12 | 28.1056 | -6.6166e-11 | 28.2086
-5.0022e-12 | 28.5176 | -6.662e-11 | 28.6146
-2.9559e-12 | 28.9216 | -6.7303e-11 | 29.0207
-2.5011e-12 | 29.3257 | -6.3892e-11 | 29.4267
-5.0022e-12 | 29.7307 | -5.8662e-11 | 29.8327
-7.5033e-12 | 30.1367 | -7.7307e-11 | 30.2387
-3.8654e-12 | 30.5407 | -7.3896e-11 | 30.6438
-1.1369e-12 | 30.9448 | -6.5484e-11 | 31.0498
-5.9117e-12 | 31.3508 | -6.3892e-11 | 31.4548
-4.7748e-12 | 31.7558 | -6.1618e-11 | 31.8598
0 | 32.1608 | -6.662e-11 | 32.2668
-4.7748e-12 | 32.5629 | -5.3888e-11 | 32.6729
-6.1391e-12 | 32.9699 | -6.6166e-11 | 33.0789
-2.2737e-12 | 33.3729 | -6.8667e-11 | 33.4849
-3.8654e-12 | 33.7779 | -6.3665e-11 | 33.8909
-3.4106e-12 | 34.183 | -7.1623e-11 | 34.296
-8.4128e-12 | 34.587 | -6.7303e-11 | 34.701
-1.5916e-12 | 34.992 | -6.8439e-11 | 35.107
-2.2737e-12 | 35.396 | -7.8899e-11 | 35.511
-4.5475e-12 | 35.8021 | -6.3892e-11 | 35.9171
-6.3665e-12 | 36.2061 | -6.7985e-11 | 36.3221
-3.638e-12 | 36.6131 | -6.2755e-11 | 36.7281
-3.1832e-12 | 37.0171 | -6.5938e-11 | 37.1341
-5.0022e-12 | 37.4221 | -7.8899e-11 | 37.5401
-5.9117e-12 | 37.8262 | -7.1168e-11 | 37.9462
-2.9559e-12 | 38.2302 | -5.6843e-11 | 38.3512
-1.819e-12 | 38.6362 | -6.4574e-11 | 38.7572
-5.457e-12 | 39.0412 | -6.1391e-11 | 39.1632
-7.276e-12 | 39.4453 | -6.6166e-11 | 39.5673
-3.4106e-12 | 39.8503 | -6.5256e-11 | 39.9723
-1.1369e-12 | 40.2553 | -6.2073e-11 | 40.3783
-5.457e-12 | 40.6603 | -6.6393e-11 | 40.7833
-5.0022e-12 | 41.0643 | -6.0709e-11 | 41.1894
-2.0464e-12 | 41.4694 | -6.9804e-11 | 41.5954
-2.9559e-12 | 41.8744 | -6.1164e-11 | 42.0014
-5.0022e-12 | 42.2784 | -6.2073e-11 | 42.4054
-3.638e-12 | 42.6834 | -7.276e-11 | 42.8115
2.2737e-13 | 43.0885 | -5.9572e-11 | 43.2165
0 | 43.4925 | -6.753e-11 | 43.6215
-3.4106e-12 | 43.8955 | -6.9349e-11 | 44.0255
-4.7748e-12 | 44.3015 | -5.5934e-11 | 44.4325
-3.1832e-12 | 44.7076 | -7.0486e-11 | 44.8376
-7.0486e-12 | 45.1126 | -7.9353e-11 | 45.2416
-6.1391e-12 | 45.5176 | -5.6389e-11 | 45.6476
-3.1832e-12 | 45.9216 | -6.7075e-11 | 46.0526
-9.0949e-13 | 46.3276 | -6.5484e-11 | 46.4647
-3.638e-12 | 46.7307 | -6.2073e-11 | 46.8697
-3.638e-12 | 47.1347 | -6.753e-11 | 47.2757
-6.3665e-12 | 47.5397 | -7.276e-11 | 47.6817
-4.7748e-12 | 47.9437 | -7.6398e-11 | 48.0857
-2.9559e-12 | 48.3498 | -7.5488e-11 | 48.4928
-5.0022e-12 | 48.7548 | -6.7303e-11 | 48.8988
-9.0949e-13 | 49.1608 | -6.8894e-11 | 49.3048
2.2737e-13 | 49.5648 | -6.3665e-11 | 49.7098
-3.1832e-12 | 49.9689 | -6.6166e-11 | 50.1159
-4.5475e-12 | 50.3739 | -6.9349e-11 | 50.5219
-6.1391e-12 | 50.7779 | -6.8667e-11 | 50.9259
-1.5916e-12 | 51.1819 | -7.1395e-11 | 51.3319
-3.1832e-12 | 51.587 | -7.2532e-11 | 51.737
-5.457e-12 | 51.992 | -7.0941e-11 | 52.143
-3.8654e-12 | 52.396 | -7.2305e-11 | 52.548
-9.0949e-13 | 52.8 | -5.7071e-11 | 52.954
-6.8212e-13 | 53.205 | -6.7985e-11 | 53.3591
-4.7748e-12 | 53.6091 | -6.0709e-11 | 53.7651
-3.1832e-12 | 54.0151 | -6.753e-11 | 54.1711
-2.5011e-12 | 54.4211 | -7.6852e-11 | 54.5761
-1.1369e-12 | 54.8261 | -6.8894e-11 | 54.9821
-5.0022e-12 | 55.2312 | -7.1395e-11 | 55.3882
9.0949e-13 | 55.6352 | -6.7075e-11 | 55.7922
-9.0949e-13 | 56.0392 | -7.276e-11 | 56.1972
-4.3201e-12 | 56.4432 | -6.7075e-11 | 56.6032
-4.7748e-12 | 56.8483 | -7.1168e-11 | 57.0083
-3.638e-12 | 57.2533 | -6.9122e-11 | 57.4133
-3.4106e-12 | 57.6583 | -6.4801e-11 | 57.8183
-2.2737e-12 | 58.0633 | -6.1618e-11 | 58.2233
-1.5916e-12 | 58.4673 | -7.3669e-11 | 58.6283
-4.5475e-13 | 58.8734 | -6.7985e-11 | 59.0364
4.5475e-13 | 59.2774 | -6.5029e-11 | 59.4414
-3.8654e-12 | 59.6804 | -6.8667e-11 | 59.8474
-4.5475e-12 | 60.0854 | -6.0936e-11 | 60.2535
-3.8654e-12 | 60.4895 | -6.753e-11 | 60.6595
-3.4106e-12 | 60.8955 | -6.0936e-11 | 61.0645
-3.1832e-12 | 61.3005 | -7.0486e-11 | 61.4695
-5.0022e-12 | 61.7055 | -6.6393e-11 | 61.8755
6.8212e-13 | 62.1106 | -6.7303e-11 | 62.2816
-2.9559e-12 | 62.5166 | -6.7303e-11 | 62.6876
-3.1832e-12 | 62.9226 | -5.6161e-11 | 63.0936
-6.8212e-12 | 63.3266 | -6.9349e-11 | 63.4986
-6.8212e-13 | 63.7326 | -7.6398e-11 | 63.9037
-1.5916e-12 | 64.1357 | -6.8439e-11 | 64.3097
-7.276e-12 | 64.5397 | -6.7303e-11 | 64.7147
-7.276e-12 | 64.9447 | -7.2305e-11 | 65.1197
-2.0464e-12 | 65.3497 | -7.0258e-11 | 65.5258
-5.6843e-12 | 65.7548 | -6.321e-11 | 65.9308
-2.0464e-12 | 66.1588 | -6.3665e-11 | 66.3358
-5.9117e-12 | 66.5638 | -8.1855e-11 | 66.7408
-3.8654e-12 | 66.9698 | -6.5029e-11 | 67.1468
-1.819e-12 | 67.3739 | -5.8662e-11 | 67.5529
-3.4106e-12 | 67.7779 | -7.0031e-11 | 67.9579
-5.9117e-12 | 68.1839 | -7.0941e-11 | 68.3649
-4.7748e-12 | 68.5869 | -6.662e-11 | 68.7699
-5.457e-12 | 68.9929 | -6.7303e-11 | 69.175
-7.0486e-12 | 69.398 | -5.8208e-11 | 69.58
-2.5011e-12 | 69.803 | -5.3433e-11 | 69.985
-6.8212e-13 | 70.209 | -6.0936e-11 | 70.391
-2.2737e-12 | 70.613 | -6.7757e-11 | 70.797
-4.3201e-12 | 71.0181 | -7.7762e-11 | 71.2031
-2.2737e-12 | 71.4221 | -7.7762e-11 | 71.6091
-1.1369e-12 | 71.8281 | -7.4806e-11 | 72.0141
-3.8654e-12 | 72.2331 | -6.1164e-11 | 72.4201
-3.638e-12 | 72.6381 | -6.2528e-11 | 72.8252
-4.5475e-13 | 73.0432 | -7.185e-11 | 73.2322
-3.4106e-12 | 73.4482 | -6.321e-11 | 73.6372
-6.8212e-12 | 73.8532 | -6.5938e-11 | 74.0422
-4.3201e-12 | 74.2572 | -6.23e-11 | 74.4483
-5.457e-12 | 74.6623 | -6.8439e-11 | 74.8543
-5.0022e-12 | 75.0663 | -6.3892e-11 | 75.2593
-4.7748e-12 | 75.4713 | -6.2073e-11 | 75.6643
-3.4106e-12 | 75.8803 | -6.3665e-11 | 76.0713
-6.1391e-12 | 76.2854 | -6.9804e-11 | 76.4774
-5.6843e-12 | 76.6914 | -8.5493e-11 | 76.8824
-6.8212e-13 | 77.1004 | -5.9799e-11 | 77.2894
-7.7307e-12 | 77.5054 | -7.3669e-11 | 77.6954
-6.1391e-12 | 77.9115 | -6.4574e-11 | 78.1015
-3.638e-12 | 78.3165 | -6.3892e-11 | 78.5065
-3.8654e-12 | 78.7215 | -6.4574e-11 | 78.9125
-4.7748e-12 | 79.1265 | -7.1168e-11 | 79.3175
-3.638e-12 | 79.5326 | -6.4574e-11 | 79.7226
-2.0464e-12 | 79.9386 | -7.6852e-11 | 80.1266
-4.7748e-12 | 80.3426 | -6.2073e-11 | 80.5326
-6.1391e-12 | 80.7486 | -6.3892e-11 | 80.9396
-4.3201e-12 | 81.1526 | -6.3665e-11 | 81.3457
2.2737e-13 | 81.5567 | -6.6166e-11 | 81.7517
-2.9559e-12 | 81.9827 | -6.7757e-11 | 82.1787
-3.4106e-12 | 82.3877 | -5.6161e-11 | 82.5847
-2.0464e-12 | 82.7927 | -6.9804e-11 | 82.9898
-5.457e-12 | 83.1988 | -7.4806e-11 | 83.3958
-4.5475e-13 | 83.6038 | -6.4119e-11 | 83.8018
-6.3665e-12 | 84.0078 | -6.3437e-11 | 84.2068
-3.1832e-12 | 84.4128 | -6.23e-11 | 84.6128
-5.9117e-12 | 84.8188 | -6.2073e-11 | 85.0179
-5.0022e-12 | 85.2239 | -6.4801e-11 | 85.4239
-5.457e-12 | 85.6279
-3.8654e-12 | 86.0339
-5.0022e-12 | 86.4389
-3.4106e-12 | 86.844
-4.3201e-12 | 87.248
-4.7748e-12 | 87.655
-5.457e-12 | 88.06

## 640uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-4.0873e-12 | 1.8483e-12 | -9.1954e-11 | 7.3798e-12
GEMs On |  | GEMs On
-6.1391e-12 | 0.31102 | -1.03e-10 | 0.31402
-3.638e-12 | 0.99506 | -9.3223e-11 | 0.99806
-1.5916e-12 | 1.4001 | -7.6398e-11 | 1.4051
-3.638e-12 | 1.8051 | -8.4128e-11 | 1.8101
-5.2296e-12 | 2.2121 | -8.7539e-11 | 2.2151
-3.638e-12 | 2.6161 | -9.959e-11 | 2.6212
-2.2737e-12 | 3.0212 | -1.1096e-10 | 3.0272
-4.3201e-12 | 3.4262 | -1.0436e-10 | 3.4322
-3.8654e-12 | 3.8312 | -8.8676e-11 | 3.8362
-3.638e-12 | 4.2362 | -8.049e-11 | 4.2432
-6.8212e-13 | 4.6413 | -9.527e-11 | 4.6473
-5.2296e-12 | 5.0463 | -9.2541e-11 | 5.0523
-5.2296e-12 | 5.4523 | -8.2991e-11 | 5.4583
-3.8654e-12 | 5.8573 | -1.0323e-10 | 5.8643
-1.5916e-12 | 6.2614 | -9.7998e-11 | 6.2694
-2.2737e-12 | 6.6664 | -1.0414e-10 | 6.6744
-6.1391e-12 | 7.0704 | -7.8444e-11 | 7.0794
-3.8654e-12 | 7.4754 | -9.3905e-11 | 7.4854
-6.8212e-12 | 7.8815 | -1.0141e-10 | 7.8905
-5.0022e-12 | 8.2865 | -9.4587e-11 | 8.2955
-4.7748e-12 | 8.6915 | -9.5724e-11 | 8.7015
-4.3201e-12 | 9.0965 | -9.6406e-11 | 9.1075
-5.6843e-12 | 9.5005 | -8.6402e-11 | 9.5135
-5.457e-12 | 9.9056 | -8.6175e-11 | 9.9196
-5.0022e-12 | 10.3096 | -8.7539e-11 | 10.3246
-2.2737e-12 | 10.7156 | -9.2086e-11 | 10.7306
-7.9581e-12 | 11.1206 | -9.436e-11 | 11.1346
-4.3201e-12 | 11.5277 | -9.2541e-11 | 11.5407
-3.638e-12 | 11.9317 | -8.1855e-11 | 11.9467
-5.6843e-12 | 12.3377 | -7.7534e-11 | 12.3507
-5.2296e-12 | 12.7427 | -9.5952e-11 | 12.7577
-3.638e-12 | 13.1478 | -9.1404e-11 | 13.1638
-2.7285e-12 | 13.5528 | -8.913e-11 | 13.5708
-2.2737e-12 | 13.9588 | -9.3905e-11 | 13.9758
-7.0486e-12 | 14.3638 | -9.5952e-11 | 14.3808
-4.0927e-12 | 14.7688 | -1.0277e-10 | 14.7858
-1.3642e-12 | 15.1729 | -8.7539e-11 | 15.1919
-3.8654e-12 | 15.5789 | -1.1119e-10 | 15.5969
-5.457e-12 | 15.9829 | -9.4587e-11 | 16.0009
-2.7285e-12 | 16.3879 | -8.5265e-11 | 16.4059
-4.3201e-12 | 16.793 | -9.436e-11 | 16.811
-7.9581e-12 | 17.198 | -1.0368e-10 | 17.216
-7.0486e-12 | 17.604 | -1.005e-10 | 17.621
-4.3201e-12 | 18.009 | -9.7771e-11 | 18.027
-5.0022e-12 | 18.413 | -9.5497e-11 | 18.4321
-4.7748e-12 | 18.8181 | -8.8903e-11 | 18.8381
-4.3201e-12 | 19.2231 | -1.0095e-10 | 19.2441
-3.8654e-12 | 19.6271 | -8.7766e-11 | 19.6481
-4.3201e-12 | 20.0312 | -9.5497e-11 | 20.0531
-5.6843e-12 | 20.4372 | -8.7539e-11 | 20.4582
-6.3665e-12 | 20.8412 | -9.6861e-11 | 20.8642
-1.1369e-12 | 21.2462 | -1.0459e-10 | 21.2702
-2.7285e-12 | 21.6512 | -1.0732e-10 | 21.6752
-7.5033e-12 | 22.0563 | -8.5265e-11 | 22.0813
-5.2296e-12 | 22.4623 | -9.959e-11 | 22.4863
-1.1369e-12 | 22.8683 | -7.9808e-11 | 22.8923
-6.8212e-12 | 23.2733 | -9.5952e-11 | 23.2983
-5.2296e-12 | 23.6774 | -9.3451e-11 | 23.7044
-4.7748e-12 | 24.0814 | -9.4587e-11 | 24.1114
-1.819e-12 | 24.4874 | -8.5265e-11 | 24.5174
-5.0022e-12 | 24.8914 | -9.8453e-11 | 24.9224
-7.0486e-12 | 25.2954 | -8.1627e-11 | 25.3295
-3.8654e-12 | 25.6995 | -7.2077e-11 | 25.7355
-3.1832e-12 | 26.1055 | -8.1855e-11 | 26.1415
-5.0022e-12 | 26.5105 | -8.7766e-11 | 26.5485
-5.0022e-12 | 26.9145 | -8.0036e-11 | 26.9555
-2.7285e-12 | 27.3206 | -9.3451e-11 | 27.3616
-2.2737e-12 | 27.7256 | -8.7539e-11 | 27.7676
-2.9559e-12 | 28.1326 | -1.0027e-10 | 28.1726
-3.638e-12 | 28.5376 | -9.436e-11 | 28.5786
2.2737e-13 | 28.9417 | -1.0027e-10 | 28.9847
-4.0927e-12 | 29.3467 | -8.9585e-11 | 29.3897
-4.7748e-12 | 29.7517 | -9.6406e-11 | 29.7947
-1.3642e-12 | 30.1567 | -9.0722e-11 | 30.2007
-4.5475e-13 | 30.5628 | -9.1177e-11 | 30.6058
-3.8654e-12 | 30.9668 | -9.3223e-11 | 31.0108
-6.1391e-12 | 31.3728 | -8.2309e-11 | 31.4158
6.8212e-13 | 31.7768 | -7.7534e-11 | 31.8208
-2.5011e-12 | 32.1808 | -9.8453e-11 | 32.2268
-6.1391e-12 | 32.5869 | -9.5042e-11 | 32.6329
-2.9559e-12 | 32.9909 | -8.7994e-11 | 33.0379
-2.2737e-12 | 33.3939 | -1.0687e-10 | 33.4439
-5.457e-12 | 33.7989 | -8.0945e-11 | 33.8499
-4.3201e-12 | 34.204 | -9.6406e-11 | 34.256
-3.638e-12 | 34.609 | -9.2768e-11 | 34.661
-2.9559e-12 | 35.014 | -9.8453e-11 | 35.066
-5.0022e-12 | 35.419 | -9.8908e-11 | 35.472
-8.1855e-12 | 35.8231 | -8.7311e-11 | 35.877
-3.8654e-12 | 36.2281 | -9.9362e-11 | 36.2821
-3.8654e-12 | 36.6361 | -8.7994e-11 | 36.6881
-4.7748e-12 | 37.0411 | -8.6175e-11 | 37.0931
-3.1832e-12 | 37.4461 | -9.8908e-11 | 37.4992
-1.819e-12 | 37.8522 | -1.0027e-10 | 37.9052
-4.3201e-12 | 38.2572 | -9.2996e-11 | 38.3112
-4.3201e-12 | 38.6622 | -7.7307e-11 | 38.7172
-2.2737e-12 | 39.0672 | -8.8676e-11 | 39.1212
-1.3642e-12 | 39.4723 | -9.004e-11 | 39.5263
-5.457e-12 | 39.8763 | -8.8221e-11 | 39.9333
-5.2296e-12 | 40.2813 | -9.527e-11 | 40.3393
-4.3201e-12 | 40.6863 | -7.5488e-11 | 40.7443
-1.3642e-12 | 41.0924 | -8.5265e-11 | 41.1514
-5.0022e-12 | 41.4974 | -9.8453e-11 | 41.5584
-4.7748e-12 | 41.9024 | -9.2768e-11 | 41.9644
-2.2737e-12 | 42.3074 | -8.7539e-11 | 42.3704
-2.2737e-12 | 42.7124 | -9.0495e-11 | 42.7764
-4.3201e-12 | 43.1165 | -8.4128e-11 | 43.1825
-2.7285e-12 | 43.5215 | -1.0095e-10 | 43.5875
-5.2296e-12 | 43.9285 | -9.3905e-11 | 43.9945
-4.3201e-12 | 44.3335 | -9.436e-11 | 44.3995
-4.7748e-12 | 44.7386 | -1.0164e-10 | 44.8066
-3.638e-12 | 45.1426 | -1.0527e-10 | 45.2116
-4.0927e-12 | 45.5476 | -9.1632e-11 | 45.6176
-7.5033e-12 | 45.9546 | -9.7543e-11 | 46.0226
-6.5938e-12 | 46.3586 | -1.0186e-10 | 46.4277
-2.5011e-12 | 46.7647 | -8.7994e-11 | 46.8327
-5.2296e-12 | 47.1687 | -8.2537e-11 | 47.2377
-2.9559e-12 | 47.5747 | -9.5952e-11 | 47.6417
-2.7285e-12 | 47.9807 | -9.5497e-11 | 48.0478
-5.0022e-12 | 48.3838 | -8.7311e-11 | 48.4528
-5.457e-12 | 48.7888 | -8.6402e-11 | 48.8588
-6.8212e-12 | 49.1928 | -9.0722e-11 | 49.2658
-1.819e-12 | 49.5978 | -1.0687e-10 | 49.6718
2.2737e-13 | 50.0039 | -9.9817e-11 | 50.0779
-4.7748e-12 | 50.4079 | -9.7089e-11 | 50.4829
-1.3642e-12 | 50.8129 | -9.4587e-11 | 50.8889
-4.7748e-12 | 51.2189 | -7.9808e-11 | 51.2929
-2.9559e-12 | 51.623 | -8.6857e-11 | 51.698
-4.3201e-12 | 52.028 | -9.3223e-11 | 52.104
-4.0927e-12 | 52.433 | -9.5724e-11 | 52.508
-1.819e-12 | 52.838 | -8.2991e-11 | 52.914
-2.2737e-13 | 53.242 | -9.0495e-11 | 53.318
-4.7748e-12 | 53.6471 | -1.0459e-10 | 53.7241
-6.3665e-12 | 54.0511 | -9.4133e-11 | 54.1291
-4.0927e-12 | 54.4561 | -8.3901e-11 | 54.5351
-5.6843e-12 | 54.8611 | -8.7994e-11 | 54.9401
-8.6402e-12 | 55.2662 | -9.1859e-11 | 55.3462
-5.0022e-12 | 55.6722 | -9.3451e-11 | 55.7532
-1.819e-12 | 56.0772 | -8.3446e-11 | 56.1582
-2.7285e-12 | 56.4812 | -9.8225e-11 | 56.5632
-3.1832e-12 | 56.8862 | -9.9362e-11 | 56.9673
-2.9559e-12 | 57.2933 | -9.8225e-11 | 57.3723
-5.2296e-12 | 57.6983 | -9.5724e-11 | 57.7783
-7.5033e-12 | 58.1033 | -8.7311e-11 | 58.1833
-1.819e-12 | 58.5084 | -9.2086e-11 | 58.5883
-1.3642e-12 | 58.9134 | -1.0459e-10 | 58.9954
-7.5033e-12 | 59.3174 | -8.7311e-11 | 59.4004
-4.0927e-12 | 59.7234 | -9.4133e-11 | 59.8064
-2.5011e-12 | 60.1274 | -8.9585e-11 | 60.2124
-2.2737e-12 | 60.5325 | -8.913e-11 | 60.6165
-1.1369e-12 | 60.9365 | -9.7543e-11 | 61.0225
-4.3201e-12 | 61.3405 | -8.7766e-11 | 61.4275
-3.638e-12 | 61.7465 | -7.5943e-11 | 61.8335
-5.457e-12 | 62.1506 | -9.4133e-11 | 62.2396
-5.0022e-12 | 62.5556 | -9.0267e-11 | 62.6456
-3.8654e-12 | 62.9606 | -7.7989e-11 | 63.0516
-6.1391e-12 | 63.3656 | -9.7543e-11 | 63.4566
-1.5916e-12 | 63.7696 | -9.004e-11 | 63.8616
-5.6843e-12 | 64.1757 | -9.7998e-11 | 64.2677
-5.457e-12 | 64.5807 | -9.7543e-11 | 64.6737
-4.5475e-13 | 64.9857 | -8.8903e-11 | 65.0807
-5.0022e-12 | 65.3887 | -8.4356e-11 | 65.4868
-3.1832e-12 | 65.7938 | -9.1632e-11 | 65.8928
-6.3665e-12 | 66.1988 | -7.8671e-11 | 66.2978
0 | 66.6028 | -8.6629e-11 | 66.7028
-4.5475e-13 | 67.0038 | -8.3446e-11 | 67.1088
-4.7748e-12 | 67.4079 | -9.959e-11 | 67.5159
-7.9581e-12 | 67.8129 | -8.6857e-11 | 67.9199
-3.1832e-12 | 68.2179 | -8.572e-11 | 68.3239
-2.5011e-12 | 68.6229 | -8.572e-11 | 68.7299
-5.0022e-12 | 69.0269 | -8.572e-11 | 69.135
-3.638e-12 | 69.431 | -8.8221e-11 | 69.54
-6.8212e-13 | 69.836 | -9.2541e-11 | 69.945
-4.0927e-12 | 70.241 | -8.9358e-11 | 70.349
-5.2296e-12 | 70.646 | -9.7543e-11 | 70.755
-5.2296e-12 | 71.0501 | -9.2996e-11 | 71.1601
-5.0022e-12 | 71.4561 | -8.9358e-11 | 71.5651
-1.5916e-12 | 71.8611 | -9.1404e-11 | 71.9701
-4.3201e-12 | 72.2661 | -8.2764e-11 | 72.3761
-4.0927e-12 | 72.6702 | -9.3223e-11 | 72.7822
-1.5916e-12 | 73.0772 | -9.2541e-11 | 73.1882
-4.3201e-12 | 73.4822 | -8.5265e-11 | 73.5922
-6.8212e-12 | 73.8862 | -8.572e-11 | 73.9982
-5.457e-12 | 74.2913 | -8.0945e-11 | 74.4043
-3.1832e-12 | 74.6963 | -8.0263e-11 | 74.8093
-8.6402e-12 | 75.1013 | -9.1404e-11 | 75.2153
-3.638e-12 | 75.5053 | -8.5265e-11 | 75.6223
-5.6843e-12 | 75.9093 | -8.8903e-11 | 76.0284
-3.638e-12 | 76.3134 | -8.1627e-11 | 76.4324
-2.7285e-12 | 76.7174 | -9.9362e-11 | 76.8374
-5.6843e-12 | 77.1224 | -9.0495e-11 | 77.2434
-4.7748e-12 | 77.5284 | -9.3451e-11 | 77.6494
-4.7748e-12 | 77.9325 | -1.0095e-10 | 78.0555
-4.0927e-12 | 78.3375 | -1.0436e-10 | 78.4605
-4.7748e-12 | 78.7435 | -8.913e-11 | 78.8665
-6.5938e-12 | 79.1495 | -9.1859e-11 | 79.2725
-5.2296e-12 | 79.5546 | -8.3674e-11 | 79.6786
-5.2296e-12 | 79.9596 | -9.6634e-11 | 80.0856
-4.0927e-12 | 80.3656 | -1.0732e-10 | 80.4916
-5.0022e-12 | 80.7696 | -9.7771e-11 | 80.8966
-5.0022e-12 | 81.1746 | -9.5497e-11 | 81.3016
-3.1832e-12 | 81.5807 | -9.004e-11 | 81.7087
-4.0927e-12 | 82.0077 | -7.8899e-11 | 82.1367
-2.9559e-12 | 82.4127 | -8.9358e-11 | 82.5427
-2.9559e-12 | 82.8167 | -9.527e-11 | 82.9487
-8.6402e-12 | 83.2228 | -9.6634e-11 | 83.3538
-3.8654e-12 | 83.6278 | -9.7543e-11 | 83.7588
-3.638e-12 | 84.0328 | -9.0267e-11 | 84.1638
-4.0927e-12 | 84.4378 | -9.959e-11 | 84.5698
-3.8654e-12 | 84.8428 | -9.0722e-11 | 84.9749
-2.7285e-12 | 85.2479 | -9.527e-11 | 85.3809
|  | -1.0186e-10 | 85.7859
|  | -9.004e-11 | 86.1899
|  | -9.6406e-11 | 86.596
|  | -9.4587e-11 | 87.001
|  | -8.2991e-11 | 87.406
|  | -8.7766e-11 | 87.812
|  | -9.7089e-11 | 88.2171
|  | -9.1632e-11 | 88.6221
|  | -9.9362e-11 | 89.0291
|  | -8.1627e-11 | 89.4351
|  | -8.7311e-11 | 89.8411

## 650uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-4.9719e-12 | 1.7634e-12 | -1.2918e-10 | 9.0658e-12
GEMs On |  | GEMs On
-6.1391e-12 | 0.31002 | -1.2392e-10 | 0.31002
-5.6843e-12 | 0.99406 | -1.3733e-10 | 0.99506
-4.3201e-12 | 1.3991 | -1.2824e-10 | 1.4021
-7.0486e-12 | 1.8041 | -1.4438e-10 | 1.8061
-7.5033e-12 | 2.2091 | -1.3551e-10 | 2.2101
-5.0022e-12 | 2.6151 | -1.3915e-10 | 2.6161
-6.1391e-12 | 3.0192 | -1.6144e-10 | 3.0212
-5.457e-12 | 3.4242 | -1.296e-10 | 3.4262
-6.5938e-12 | 3.8302 | -1.3551e-10 | 3.8312
-6.5938e-12 | 4.2352 | -1.1732e-10 | 4.2362
-3.1832e-12 | 4.6403 | -1.1096e-10 | 4.6413
-4.3201e-12 | 5.0443 | -1.4211e-10 | 5.0473
-5.457e-12 | 5.4503 | -1.1846e-10 | 5.4633
-2.5011e-12 | 5.8553 | -1.3711e-10 | 5.8683
-4.0927e-12 | 6.2614 | -1.4029e-10 | 6.2724
-3.1832e-12 | 6.6644 | -1.3642e-10 | 6.6784
-9.3223e-12 | 7.0694 | -1.4074e-10 | 7.0834
-6.1391e-12 | 7.4734 | -1.3051e-10 | 7.4884
-4.5475e-12 | 7.8785 | -1.2756e-10 | 7.8925
-3.638e-12 | 8.2845 | -1.1119e-10 | 8.3015
-6.3665e-12 | 8.6895 | -1.2119e-10 | 8.7075
-5.2296e-12 | 9.0945 | -1.3938e-10 | 9.1115
-2.9559e-12 | 9.4985 | -1.3006e-10 | 9.5165
-2.9559e-12 | 9.9036 | -1.3142e-10 | 9.9206
-3.8654e-12 | 10.3096 | -1.2756e-10 | 10.3256
-6.1391e-12 | 10.7146 | -1.4074e-10 | 10.9336
-4.3201e-12 | 11.1196 | -1.3461e-10 | 11.4056
-7.5033e-12 | 11.5247 | -1.3529e-10 | 11.8117
-4.0927e-12 | 11.9297 | -1.2187e-10 | 12.2167
-6.8212e-12 | 12.3357 | -1.296e-10 | 12.6447
-7.0486e-12 | 12.7407 | -1.4438e-10 | 13.0587
-5.2296e-12 | 13.1457 | -1.3029e-10 | 13.4678
-2.7285e-12 | 13.5518 | -1.3006e-10 | 13.9088
-6.8212e-12 | 13.9568 | -1.1482e-10 | 14.3158
-6.1391e-12 | 14.3628 | -1.2437e-10 | 14.7208
-2.7285e-12 | 14.7678 | -1.2687e-10 | 15.1259
-4.5475e-12 | 15.1729 | -1.271e-10 | 15.5309
-6.3665e-12 | 15.5799 | -1.3347e-10 | 15.9359
-7.0486e-12 | 15.9849 | -1.1937e-10 | 16.3429
-7.5033e-12 | 16.3899 | -1.3438e-10 | 16.748
-1.819e-12 | 16.795 | -1.1983e-10 | 17.152
-6.1391e-12 | 17.2 | -1.3574e-10 | 17.556
-2.2737e-12 | 17.605 | -1.3256e-10 | 17.96
-2.7285e-12 | 18.017 | -1.4097e-10 | 18.3661
-6.8212e-12 | 18.4271 | -1.1892e-10 | 18.7711
-5.2296e-12 | 18.8311 | -1.2619e-10 | 19.1761
-6.3665e-12 | 19.2351 | -1.2187e-10 | 19.5811
-4.3201e-12 | 19.6421 | -1.321e-10 | 19.9861
-6.3665e-12 | 20.0471 | -1.3301e-10 | 20.3912
-3.1832e-12 | 20.4532 | -1.2869e-10 | 20.7962
-6.8212e-12 | 20.8582 | -1.2665e-10 | 21.2022
-7.0486e-12 | 21.2642 | -1.3665e-10 | 21.6062
-6.1391e-12 | 21.6732 | -1.2915e-10 | 22.0103
-2.7285e-12 | 22.0813 | -1.3551e-10 | 22.4163
-4.3201e-12 | 22.4863 | -1.3392e-10 | 22.8203
-7.9581e-12 | 22.8943 | -1.4438e-10 | 23.2253
-4.0927e-12 | 23.3003 | -1.2778e-10 | 23.6303
-5.0022e-12 | 23.7044 | -1.2301e-10 | 24.0354
-6.5938e-12 | 24.1084 | -1.3688e-10 | 24.4394
-6.1391e-12 | 24.5154 | -1.2915e-10 | 24.8444
-4.3201e-12 | 24.9204 | -1.3642e-10 | 25.2474
-3.638e-12 | 25.3285 | -1.2687e-10 | 25.6535
-5.2296e-12 | 25.7335 | -1.4506e-10 | 26.0575
-5.2296e-12 | 26.1385 | -1.2574e-10 | 26.4605
-5.2296e-12 | 26.5435 | -1.3802e-10 | 26.8655
-6.1391e-12 | 26.9495 | -1.2619e-10 | 27.2716
-6.8212e-12 | 27.3556 | -1.1869e-10 | 27.6756
-1.0459e-11 | 27.7626 | -1.4734e-10 | 28.0806
-3.638e-12 | 28.1686 | -1.2892e-10 | 28.4856
-5.457e-12 | 28.5736 | -1.3506e-10 | 28.8896
-7.7307e-12 | 28.9797 | -1.2233e-10 | 29.2937
-6.8212e-12 | 29.3897 | -1.1755e-10 | 29.6997
-3.1832e-12 | 29.7957 | -1.3142e-10 | 30.1037
-9.0949e-12 | 30.2007 | -1.3097e-10 | 30.5077
-4.3201e-12 | 30.6058 | -1.2778e-10 | 30.9118
-2.9559e-12 | 31.0198 | -1.1141e-10 | 31.3168
-3.638e-12 | 31.4248 | -1.1755e-10 | 31.7198
-4.5475e-12 | 31.8318 | -1.1278e-10 | 32.1258
-3.8654e-12 | 32.2378 | -1.2483e-10 | 32.5309
-5.6843e-12 | 32.6429 | -1.1482e-10 | 32.9359
-5.2296e-12 | 33.0499 | -1.2164e-10 | 33.3419
-1.3642e-12 | 33.4559 | -1.2142e-10 | 33.7469
-5.6843e-12 | 33.8599 | -1.3438e-10 | 34.151
-3.638e-12 | 34.266 | -1.0914e-10 | 34.556
-7.9581e-12 | 34.672 | -1.4097e-10 | 34.962
-1.819e-12 | 35.077 | -1.1983e-10 | 35.367
-6.8212e-12 | 35.481 | -1.4188e-10 | 35.7721
-6.3665e-12 | 35.886 | -1.1391e-10 | 36.1781
-6.5938e-12 | 36.2931 | -1.2369e-10 | 36.5831
-4.5475e-12 | 36.6971 | -1.271e-10 | 36.9871
-5.457e-12 | 37.1011 | -1.146e-10 | 37.3921
-5.457e-12 | 37.5061 | -1.3961e-10 | 37.7962
-5.2296e-12 | 37.9122 | -1.2778e-10 | 38.2002
-6.3665e-12 | 38.3182 | -1.2642e-10 | 38.6062
-5.6843e-12 | 38.7242 | -1.2278e-10 | 39.0102
-5.0022e-12 | 39.1292 | -1.121e-10 | 39.4153
-3.638e-12 | 39.5353 | -1.2801e-10 | 39.8213
-6.5938e-12 | 39.9403 | -1.3256e-10 | 40.2253
-7.0486e-12 | 40.3473 | -1.3529e-10 | 40.6303
-6.8212e-12 | 40.7523 | -1.3893e-10 | 41.0354
-4.5475e-12 | 41.1594 | -1.2096e-10 | 41.4414
-3.8654e-12 | 41.5654 | -1.0914e-10 | 41.8444
-3.8654e-12 | 41.9704 | -1.2483e-10 | 42.2494
-8.4128e-12 | 42.3794 | -1.3506e-10 | 42.6524
-5.6843e-12 | 42.7844 | -1.3665e-10 | 43.0565
-5.6843e-12 | 43.2015 | -1.2483e-10 | 43.4615
-6.8212e-12 | 43.6055 | -1.2278e-10 | 43.8655
-7.5033e-12 | 44.0115 | -1.4802e-10 | 44.2695
-8.4128e-12 | 44.4165 | -1.2665e-10 | 44.6735
-8.4128e-12 | 44.8216 | -1.2869e-10 | 45.0796
-2.9559e-12 | 45.2276 | -1.3006e-10 | 45.4846
-5.0022e-12 | 45.6326 | -1.2824e-10 | 45.8896
-4.3201e-12 | 46.0366 | -1.3983e-10 | 46.2936
-7.7307e-12 | 46.4417 | -1.3961e-10 | 46.6987
-6.1391e-12 | 46.8477 | -1.3392e-10 | 47.1027
-2.7285e-12 | 47.2527 | -1.4188e-10 | 47.5067
-6.1391e-12 | 47.6587 | -1.3324e-10 | 47.9127
-2.5011e-12 | 48.0628 | -1.2324e-10 | 48.3158
-3.8654e-12 | 48.4678 | -1.2369e-10 | 48.7218
-5.2296e-12 | 48.8748 | -1.3711e-10 | 49.1248
-6.3665e-12 | 49.2818 | -1.3074e-10 | 49.5308
-7.0486e-12 | 49.6938 | -1.2892e-10 | 49.9349
-7.5033e-12 | 50.1009 | -1.0959e-10 | 50.3399
-2.7285e-12 | 50.5059 | -1.3711e-10 | 50.7449
-2.7285e-12 | 50.9119 | -1.3529e-10 | 51.1519
-6.8212e-12 | 51.3159 | -1.3097e-10 | 51.5579
-4.5475e-12 | 51.721 | -1.3642e-10 | 51.963
-4.0927e-12 | 52.126 | -1.3188e-10 | 52.367
-4.3201e-12 | 52.533 | -1.3188e-10 | 52.773
-2.9559e-12 | 52.942 | -1.296e-10 | 53.179
-5.2296e-12 | 53.3481 | -1.2619e-10 | 53.5841
-6.5938e-12 | 53.7531 | -1.2619e-10 | 53.9891
-3.1832e-12 | 54.1591 | -1.0073e-10 | 54.3951
-2.2737e-12 | 54.5641 | -1.3915e-10 | 54.7991
-5.2296e-12 | 54.9701 | -1.2483e-10 | 55.2022
-4.3201e-12 | 55.3782 | -1.2255e-10 | 55.6072
-5.457e-12 | 55.7832 | -1.3324e-10 | 56.0112
-2.5011e-12 | 56.1882 | -1.0687e-10 | 56.4162
-3.1832e-12 | 56.5932 | -1.3188e-10 | 56.8212
-4.0927e-12 | 56.9973 | -1.2301e-10 | 57.2263
-5.6843e-12 | 57.4023 | -1.3438e-10 | 57.6303
-5.2296e-12 | 57.8063 | -1.3961e-10 | 58.0353
-6.5938e-12 | 58.2123 | -1.1801e-10 | 58.4393
-4.0927e-12 | 58.6174 | -1.2619e-10 | 58.8434
-7.0486e-12 | 59.0234 | -1.3074e-10 | 59.2474
-2.7285e-12 | 59.4274 | -1.1141e-10 | 59.6524
-5.2296e-12 | 59.8324 | -1.3461e-10 | 60.0574
-2.5011e-12 | 60.2385 | -1.1141e-10 | 60.4625
-2.5011e-12 | 60.6435 | -1.3983e-10 | 60.8675
-6.1391e-12 | 61.0495 | -1.1619e-10 | 61.2715
-5.6843e-12 | 61.4555 | -1.2255e-10 | 61.6755
-5.2296e-12 | 61.8615 | -1.2892e-10 | 62.0795
-5.6843e-12 | 62.2656 | -1.2824e-10 | 62.4846
-5.457e-12 | 62.6716 | -1.3165e-10 | 62.8886
-2.9559e-12 | 63.0766 | -1.321e-10 | 63.2936
-4.0927e-12 | 63.4806 | -1.2574e-10 | 63.6976
-5.6843e-12 | 63.8856 | -1.4052e-10 | 64.1037
-8.4128e-12 | 64.2907 | -1.2324e-10 | 64.5077
-2.9559e-12 | 64.6967 | -1.1869e-10 | 64.9127
-3.638e-12 | 65.1017 | -1.2119e-10 | 65.3177
-1.819e-12 | 65.5067 | -1.271e-10 | 65.7238
-5.457e-12 | 65.9118 | -1.2187e-10 | 66.1278
-4.5475e-12 | 66.3168 | -1.3529e-10 | 66.5338
-2.5011e-12 | 66.7228 | -1.3029e-10 | 66.9398
-6.8212e-12 | 67.1268 | -1.2892e-10 | 67.3458
-6.3665e-12 | 67.5319 | -1.2756e-10 | 67.7509
-2.9559e-12 | 67.9379 | -1.3301e-10 | 68.1549
-5.2296e-12 | 68.3429 | -1.4802e-10 | 68.5599
-6.3665e-12 | 68.7489 | -1.3029e-10 | 68.9649
-4.3201e-12 | 69.155 | -1.3711e-10 | 69.369
-2.9559e-12 | 69.559 | -1.3461e-10 | 69.773
-4.0927e-12 | 69.965 | -1.3915e-10 | 70.177
-5.0022e-12 | 70.369 | -1.0868e-10 | 70.583
-4.5475e-12 | 70.775 | -1.321e-10 | 70.9891
-4.3201e-12 | 71.1791 | -1.4234e-10 | 71.3931
-2.9559e-12 | 71.5851 | -1.3824e-10 | 71.7981
-3.638e-12 | 71.9891 | -1.3733e-10 | 72.2031
-4.3201e-12 | 72.3951 | -1.2869e-10 | 72.6072
-4.3201e-12 | 72.8012 | -1.3029e-10 | 73.0132
-3.1832e-12 | 73.2062 | -1.2892e-10 | 73.4172
-4.0927e-12 | 73.6102 | -1.3097e-10 | 73.8202
-4.3201e-12 | 74.0142 | -1.3097e-10 | 74.2262
-2.2737e-12 | 74.4193 | -1.2483e-10 | 74.6313
-4.0927e-12 | 74.8243 | -1.3051e-10 | 75.0363
-5.457e-12 | 75.2303 | -1.3461e-10 | 75.4403
-3.1832e-12 | 75.6343 | -1.4302e-10 | 75.8453
-6.5938e-12 | 76.0404 | -1.3165e-10 | 76.2514
-7.7307e-12 | 76.4454 | -1.3529e-10 | 76.6554
-4.5475e-12 | 76.8494 | -1.2392e-10 | 77.0614
-2.7285e-12 | 77.2544 | -1.2642e-10 | 77.4654
-5.0022e-12 | 77.6584 | -1.2642e-10 | 77.8684
-5.6843e-12 | 78.0635 | -1.3506e-10 | 78.2735
-3.638e-12 | 78.4695 | -1.3779e-10 | 78.6795
-3.1832e-12 | 78.8755 | -1.2005e-10 | 79.0835
-3.1832e-12 | 79.2805 | -1.3097e-10 | 79.4895
-3.1832e-12 | 79.6846 | -1.3051e-10 | 79.8946
-4.5475e-12 | 80.0896 | -1.1732e-10 | 80.3006
-2.2737e-13 | 80.4946 | -1.2278e-10 | 80.7046
-2.9559e-12 | 80.8996 | -1.2142e-10 | 81.1106
-7.9581e-12 | 81.3026 | -1.3802e-10 | 81.5147
-3.8654e-12 | 81.7067 | -1.271e-10 | 81.9197
-3.8654e-12 | 82.1237 | -1.13e-10 | 82.3357
-5.2296e-12 | 82.5307 | -1.3961e-10 | 82.7407
-1.819e-12 | 82.9347 | -1.2301e-10 | 83.1468
-7.5033e-12 | 83.3398 | -1.3188e-10 | 83.5508
-5.0022e-12 | 83.7458 | -1.2778e-10 | 83.9548
-1.3642e-12 | 84.1498 | -1.3551e-10 | 84.3608
-3.638e-12 | 84.5538 | -1.2915e-10 | 84.7669
-6.1391e-12 | 84.9589
-6.1391e-12 | 85.3649

## 660uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-6.6092e-12 | 2.9071e-12 | -1.8263e-10 | 1.3303e-11
GEMs On |  | GEMs On
-8.1855e-12 | 0.31002 | -1.8849e-10 | 0.31002
-4.3201e-12 | 0.99606 | -1.8053e-10 | 0.99406
-8.6402e-12 | 1.4011 | -1.6394e-10 | 1.4001
-7.9581e-12 | 1.8071 | -2.0555e-10 | 1.8051
0 | 2.2121 | -1.6325e-10 | 2.2101
-2.2737e-11 | 2.6141 | -1.8326e-10 | 2.6151
-8.1855e-12 | 3.0192 | -2.0327e-10 | 3.0202
-8.6402e-12 | 3.4252 | -1.8122e-10 | 3.4242
-3.8654e-12 | 3.8302 | -1.9122e-10 | 3.8292
-3.638e-12 | 4.2372 | -1.9145e-10 | 4.2342
-4.3201e-12 | 4.6423 | -1.7576e-10 | 4.6393
-5.2296e-12 | 5.0463 | -1.7462e-10 | 5.0433
-9.0949e-12 | 5.4513 | -1.5825e-10 | 5.4483
-1.5007e-11 | 5.8573 | -1.9486e-10 | 5.8523
-4.7748e-12 | 6.2614 | -2.01e-10 | 6.2564
-6.8212e-12 | 6.6674 | -1.7099e-10 | 6.6614
-5.2296e-12 | 7.0714 | -1.794e-10 | 7.0654
-5.2296e-12 | 7.4764 | -2.1168e-10 | 7.4714
-8.8676e-12 | 7.8815 | -1.8144e-10 | 7.8755
-1.5916e-12 | 8.2875 | -1.6598e-10 | 8.2805
-7.7307e-12 | 8.6925 | -1.8781e-10 | 8.6845
-7.7307e-12 | 9.0975 | -1.678e-10 | 9.0895
-5.6843e-12 | 9.5035 | -1.7553e-10 | 9.4935
-3.8654e-12 | 9.9076 | -2.0032e-10 | 9.8986
-3.4106e-12 | 10.3126 | -1.8758e-10 | 10.3036
-8.1855e-12 | 10.7166 | -1.9986e-10 | 10.7076
-1.0004e-11 | 11.1226 | -1.6576e-10 | 11.1126
-3.638e-12 | 11.5277 | -2.0123e-10 | 11.5167
-6.8212e-12 | 11.9337 | -1.8758e-10 | 11.9217
-5.2296e-12 | 12.3387 | -1.8122e-10 | 12.3277
-6.8212e-12 | 12.7437 | -1.7894e-10 | 12.7337
-3.4106e-12 | 13.1497 | -1.7758e-10 | 13.1397
-5.0022e-12 | 13.5548 | -1.6462e-10 | 13.5448
-4.0927e-12 | 13.9608 | -1.8849e-10 | 13.9488
-4.3201e-12 | 14.3648 | -1.8099e-10 | 14.3538
-5.2296e-12 | 14.7689 | -1.8144e-10 | 14.7588
-1.1369e-11 | 15.1749 | -1.8395e-10 | 15.1639
-1.819e-12 | 15.5809 | -1.6939e-10 | 15.5689
-3.638e-12 | 15.9859 | -1.8872e-10 | 15.9749
-8.8676e-12 | 16.3909 | -2.1669e-10 | 16.3789
-8.1855e-12 | 16.795 | -1.794e-10 | 16.784
-6.3665e-12 | 17.2 | -1.7849e-10 | 17.188
-5.0022e-12 | 17.605 | -1.6689e-10 | 17.593
-1.1369e-11 | 18.011 | -1.8645e-10 | 18
-6.3665e-12 | 18.416 | -1.9122e-10 | 18.4061
-4.0927e-12 | 18.8211 | -1.7985e-10 | 18.8111
-5.2296e-12 | 19.2261 | -1.7599e-10 | 19.2151
-5.2296e-12 | 19.6311 | -1.9668e-10 | 19.6221
-8.1855e-12 | 20.0361 | -1.678e-10 | 20.0271
-9.3223e-12 | 20.4402 | -1.9304e-10 | 20.4312
-5.2296e-12 | 20.8442 | -1.7303e-10 | 20.8342
-7.276e-12 | 21.2482 | -1.8213e-10 | 21.2382
-1.0004e-11 | 21.6532 | -1.6462e-10 | 21.6422
-7.5033e-12 | 22.0583 | -2.0009e-10 | 22.0473
-5.457e-12 | 22.4633 | -1.6666e-10 | 22.4523
-7.276e-12 | 22.8673 | -1.769e-10 | 22.8593
-5.457e-12 | 23.2723 | -2.0532e-10 | 23.2643
-6.8212e-12 | 23.6783 | -1.9486e-10 | 23.6684
-5.6843e-12 | 24.0834 | -1.7803e-10 | 24.0734
-6.5938e-12 | 24.4884 | -1.7621e-10 | 24.4774
-8.8676e-12 | 24.8934 | -1.6325e-10 | 24.8824
-3.4106e-12 | 25.2975 | -1.6644e-10 | 25.2885
-7.5033e-12 | 25.7015 | -1.7349e-10 | 25.6925
-1.0687e-11 | 26.1065 | -1.8485e-10 | 26.0985
-6.8212e-12 | 26.5125 | -1.7349e-10 | 26.5045
-3.638e-12 | 26.9175 | -1.9622e-10 | 26.9085
-5.0022e-12 | 27.3236 | -1.6416e-10 | 27.3156
-6.3665e-12 | 27.7276 | -1.7872e-10 | 27.7206
-8.6402e-12 | 28.1326 | -1.7758e-10 | 28.1246
-3.638e-12 | 28.5366 | -1.8349e-10 | 28.5296
-4.7748e-12 | 28.9417 | -1.8463e-10 | 28.9347
-3.8654e-12 | 29.3457 | -1.819e-10 | 29.3397
-2.2737e-12 | 29.7517 | -1.8599e-10 | 29.7447
-6.1391e-12 | 30.1557 | -1.7053e-10 | 30.1487
-5.0022e-12 | 30.5598 | -1.985e-10 | 30.5538
-9.7771e-12 | 30.9648 | -1.8622e-10 | 30.9578
-3.8654e-12 | 31.3698 | -2.0123e-10 | 31.3628
-8.6402e-12 | 31.7758 | -1.8827e-10 | 31.7678
-1.0459e-11 | 32.1798 | -1.819e-10 | 32.1718
-7.5033e-12 | 32.5859 | -1.8008e-10 | 32.5779
-2.7285e-12 | 32.9899 | -1.8372e-10 | 32.9829
-4.0927e-12 | 33.3939 | -1.8053e-10 | 33.3879
-7.7307e-12 | 33.7989 | -1.703e-10 | 33.7919
-6.5938e-12 | 34.203 | -1.9509e-10 | 34.196
-4.0927e-12 | 34.608 | -1.8781e-10 | 34.601
-4.3201e-12 | 35.012 | -1.8372e-10 | 35.005
-6.8212e-12 | 35.416 | -1.9554e-10 | 35.412
-6.3665e-12 | 35.821 | -1.9486e-10 | 35.8171
-6.8212e-12 | 36.2261 | -1.9145e-10 | 36.2221
-5.6843e-12 | 36.6311 | -1.7553e-10 | 36.6271
-8.1855e-12 | 37.0351 | -1.5189e-10 | 37.0311
-7.9581e-12 | 37.4411 | -1.8235e-10 | 37.4361
-5.2296e-12 | 37.8462 | -1.9418e-10 | 37.8412
-9.3223e-12 | 38.2502 | -1.7894e-10 | 38.2462
-7.5033e-12 | 38.6552 | -1.7189e-10 | 38.6502
-8.1855e-12 | 39.0592 | -1.5302e-10 | 39.0552
-4.7748e-12 | 39.4643 | -1.7849e-10 | 39.4593
-9.7771e-12 | 39.8683 | -1.8781e-10 | 39.8643
-6.8212e-12 | 40.2733 | -1.9622e-10 | 40.2683
-7.276e-12 | 40.6783 | -1.9736e-10 | 40.6733
-6.8212e-12 | 41.0834 | -1.9804e-10 | 41.0774
-5.2296e-12 | 41.4884 | -1.8372e-10 | 41.4844
-7.5033e-12 | 41.8924 | -1.7121e-10 | 41.8894
-9.7771e-12 | 42.2984 | -1.794e-10 | 42.2934
-5.6843e-12 | 42.7044 | -1.9804e-10 | 42.6984
-6.1391e-12 | 43.1095 | -1.769e-10 | 43.1025
-6.1391e-12 | 43.5145 | -1.5712e-10 | 43.5075
-7.276e-12 | 43.9195 | -1.7349e-10 | 43.9125
-8.6402e-12 | 44.3245 | -2.0032e-10 | 44.3165
-1.0232e-11 | 44.7306 | -1.8645e-10 | 44.7226
-9.3223e-12 | 45.1356 | -1.9395e-10 | 45.1276
-8.6402e-12 | 45.5386 | -1.7553e-10 | 45.5316
-5.2296e-12 | 45.9446 | -1.9509e-10 | 45.9366
-6.5938e-12 | 46.3496 | -1.9668e-10 | 46.3417
-5.457e-12 | 46.7537 | -1.8872e-10 | 46.7457
-5.457e-12 | 47.1587 | -1.703e-10 | 47.1507
-7.276e-12 | 47.5637 | -1.8963e-10 | 47.5557
-7.5033e-12 | 47.9687 | -1.8576e-10 | 47.9617
-5.6843e-12 | 48.3728 | -1.8895e-10 | 48.3658
-5.0022e-12 | 48.7778 | -2.0282e-10 | 48.7708
-4.7748e-12 | 49.1818 | -1.8895e-10 | 49.1768
-8.8676e-12 | 49.5868 | -1.7803e-10 | 49.5838
-7.7307e-12 | 49.9909 | -1.9463e-10 | 49.9879
-1.1369e-12 | 50.3969 | -1.703e-10 | 50.3929
-8.6402e-12 | 50.8019 | -2.01e-10 | 50.7969
-7.9581e-12 | 51.2059 | -1.8144e-10 | 51.2019
-6.3665e-12 | 51.611 | -1.7849e-10 | 51.608
-7.5033e-12 | 52.016 | -1.8917e-10 | 52.014
-7.276e-12 | 52.421 | -1.7599e-10 | 52.42
-8.6402e-12 | 52.826 | -1.8372e-10 | 52.825
-6.8212e-12 | 53.231 | -1.728e-10 | 53.229
-3.8654e-12 | 53.6361 | -1.8963e-10 | 53.6331
-8.8676e-12 | 54.0421 | -1.4552e-10 | 54.0391
-7.276e-12 | 54.4501 | -1.6826e-10 | 54.4441
-5.6843e-12 | 54.8541 | -1.7758e-10 | 54.8491
-6.5938e-12 | 55.2592 | -1.7621e-10 | 55.2552
-9.7771e-12 | 55.6642 | -1.8736e-10 | 55.6612
-4.7748e-12 | 56.0692 | -1.8622e-10 | 56.0652
-5.6843e-12 | 56.4752 | -1.7371e-10 | 56.4712
-7.7307e-12 | 56.8792 | -1.8485e-10 | 56.8772
-5.2296e-12 | 57.2853 | -1.9622e-10 | 57.2813
-2.5011e-12 | 57.6893 | -1.7894e-10 | 57.6843
-6.8212e-12 | 58.0943 | -1.7803e-10 | 58.0903
-6.3665e-12 | 58.4993 | -1.6439e-10 | 58.4954
-7.9581e-12 | 58.9034 | -1.8827e-10 | 58.8994
-6.5938e-12 | 59.3064 | -1.9077e-10 | 59.3034
-6.8212e-12 | 59.7124 | -1.6917e-10 | 59.7074
-6.5938e-12 | 60.1164 | -1.8485e-10 | 60.1124
-1.1369e-11 | 60.5225 | -1.6598e-10 | 60.5175
-5.457e-12 | 60.9295 | -1.6462e-10 | 60.9225
-5.6843e-12 | 61.3355 | -1.7485e-10 | 61.3265
-3.8654e-12 | 61.7395 | -1.9668e-10 | 61.7315
-7.5033e-12 | 62.1476 | -2.0282e-10 | 62.1365
-6.3665e-12 | 62.5516 | -1.8372e-10 | 62.5416
-5.6843e-12 | 62.9576 | -1.8463e-10 | 62.9456
-4.3201e-12 | 63.3616 | -1.6917e-10 | 63.3496
-5.6843e-12 | 63.7687 | -1.6803e-10 | 63.7557
-8.1855e-12 | 64.1727 | -1.5484e-10 | 64.1587
-6.5938e-12 | 64.5767 | -2.1328e-10 | 64.5637
-3.638e-12 | 64.9817 | -1.9941e-10 | 64.9667
-5.6843e-12 | 65.3877 | -1.8986e-10 | 65.3717
-9.0949e-12 | 65.7918 | -1.9941e-10 | 65.7778
-9.3223e-12 | 66.1958 | -1.7462e-10 | 66.1818
-5.457e-12 | 66.6018 | -1.7621e-10 | 66.5868
-5.6843e-12 | 67.0068 | -1.9509e-10 | 66.9908
-9.3223e-12 | 67.4129 | -1.9236e-10 | 67.3949
-2.5011e-12 | 67.8179 | -1.6303e-10 | 67.7999
-7.276e-12 | 68.2219 | -1.7872e-10 | 68.2049
-1.0004e-11 | 68.6259 | -1.8622e-10 | 68.6099
-7.9581e-12 | 69.0329 | -1.9531e-10 | 69.0149
-6.3665e-12 | 69.438 | -1.7508e-10 | 69.419
-5.457e-12 | 69.843 | -1.8917e-10 | 69.824
-6.3665e-12 | 70.248 | -1.5575e-10 | 70.229
-7.7307e-12 | 70.652 | -1.8645e-10 | 70.635
-4.0927e-12 | 71.0561 | -2.1805e-10 | 71.0391
-4.0927e-12 | 71.4601 | -1.678e-10 | 71.4431
-8.6402e-12 | 71.8641 | -1.7485e-10 | 71.8481
-5.2296e-12 | 72.2691 | -1.7485e-10 | 72.2531
-5.6843e-12 | 72.6752 | -1.7667e-10 | 72.6572
-5.2296e-12 | 73.0792 | -2.0646e-10 | 73.0622
-7.7307e-12 | 73.4822 | -1.678e-10 | 73.4672
-7.9581e-12 | 73.8862 | -1.8827e-10 | 73.8732
-7.5033e-12 | 74.2913 | -1.8736e-10 | 74.2793
-6.1391e-12 | 74.6973 | -1.9782e-10 | 74.6843
-5.457e-12 | 75.1023 | -1.9759e-10 | 75.0873
-7.5033e-12 | 75.5063 | -2.0304e-10 | 75.4923
-8.1855e-12 | 75.9103 | -1.578e-10 | 75.8973
-6.5938e-12 | 76.3154 | -1.7985e-10 | 76.3034
-6.5938e-12 | 76.7214 | -1.7371e-10 | 76.7084
-8.8676e-12 | 77.1264 | -1.935e-10 | 77.1134
-3.638e-12 | 77.5304 | -1.9918e-10 | 77.5194
-4.3201e-12 | 77.9345 | -1.9259e-10 | 77.9235
-8.6402e-12 | 78.3395 | -1.5643e-10 | 78.3285
-7.7307e-12 | 78.7435 | -1.9645e-10 | 78.7335
-7.276e-12 | 79.1495 | -1.985e-10 | 79.1385
-5.457e-12 | 79.5525 | -1.8872e-10 | 79.5435
-8.8676e-12 | 79.9586 | -2.1737e-10 | 79.9486
-1.0232e-11 | 80.3636 | -1.6485e-10 | 80.3536
-7.9581e-12 | 80.7696 | -1.8372e-10 | 80.7586
-6.8212e-12 | 81.1746 | -1.5871e-10 | 81.1626
-8.8676e-12 | 81.5787 | -1.5302e-10 | 81.5687
-7.9581e-12 | 81.9937 | -1.8053e-10 | 81.9827
-6.3665e-12 | 82.3987 | -1.9145e-10 | 82.3877
-1.0232e-11 | 82.8037 | -1.7735e-10 | 82.7937
-4.0927e-12 | 83.2098 | -1.8099e-10 | 83.1988
-6.1391e-12 | 83.6138 | -1.728e-10 | 83.6038
-8.8676e-12 | 84.0188 | -1.7963e-10 | 84.0088
-7.5033e-12 | 84.4238 | -1.7667e-10 | 84.4128
-5.0022e-12 | 84.8289 | -1.7894e-10 | 84.8179
-6.3665e-12 | 85.2349 | -1.8827e-10 | 85.2229
-4.7748e-12 | 85.6399 | -1.7462e-10 | 85.6279
-7.9581e-12 | 86.0439
-5.6843e-12 | 86.4509
-8.8676e-12 | 86.855
-3.4106e-12 | 87.261
-4.3201e-12 | 87.666
-9.7771e-12 | 88.071
-9.0949e-12 | 88.4771
-7.7307e-12 | 88.8831
-3.4106e-12 | 89.2891
-6.8212e-12 | 89.6951
-7.276e-12 | 90.0991
-9.0949e-12 | 90.5042
-8.8676e-12 | 90.9092
-8.8676e-12 | 91.3142
-8.8676e-12 | 91.7193
-5.6843e-12 | 92.1233
-7.9581e-12 | 92.5283
-6.5938e-12 | 92.9333
-3.638e-12 | 93.3393
-7.5033e-12 | 93.7434
-9.3223e-12 | 94.1484
-9.0949e-12 | 94.5534
-4.7748e-12 | 94.9584
-8.6402e-12 | 95.3635
-1.0232e-11 | 95.7675
-5.2296e-12 | 96.1745
-6.3665e-12 | 96.5785
-3.638e-12 | 96.9845
-9.0949e-12 | 97.3886
-5.2296e-12 | 97.7926
-8.8676e-12 | 98.1996
-5.457e-12 | 98.6036
-1.0004e-11 | 99.0087
-8.1855e-12 | 99.4137
-7.276e-12 | 99.8177
-6.8212e-12 | 100.2227
-5.2296e-12 | 100.6298
-5.6843e-12 | 101.0348
-6.8212e-12 | 101.4398
-8.8676e-12 | 101.8448
-6.8212e-12 | 102.2508
-7.276e-12 | 102.6549
-5.6843e-12 | 103.0599
-5.6843e-12 | 103.4639
-8.1855e-12 | 103.8699
-5.6843e-12 | 104.275
-6.3665e-12 | 104.68
-8.1855e-12 | 105.085
-5.0022e-12 | 105.49
-7.5033e-12 | 105.8941
-3.4106e-12 | 106.2981
-8.1855e-12 | 106.7031
-7.5033e-12 | 107.1091
-5.2296e-12 | 107.5141
-7.276e-12 | 107.9192
-4.0927e-12 | 108.3242
-6.8212e-12 | 108.7282
-2.9559e-12 | 109.1332
-9.7771e-12 | 109.5383
-5.6843e-12 | 109.9433
-6.3665e-12 | 110.3483
-7.5033e-12 | 110.7543
-9.0949e-12 | 111.1594
-6.1391e-12 | 111.5644
-5.0022e-12 | 111.9704
-7.7307e-12 | 112.3744
-1.0687e-11 | 112.7795
-2.9559e-12 | 113.1855
-7.7307e-12 | 113.5905
-6.3665e-12 | 113.9955
-1.1369e-11 | 114.4005
-7.7307e-12 | 114.8056
-5.6843e-12 | 115.2116
-5.0022e-12 | 115.6156
-7.9581e-12 | 116.0206
-8.1855e-12 | 116.4257
-6.8212e-12 | 116.8297
-8.6402e-12 | 117.2347
-9.3223e-12 | 117.6387
-3.638e-12 | 118.0438
-5.2296e-12 | 118.4488
-4.0927e-12 | 118.8528
-4.3201e-12 | 119.2588
-6.8212e-12 | 119.6638
-6.3665e-12 | 120.0689
-5.2296e-12 | 120.4749
-6.1391e-12 | 120.8799
-6.5938e-12 | 121.2849
-3.8654e-12 | 121.69
-7.9581e-12 | 122.094
-6.1391e-12 | 122.5
-6.3665e-12 | 122.903
-4.3201e-12 | 123.3091
-7.276e-12 | 123.7151
-6.3665e-12 | 124.1201
-6.1391e-12 | 124.5251
-7.276e-12 | 124.9311
-1.1369e-11 | 125.3352
-8.6402e-12 | 125.7402
2.3874e-11 | 126.1462
-3.638e-12 | 126.5512
-7.5033e-12 | 126.9563
-6.1391e-12 | 127.3623
-7.5033e-12 | 127.7663
-8.1855e-12 | 128.1713
-6.8212e-12 | 128.5764
-7.9581e-12 | 128.9814
-8.1855e-12 | 129.3864
5.9117e-12 | 129.7914
-4.7748e-12 | 130.1964
-1.0459e-11 | 130.6015
-7.9581e-12 | 131.0075
-9.7771e-12 | 131.4115
-8.6402e-12 | 131.8175
-6.1391e-12 | 132.2226

## 670uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-6.9893e-12 | 2.0441e-12 | -2.6041e-10 | 2.0503e-11
GEMs On |  | GEMs On
-5.6843e-12 | 0.31202 | -2.7035e-10 | 0.31002
-3.1832e-12 | 1.0081 | -2.8194e-10 | 0.99506
-5.2296e-12 | 1.4131 | -3.0423e-10 | 1.4011
-1.2733e-11 | 1.8191 | -2.658e-10 | 1.8061
-4.3201e-12 | 2.2231 | -2.4534e-10 | 2.2101
-3.8654e-12 | 2.6282 | -2.5079e-10 | 2.6202
-5.457e-12 | 3.0332 | -2.8081e-10 | 3.0262
-7.0486e-12 | 3.4372 | -2.733e-10 | 3.4302
-9.3223e-12 | 3.8422 | -2.8058e-10 | 3.8342
-1.0232e-11 | 4.2482 | -2.6785e-10 | 4.2382
-6.5938e-12 | 4.6533 | -2.5784e-10 | 4.6433
-5.0022e-12 | 5.0583 | -2.9399e-10 | 5.0493
-7.7307e-12 | 5.4633 | -3.04e-10 | 5.4543
-6.8212e-12 | 5.8683 | -2.276e-10 | 5.8593
-2.7285e-12 | 6.2744 | -2.1805e-10 | 6.2654
-3.638e-12 | 6.6784 | -2.8763e-10 | 6.6694
-8.1855e-12 | 7.0844 | -2.5443e-10 | 7.0764
-9.5497e-12 | 7.4894 | -2.8604e-10 | 7.4804
-5.457e-12 | 7.8945 | -2.3783e-10 | 7.8845
-5.2296e-12 | 8.2995 | -3.0423e-10 | 8.2895
-2.9559e-12 | 8.7045 | -2.7057e-10 | 8.6945
-6.3665e-12 | 9.1095 | -2.5148e-10 | 9.0995
-6.5938e-12 | 9.5135 | -2.4261e-10 | 9.5025
-7.7307e-12 | 9.9186 | -2.7944e-10 | 9.9066
-6.8212e-12 | 10.3306 | -2.5238e-10 | 10.3126
-7.5033e-12 | 10.7346 | -2.4215e-10 | 10.7166
-7.5033e-12 | 11.1386 | -3.0627e-10 | 11.1206
-5.0022e-12 | 11.5447 | -2.4602e-10 | 11.5247
-5.2296e-12 | 11.9497 | -2.3738e-10 | 11.9317
-1.0459e-11 | 12.3537 | -2.6262e-10 | 12.3367
-5.2296e-12 | 12.7597 | -2.4693e-10 | 12.7407
-3.638e-12 | 13.1647 | -2.6262e-10 | 13.1457
-7.0486e-12 | 13.5698 | -2.5784e-10 | 13.5498
-3.8654e-12 | 13.9738 | -2.7899e-10 | 13.9548
-9.3223e-12 | 14.3778 | -2.542e-10 | 14.3598
-6.3665e-12 | 14.7828 | -2.6807e-10 | 14.7638
-9.5497e-12 | 15.1879 | -2.9945e-10 | 15.1689
-1.0004e-11 | 15.5939 | -2.6694e-10 | 15.5739
-7.0486e-12 | 15.9999 | -2.5307e-10 | 15.9779
-5.9117e-12 | 16.4039 | -2.4124e-10 | 16.3839
-7.9581e-12 | 16.81 | -2.4511e-10 | 16.788
-5.0022e-12 | 17.215 | -2.7035e-10 | 17.191
-7.0486e-12 | 17.619 | -2.4374e-10 | 17.598
-4.5475e-12 | 18.024 | -2.4352e-10 | 18.003
-7.5033e-12 | 18.4291 | -2.4602e-10 | 18.4091
-1.0687e-11 | 18.8341 | -2.2851e-10 | 18.8141
-5.0022e-12 | 19.2381 | -2.6284e-10 | 19.2191
-5.0022e-12 | 19.6441 | -2.6557e-10 | 19.6241
-8.1855e-12 | 20.0492 | -2.3783e-10 | 20.0282
-1.0459e-11 | 20.4542 | -2.7035e-10 | 20.4372
-8.4128e-12 | 20.8592 | -2.5057e-10 | 20.8782
-6.5938e-12 | 21.2642 | -2.4602e-10 | 21.2872
-8.1855e-12 | 21.6692 | -2.326e-10 | 21.6912
-8.1855e-12 | 22.0743 | -2.2874e-10 | 22.0983
-5.6843e-12 | 22.4793 | -2.5489e-10 | 22.5023
-9.3223e-12 | 22.8853 | -2.542e-10 | 22.9073
-1.0004e-11 | 23.2893 | -2.4215e-10 | 23.3123
-7.5033e-12 | 23.6954 | -2.7217e-10 | 23.7184
-6.3665e-12 | 24.1004 | -2.9149e-10 | 24.1244
-7.0486e-12 | 24.5054 | -2.5148e-10 | 24.5304
-9.0949e-12 | 24.9104 | -2.799e-10 | 24.9374
-7.7307e-12 | 25.3154 | -2.4215e-10 | 25.3424
-5.2296e-12 | 25.7205 | -2.4943e-10 | 25.7475
-1.2733e-11 | 26.1255 | -2.658e-10 | 26.1615
-1.387e-11 | 26.5305 | -2.7831e-10 | 26.5725
-5.2296e-12 | 26.9355 | -3.0286e-10 | 26.9805
-4.3201e-12 | 27.3406 | -2.7126e-10 | 27.3846
-7.0486e-12 | 27.7456 | -2.6102e-10 | 27.7906
-9.0949e-12 | 28.1496 | -2.6353e-10 | 28.1946
-6.3665e-12 | 28.5556 | -2.4693e-10 | 28.5996
-6.5938e-12 | 28.9607 | -2.9263e-10 | 29.0047
-5.457e-12 | 29.3657 | -2.542e-10 | 29.4077
-7.9581e-12 | 29.7717 | -2.7808e-10 | 29.8137
-3.638e-12 | 30.1777 | -2.5375e-10 | 30.2177
-8.1855e-12 | 30.5817 | -2.5943e-10 | 30.6238
-9.0949e-12 | 30.9878 | -2.567e-10 | 31.0288
-7.5033e-12 | 31.3928 | -2.633e-10 | 31.4348
-5.0022e-12 | 31.7988 | -2.758e-10 | 31.8398
-7.9581e-12 | 32.2038 | -2.6262e-10 | 32.2458
-5.6843e-12 | 32.6089 | -2.6353e-10 | 32.6509
-1.1823e-11 | 33.0139 | -2.6284e-10 | 33.0559
-7.5033e-12 | 33.4189 | -2.1942e-10 | 33.4599
-6.5938e-12 | 33.8239 | -2.8194e-10 | 33.8649
-2.9559e-12 | 34.23 | -2.392e-10 | 34.27
-5.6843e-12 | 34.635 | -2.9104e-10 | 34.675
-6.8212e-12 | 35.04 | -2.2874e-10 | 35.081
-5.6843e-12 | 35.447 | -2.7831e-10 | 35.485
-5.0022e-12 | 35.852 | -2.7126e-10 | 35.8901
-5.6843e-12 | 36.2581 | -2.4033e-10 | 36.7591
-8.4128e-12 | 36.6631 | -2.5761e-10 | 37.1651
-7.5033e-12 | 37.0681 | -2.5921e-10 | 37.5731
-8.4128e-12 | 37.4731 | -2.5716e-10 | 37.9832
-7.9581e-12 | 37.8782 | -2.326e-10 | 38.6302
-1.1823e-11 | 38.2822 | -2.7944e-10 | 39.4102
-5.0022e-12 | 38.6882 | -2.9172e-10 | 39.8163
-4.5475e-12 | 39.0942 | -2.1532e-10 | 40.2213
-5.6843e-12 | 39.4983 | -2.5307e-10 | 40.6263
-7.5033e-12 | 39.9023 | -2.3579e-10 | 41.0314
-5.6843e-12 | 40.3073 | -2.7831e-10 | 41.4364
-3.638e-12 | 40.7133 | -2.6557e-10 | 41.8924
-7.9581e-12 | 41.1174 | -2.4011e-10 | 42.3714
-8.1855e-12 | 41.5214 | -2.5921e-10 | 42.7775
-6.3665e-12 | 41.9264 | -2.2965e-10 | 43.1825
-6.3665e-12 | 42.3324 | -2.4488e-10 | 43.6095
-6.8212e-12 | 42.7384 | -2.8217e-10 | 44.1195
-7.5033e-12 | 43.1435 | -2.8058e-10 | 44.6495
-5.2296e-12 | 43.5485 | -2.4215e-10 | 45.0656
-5.6843e-12 | 43.9525 | -2.4147e-10 | 45.4766
-5.9117e-12 | 44.3565 | -2.6876e-10 | 45.8826
-4.5475e-12 | 44.7616 | -2.658e-10 | 46.2886
-5.0022e-12 | 45.1676 | -2.1191e-10 | 46.6927
-4.5475e-12 | 45.5726 | -2.5557e-10 | 47.0977
-7.7307e-12 | 45.9776 | -2.7671e-10 | 47.5017
-7.0486e-12 | 46.3826 | -2.4352e-10 | 47.9057
-6.8212e-12 | 46.7867 | -2.3601e-10 | 48.3118
-8.1855e-12 | 47.1917 | -2.6102e-10 | 48.7178
-7.9581e-12 | 47.5977 | -2.3442e-10 | 49.1268
-4.0927e-12 | 48.0027 | -2.9399e-10 | 49.5338
-4.5475e-12 | 48.4068 | -2.6989e-10 | 49.9489
-9.0949e-12 | 48.8118 | -2.9127e-10 | 50.3869
-1.0687e-11 | 49.2158 | -2.583e-10 | 50.7919
-6.8212e-12 | 49.6218 | -2.3851e-10 | 51.1979
-5.0022e-12 | 50.0269 | -2.5875e-10 | 51.6029
-5.6843e-12 | 50.4319 | -2.4215e-10 | 52.013
-7.5033e-12 | 50.8369 | -2.7467e-10 | 52.419
-7.0486e-12 | 51.2429 | -2.5534e-10 | 52.829
-4.5475e-12 | 51.6489 | -2.558e-10 | 53.2341
-1.3188e-11 | 52.053 | -2.5034e-10 | 53.6391
-1.0004e-11 | 52.458 | -2.8331e-10 | 54.0461
-6.5938e-12 | 52.864 | -2.9081e-10 | 54.4541
-9.0949e-12 | 53.2691 | -2.7808e-10 | 54.8591
-1.0459e-11 | 53.6731 | -2.5193e-10 | 55.2642
-7.7307e-12 | 54.0781 | -2.3533e-10 | 55.6692
-9.5497e-12 | 54.4841 | -2.608e-10 | 56.0822
-9.3223e-12 | 54.8891 | -2.7057e-10 | 56.4872
-7.5033e-12 | 55.2942 | -2.7308e-10 | 56.8972
-7.7307e-12 | 55.6992 | -2.7603e-10 | 57.3033
-6.5938e-12 | 56.1032 | -2.6262e-10 | 57.7073
-7.7307e-12 | 56.5082 | -2.9627e-10 | 58.1333
-9.3223e-12 | 56.9133 | -2.1055e-10 | 58.5513
-5.6843e-12 | 57.3183 | -2.5238e-10 | 59.0334
-6.8212e-12 | 57.7243 | -2.658e-10 | 59.4404
-6.8212e-12 | 58.1283 | -2.3533e-10 | 59.8464
-6.5938e-12 | 58.5344 | -2.6489e-10 | 60.2525
-5.0022e-12 | 58.9384 | -2.8126e-10 | 60.6595
-8.8676e-12 | 59.3424 | -2.824e-10 | 61.0655
-6.8212e-12 | 59.7474 | -2.3101e-10 | 61.4835
-6.8212e-12 | 60.1514 | -2.7899e-10 | 61.8985
-5.457e-12 | 60.5555 | -2.3647e-10 | 62.3076
-6.5938e-12 | 60.9605 | -2.6216e-10 | 62.7146
-7.5033e-12 | 61.3665 | -2.5489e-10 | 63.1196
-4.3201e-12 | 61.7715 | -2.4306e-10 | 63.5326
-6.3665e-12 | 62.1766 | -2.6785e-10 | 63.9397
-7.7307e-12 | 62.5816 | -2.4306e-10 | 64.3647
-1.0232e-11 | 62.9866 | -2.6535e-10 | 64.7707
-6.8212e-12 | 63.3936 | -2.7558e-10 | 65.1747
-5.6843e-12 | 63.7976 | -2.924e-10 | 65.6088
-8.4128e-12 | 64.2037 | -2.2874e-10 | 66.0158
-6.8212e-12 | 64.6087 | -2.7057e-10 | 66.4208
-6.5938e-12 | 65.0127 | -2.9695e-10 | 66.8408
-1.1596e-11 | 65.4177 | -2.8786e-10 | 67.2459
-9.5497e-12 | 65.8248 | -2.7239e-10 | 67.6529
-5.6843e-12 | 66.2278 | -2.4511e-10 | 68.0579
-6.5938e-12 | 66.6328 | -2.2328e-10 | 68.4639
-7.9581e-12 | 67.0358 | -2.658e-10 | 69.0619
-5.6843e-12 | 67.4409 | -2.8103e-10 | 69.575
-7.7307e-12 | 67.8469 | -2.7035e-10 | 70.013
-4.0927e-12 | 68.2529 | -2.5943e-10 | 70.418
-6.5938e-12 | 68.6569 | -2.7626e-10 | 71.2421
-1.0687e-11 | 69.0619 | -2.6216e-10 | 71.6501
-9.0949e-12 | 69.467 | -2.8581e-10 | 72.0601
-6.3665e-12 | 69.872 | -2.7831e-10 | 72.4691
-1.0459e-11 | 70.277 | -2.467e-10 | 72.9372
-7.7307e-12 | 70.683 | -2.5329e-10 | 73.3942
-4.3201e-12 | 71.0891 | -2.2874e-10 | 73.7992
-7.9581e-12 | 71.4931 | -2.4238e-10 | 74.2042
-5.9117e-12 | 71.8971 | -2.8353e-10 | 74.6103
-4.5475e-12 | 72.3021 | -2.4897e-10 | 75.0183
-7.0486e-12 | 72.7062 | -2.8831e-10 | 75.4233
-7.7307e-12 | 73.1122 | -2.758e-10 | 75.9754
-5.6843e-12 | 73.5172 | -2.3056e-10 | 76.4674
-6.3665e-12 | 73.9222 | -2.658e-10 | 76.8884
-4.3201e-12 | 74.3273 | -2.6807e-10 | 77.2934
-7.0486e-12 | 74.7323 | -2.6694e-10 | 77.6994
-7.5033e-12 | 75.1373 | -2.3192e-10 | 78.1045
-7.9581e-12 | 75.5423 | -2.3215e-10 | 78.5095
-7.9581e-12 | 75.9483 | -2.7103e-10 | 78.9145
-6.3665e-12 | 76.3534 | -2.558e-10 | 79.3195
-5.2296e-12 | 76.7574 | -2.7512e-10 | 79.7256
-4.5475e-12 | 77.1624 | -2.5648e-10 | 80.1296
-4.0927e-12 | 77.5664 | -2.6967e-10 | 80.5346
-5.0022e-12 | 77.9725 | -2.4374e-10 | 80.9396
-1.0232e-11 | 78.3785 | -2.4988e-10 | 81.3436
-7.0486e-12 | 78.7835 | -2.2715e-10 | 81.7487
-8.8676e-12 | 79.1885 | -2.6148e-10 | 82.1517
-5.457e-12 | 79.5935 | -3.0036e-10 | 82.5577
-6.5938e-12 | 79.9996 | -2.7512e-10 | 82.9638
-7.5033e-12 | 80.4056 | -2.758e-10 | 83.3688
-7.7307e-12 | 80.8106 | -2.4352e-10 | 83.7738
-6.3665e-12 | 81.2156 | -2.8194e-10 | 84.1768
-7.0486e-12 | 81.6217 | -2.4306e-10 | 84.5828
-5.6843e-12 | 82.0387 | -2.7171e-10 | 84.9979
-5.6843e-12 | 82.4447 | -2.5966e-10 | 85.4039
-7.5033e-12 | 82.8497 | -2.2192e-10 | 85.8089
-6.8212e-12 | 83.2558 | -2.9013e-10 | 86.2139
-5.457e-12 | 83.6628 | -2.6034e-10 | 86.617
-9.0949e-12 | 84.0678 | -2.4102e-10 | 87.021
|  | -2.2624e-10 | 87.425

## 680uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-9.7287e-12 | 2.5061e-12 | -3.669e-10 | 2.7818e-11
GEMs On |  | GEMs On
-1.0232e-11 | 0.31202 | -3.5379e-10 | 0.31202
-1.2506e-11 | 0.99706 | -3.3197e-10 | 0.99506
-1.1141e-11 | 1.4041 | -4.095e-10 | 1.4001
-1.7735e-11 | 1.8081 | -3.5789e-10 | 1.8051
-6.3665e-12 | 2.2151 | -3.4652e-10 | 2.2241
-1.2506e-11 | 2.6202 | -3.5607e-10 | 2.6312
-7.0486e-12 | 3.0722 | -3.5379e-10 | 3.0362
-7.9581e-12 | 3.4782 | -3.4629e-10 | 3.4422
-4.3201e-12 | 3.8832 | -3.331e-10 | 3.8472
-1.0232e-11 | 4.2872 | -3.4629e-10 | 4.2512
-9.0949e-12 | 4.6923 | -3.331e-10 | 4.6573
-1.0459e-11 | 5.0983 | -3.7107e-10 | 5.0633
-1.0687e-11 | 5.5043 | -3.3697e-10 | 5.4683
-7.7307e-12 | 5.9093 | -4.0313e-10 | 5.8723
-6.1391e-12 | 6.3134 | -3.7744e-10 | 6.2754
-7.5033e-12 | 6.7194 | -3.9267e-10 | 6.6994
-1.2051e-11 | 7.1244 | -3.3651e-10 | 7.1424
-7.5033e-12 | 7.5294 | -3.4765e-10 | 7.5544
-9.5497e-12 | 7.9335 | -3.2287e-10 | 7.9605
-5.0022e-12 | 8.3385 | -3.2924e-10 | 8.3665
-9.5497e-12 | 8.7435 | -3.8199e-10 | 8.9255
-8.6402e-12 | 9.1475 | -3.6812e-10 | 9.3485
-6.5938e-12 | 9.5525 | -3.4515e-10 | 9.7536
-1.0459e-11 | 9.9586 | -4.1814e-10 | 10.1576
-1.1596e-11 | 10.3636 | -3.2196e-10 | 10.5626
-6.3665e-12 | 10.7686 | -4.2041e-10 | 10.9676
-8.1855e-12 | 11.1746 | -3.8085e-10 | 11.3727
-8.1855e-12 | 11.5857 | -3.1946e-10 | 11.7787
-1.0232e-11 | 11.9917 | -3.3469e-10 | 12.1827
-6.1391e-12 | 12.3987 | -3.7585e-10 | 12.5887
-8.1855e-12 | 12.8087 | -3.4606e-10 | 12.9947
-9.5497e-12 | 13.2168 | -4.0563e-10 | 13.4008
-6.3665e-12 | 13.6228 | -3.9086e-10 | 13.8068
-6.8212e-12 | 14.0308 | -3.7448e-10 | 14.2878
-7.9581e-12 | 14.4358 | -3.6721e-10 | 14.8079
-1.0232e-11 | 14.8408 | -3.7039e-10 | 15.2129
-7.0486e-12 | 15.2499 | -3.181e-10 | 15.6159
-1.1823e-11 | 15.6549 | -3.6675e-10 | 16.0199
-7.5033e-12 | 16.0609 | -3.231e-10 | 16.4259
-8.1855e-12 | 16.4989 | -3.9722e-10 | 16.831
-1.0459e-11 | 16.917 | -3.4333e-10 | 17.236
-8.8676e-12 | 17.323 | -3.7926e-10 | 17.641
-6.1391e-12 | 17.747 | -3.5061e-10 | 18.045
-1.1369e-11 | 18.192 | -3.904e-10 | 18.4501
-1.0004e-11 | 18.5981 | -4.0245e-10 | 18.8551
-9.0949e-12 | 19.0051 | -3.9654e-10 | 19.2591
-1.1141e-11 | 19.4151 | -3.6221e-10 | 19.6641
-1.1596e-11 | 19.8211 | -3.256e-10 | 20.0681
-1.0232e-11 | 20.2262 | -3.8813e-10 | 20.4722
-9.3223e-12 | 20.6312 | -4.3951e-10 | 20.8772
-1.0004e-11 | 21.0372 | -3.4333e-10 | 21.2832
-9.3223e-12 | 21.4422 | -3.1127e-10 | 21.6902
-6.3665e-12 | 21.8472 | -3.8654e-10 | 22.0973
-1.728e-11 | 22.2533 | -3.2787e-10 | 22.5023
-1.0004e-11 | 22.6583 | -3.7608e-10 | 22.9083
-5.6843e-12 | 23.0653 | -3.5561e-10 | 23.3133
-8.6402e-12 | 23.4713 | -3.2401e-10 | 23.7184
-1.296e-11 | 23.8754 | -3.804e-10 | 24.1234
-8.1855e-12 | 24.2814 | -4.1769e-10 | 24.5284
-1.1596e-11 | 24.6874 | -3.6221e-10 | 24.9354
-1.0687e-11 | 25.0934 | -3.6289e-10 | 25.3405
-9.3223e-12 | 25.4985 | -4.2223e-10 | 25.7525
-9.3223e-12 | 25.9025 | -3.5607e-10 | 26.1575
-1.2051e-11 | 26.3085 | -3.3697e-10 | 26.5615
-9.0949e-12 | 26.7145 | -3.7312e-10 | 26.9675
-1.387e-11 | 27.1185 | -3.6562e-10 | 27.3726
-1.0687e-11 | 27.5236 | -3.5948e-10 | 27.7766
-1.0687e-11 | 27.9286 | -3.8085e-10 | 28.1826
-1.728e-11 | 28.3336 | -3.5539e-10 | 28.5876
-1.0004e-11 | 28.7386 | -3.3651e-10 | 28.9907
-7.7307e-12 | 29.1437 | -3.6471e-10 | 29.3967
-7.0486e-12 | 29.5477 | -3.9358e-10 | 29.8017
-9.5497e-12 | 29.9547 | -4.4633e-10 | 30.2067
-1.0459e-11 | 30.3597 | -3.9086e-10 | 30.6107
-1.5007e-11 | 30.7658 | -4.0745e-10 | 31.0158
-7.0486e-12 | 31.1708 | -3.7812e-10 | 31.4198
-4.3201e-12 | 31.5788 | -3.8358e-10 | 31.8258
-9.0949e-12 | 31.9838 | -3.5652e-10 | 32.2308
-8.6402e-12 | 32.3899 | -3.5629e-10 | 32.6359
-1.0232e-11 | 32.7959 | -3.9381e-10 | 33.0409
-1.1141e-11 | 33.2039 | -3.0241e-10 | 33.4449
-7.5033e-12 | 33.7189 | -4.0882e-10 | 33.8499
-6.1391e-12 | 34.1279 | -3.7971e-10 | 34.254
-1.2051e-11 | 34.533 | -3.7016e-10 | 34.658
-9.3223e-12 | 34.939 | -3.9131e-10 | 35.062
-1.1369e-11 | 35.344 | -3.5561e-10 | 35.467
-7.7307e-12 | 35.76 | -3.7153e-10 | 35.8721
-1.0459e-11 | 36.1661 | -3.4038e-10 | 36.2771
-9.0949e-12 | 36.5711 | -3.9859e-10 | 36.6821
-1.1141e-11 | 36.9761 | -3.6516e-10 | 37.0861
-6.5938e-12 | 37.3811 | -3.9654e-10 | 37.4921
-1.387e-11 | 37.7862 | -3.8449e-10 | 37.8972
-7.9581e-12 | 38.1992 | -3.8381e-10 | 38.3022
-1.1596e-11 | 38.6052 | -3.4538e-10 | 38.7122
-1.0459e-11 | 39.0102 | -3.5448e-10 | 39.1212
-1.3188e-11 | 39.4203 | -3.6289e-10 | 39.5263
-1.0459e-11 | 39.8293 | -3.8813e-10 | 39.9323
-1.1141e-11 | 40.2353 | -3.7153e-10 | 40.3363
-9.5497e-12 | 40.6413 | -4.1427e-10 | 40.7413
-1.1823e-11 | 41.0463 | -3.6175e-10 | 41.1474
-1.0459e-11 | 41.4514 | -3.8654e-10 | 41.5524
-9.5497e-12 | 41.8564 | -3.9267e-10 | 41.9574
-7.0486e-12 | 42.2614 | -3.7062e-10 | 42.3634
-1.1823e-11 | 42.6664 | -3.9017e-10 | 42.7674
-1.0687e-11 | 43.0715 | -4.1928e-10 | 43.1735
-1.1141e-11 | 43.4765 | -3.4879e-10 | 43.5785
-6.3665e-12 | 43.8805 | -3.4083e-10 | 43.9825
-7.9581e-12 | 44.2855 | -3.9222e-10 | 44.3895
-9.5497e-12 | 44.6896 | -3.5061e-10 | 44.7946
-1.1369e-11 | 45.0946 | -3.7198e-10 | 45.1986
-7.0486e-12 | 45.5016 | -3.7289e-10 | 45.6036
-6.8212e-12 | 45.9056 | -3.7608e-10 | 46.0086
-6.8212e-12 | 46.3096 | -3.5266e-10 | 46.4126
-5.0022e-12 | 46.7147 | -3.2173e-10 | 46.8187
-1.1596e-11 | 47.1187 | -3.4106e-10 | 47.2237
-6.8212e-12 | 47.5247 | -3.9267e-10 | 47.6297
-9.0949e-12 | 47.9307 | -3.8608e-10 | 48.0337
-1.4325e-11 | 48.3368 | -4.045e-10 | 48.4388
-1.4325e-11 | 48.7418 | -3.4902e-10 | 48.8448
-1.2051e-11 | 49.1478 | -3.8517e-10 | 49.2498
-1.1141e-11 | 49.5538 | -3.5834e-10 | 49.6548
-1.1823e-11 | 49.9589 | -3.4811e-10 | 50.0629
-8.6402e-12 | 50.3659 | -4.1473e-10 | 50.4669
-7.7307e-12 | 50.7729 | -3.6675e-10 | 50.8719
-1.2506e-11 | 51.1789 | -3.472e-10 | 51.2769
-1.1823e-11 | 51.584 | -3.9108e-10 | 51.682
-1.1141e-11 | 51.989 | -3.8835e-10 | 52.086
-1.4097e-11 | 52.394 | -3.5243e-10 | 52.492
-7.9581e-12 | 52.799 | -3.065e-10 | 52.898
-9.0949e-12 | 53.204 | -3.8631e-10 | 53.304
-1.0687e-11 | 53.6091 | -3.4811e-10 | 53.7111
-6.3665e-12 | 54.0151 | -3.4538e-10 | 54.1151
-8.1855e-12 | 54.4191 | -3.5448e-10 | 54.5191
-1.4325e-11 | 54.8241 | -3.4584e-10 | 54.9231
-9.0949e-12 | 55.2282 | -3.8949e-10 | 55.3272
-6.3665e-12 | 55.6342 | -3.4674e-10 | 55.7332
-9.5497e-12 | 56.0392 | -3.638e-10 | 56.1372
-1.2733e-11 | 56.4442 | -3.5288e-10 | 56.5422
-1.3188e-11 | 56.8503 | -3.9904e-10 | 56.9463
-9.3223e-12 | 57.2553 | -3.5357e-10 | 57.3523
-1.7053e-11 | 57.6593 | -3.8494e-10 | 57.7563
-1.3188e-11 | 58.0653 | -3.4788e-10 | 58.1603
-7.5033e-12 | 58.4723 | -3.7107e-10 | 58.5654
-1.0004e-11 | 58.8804 | -4.0291e-10 | 58.9704
-1.0459e-11 | 59.2904 | -3.4538e-10 | 59.3784
-9.5497e-12 | 59.6984 | -3.7608e-10 | 59.8814
-3.8654e-12 | 60.1104 | -3.181e-10 | 60.2874
-1.2733e-11 | 60.5165 | -3.5902e-10 | 60.6925
-8.1855e-12 | 60.9225 | -3.7721e-10 | 61.0975
-8.8676e-12 | 61.3275 | -3.8608e-10 | 61.9505
-6.5938e-12 | 61.7345 | -3.9677e-10 | 62.4066
-1.296e-11 | 62.1416 | -4.0291e-10 | 62.8126
-6.1391e-12 | 62.5486 | -3.2924e-10 | 63.2186
-7.7307e-12 | 62.9536 | -3.7835e-10 | 63.6246
-1.0004e-11 | 63.3736 | -3.356e-10 | 64.0307
-1.2506e-11 | 63.7796 | -3.8813e-10 | 64.4367
-9.3223e-12 | 64.1877 | -3.6175e-10 | 64.8417
-8.6402e-12 | 64.5927 | -3.6061e-10 | 65.2467
-1.2051e-11 | 64.9987 | -3.4993e-10 | 65.6528
-9.0949e-12 | 65.4047 | -3.6425e-10 | 66.0578
-7.5033e-12 | 65.8098 | -3.6061e-10 | 66.4658
-8.1855e-12 | 66.2218 | -4.0563e-10 | 66.8708
-6.8212e-12 | 66.6268 | -3.8199e-10 | 67.2759
-7.5033e-12 | 67.0318 | -3.5925e-10 | 67.6799
-1.2733e-11 | 67.4379 | -3.8835e-10 | 68.0839
-7.7307e-12 | 67.8449 | -4.0063e-10 | 68.4889
-7.9581e-12 | 68.2499 | -3.7085e-10 | 68.8959
-5.0022e-12 | 68.6539 | -3.4902e-10 | 69.299
-8.6402e-12 | 69.0589 | -3.613e-10 | 69.705
-1.387e-11 | 69.464 | -4.0654e-10 | 70.109
-1.2733e-11 | 69.871 | -3.4743e-10 | 70.514
-1.1141e-11 | 70.276 | -3.1832e-10 | 70.9261
-9.3223e-12 | 70.682 | -4.1473e-10 | 71.3321
-8.6402e-12 | 71.0861 | -3.3197e-10 | 71.7371
-9.3223e-12 | 71.4911 | -3.3333e-10 | 72.1431
-1.1596e-11 | 71.8971 | -3.6084e-10 | 72.5482
-1.0232e-11 | 72.3031 | -3.497e-10 | 72.9542
-8.8676e-12 | 72.7072 | -3.8813e-10 | 73.3582
-1.0459e-11 | 73.1132 | -3.3742e-10 | 73.7622
-7.0486e-12 | 73.5182 | -3.6835e-10 | 74.1672
-8.8676e-12 | 73.9232 | -3.1923e-10 | 74.5723
-9.0949e-12 | 74.3282 | -3.5948e-10 | 74.9783
-8.8676e-12 | 74.7323 | -3.7107e-10 | 75.3833
-9.0949e-12 | 75.1513 | -3.9108e-10 | 75.7893
-1.1141e-11 | 75.5563 | -4.1427e-10 | 76.1944
-1.1141e-11 | 75.9633 | -3.2537e-10 | 76.5994
-1.0232e-11 | 76.3684 | -3.381e-10 | 77.0064
-7.9581e-12 | 76.7724 | -3.6789e-10 | 77.4124
-1.5916e-11 | 77.1764 | -3.2037e-10 | 77.8165
-1.0687e-11 | 77.5814 | -3.5993e-10 | 78.2215
-9.0949e-12 | 77.9935 | -4.2019e-10 | 78.6255
-1.0687e-11 | 78.4035 | -3.3469e-10 | 79.0315
-7.9581e-12 | 78.8105 | -4.1427e-10 | 79.4365
-1.0004e-11 | 79.2165 | -4.0427e-10 | 79.8416
-1.4097e-11 | 79.6215 | -3.588e-10 | 80.2456
-1.1369e-11 | 80.0266 | -3.8176e-10 | 80.6496
-1.1823e-11 | 80.4336 | -3.6016e-10 | 81.0556
-7.0486e-12 | 80.8386 | -3.7107e-10 | 81.4597
-1.1596e-11 | 81.2446 | -3.6403e-10 | 81.8647
-1.1823e-11 | 81.6487 | -3.5766e-10 | 82.2687
-9.3223e-12 | 82.0547 | -3.6789e-10 | 82.6737
-9.3223e-12 | 82.4687 | -3.5743e-10 | 83.0897
-9.0949e-12 | 82.8747 | -3.4674e-10 | 83.4958
-1.1369e-11 | 83.2798 | -2.9854e-10 | 83.9008
-1.2051e-11 | 83.6858 | -3.6925e-10 | 84.3058
-1.1369e-11 | 84.0908 | -3.7221e-10 | 84.7108
-7.0486e-12 | 84.4958 | -3.356e-10 | 85.1199
|  | -3.8449e-10 | 85.5249
|  | -3.8244e-10 | 85.9299
|  | -3.929e-10 | 86.3439
|  | -3.5607e-10 | 86.75
|  | -3.1605e-10 | 87.155
|  | -3.5061e-10 | 87.559
|  | -3.979e-10 | 87.964
|  | -4.1018e-10 | 88.3701
|  | -3.6084e-10 | 88.7741

## 690uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.1172e-11 | 3.6573e-12 | -5.1993e-10 | 4.018e-11
GEMs On |  | GEMs On
-9.3223e-12 | 0.31402 | -5.4911e-10 | 0.31202
-1.0687e-11 | 1.0001 | -5.4752e-10 | 0.99606
-1.0004e-11 | 1.4061 | -4.9999e-10 | 1.4021
-1.4552e-11 | 1.8131 | -4.8021e-10 | 1.8061
-6.8212e-12 | 2.2181 | -4.9181e-10 | 2.2111
-1.1369e-11 | 2.6221 | -5.1864e-10 | 2.6172
-7.9581e-12 | 3.0292 | -4.4997e-10 | 3.0222
-7.9581e-12 | 3.4342 | -5.4683e-10 | 3.4272
-9.0949e-12 | 3.8392 | -5.6116e-10 | 3.8322
-1.0004e-11 | 4.2442 | -5.009e-10 | 4.2382
-7.276e-12 | 4.6503 | -5.0227e-10 | 4.8023
-8.1855e-12 | 5.0553 | -5.6434e-10 | 5.2633
-7.5033e-12 | 5.4603 | -5.1068e-10 | 5.6683
-8.6402e-12 | 5.8653 | -5.7867e-10 | 6.1654
-1.0459e-11 | 6.2704 | -4.7748e-10 | 6.5744
-1.5916e-11 | 6.6764 | -5.6593e-10 | 6.9944
-7.0486e-12 | 7.0814 | -5.1227e-10 | 7.3994
-8.6402e-12 | 7.4854 | -5.457e-10 | 7.8044
-1.7508e-11 | 7.8895 | -4.7521e-10 | 8.2095
-1.5689e-11 | 8.2955 | -5.1818e-10 | 8.6145
-9.7771e-12 | 8.7005 | -4.9999e-10 | 9.0205
-1.1823e-11 | 9.1045 | -5.8276e-10 | 9.4265
-9.5497e-12 | 9.5105 | -5.964e-10 | 9.8316
-1.2506e-11 | 9.9146 | -4.7339e-10 | 10.2356
-1.0914e-11 | 10.3206 | -5.0431e-10 | 10.6426
-5.457e-12 | 10.7266 | -4.9909e-10 | 11.0486
-1.2051e-11 | 11.1306 | -5.5388e-10 | 11.4557
-6.8212e-12 | 11.5357 | -5.3251e-10 | 11.8607
-9.7771e-12 | 11.9397 | -5.1023e-10 | 12.2657
-8.4128e-12 | 12.3457 | -5.1614e-10 | 12.6687
-2.9559e-12 | 12.7497 | -4.8931e-10 | 13.0747
-1.1823e-11 | 13.1538 | -5.3501e-10 | 13.4818
-1.0459e-11 | 13.5598 | -5.2364e-10 | 13.8868
-1.7735e-11 | 13.9638 | -5.1682e-10 | 14.2928
-2.0691e-11 | 14.3698 | -5.4138e-10 | 14.6968
-9.0949e-12 | 14.7759 | -5.2773e-10 | 15.1139
-7.0486e-12 | 15.1809 | -5.3637e-10 | 15.5189
-9.5497e-12 | 15.5849 | -5.2478e-10 | 15.9249
-9.7771e-12 | 15.9889 | -4.7635e-10 | 16.3309
-7.0486e-12 | 16.3939 | -5.5388e-10 | 16.75
-8.6402e-12 | 16.799 | -4.6748e-10 | 17.156
-7.9581e-12 | 17.205 | -5.0454e-10 | 17.561
-3.2969e-11 | 17.609 | -5.7821e-10 | 17.966
-1.6371e-11 | 18.015 | -5.1591e-10 | 18.3741
-9.0949e-12 | 18.42 | -4.8885e-10 | 18.7801
-1.9099e-11 | 18.8261 | -5.5093e-10 | 19.1881
-1.0004e-11 | 19.2301 | -5.5934e-10 | 19.5941
-9.0949e-13 | 19.6351 | -4.9226e-10 | 19.9981
-7.276e-12 | 20.0382 | -4.8135e-10 | 20.4042
-5.2296e-12 | 20.4432 | -4.6884e-10 | 20.8092
-1.0004e-11 | 20.8482 | -4.6202e-10 | 21.2162
-1.3188e-11 | 21.2522 | -5.4706e-10 | 21.6222
-1.0004e-11 | 21.6572 | -5.4706e-10 | 22.0273
-1.2051e-11 | 22.0623 | -5.1841e-10 | 22.4363
-9.7771e-12 | 22.4683 | -4.7953e-10 | 22.8423
-8.4128e-12 | 22.8733 | -6.0845e-10 | 23.2473
-1.0004e-11 | 23.2763 | -4.9749e-10 | 23.6523
-1.6144e-11 | 23.6834 | -5.0773e-10 | 24.0574
-1.7053e-11 | 24.0874 | -4.9886e-10 | 24.4674
-9.5497e-12 | 24.4934 | -5.0227e-10 | 24.8914
-9.3223e-12 | 24.8974 | -5.1227e-10 | 25.2975
-8.1855e-12 | 25.3044 | -4.9249e-10 | 25.7035
-7.9581e-12 | 25.7085 | -5.6207e-10 | 26.1085
-8.6402e-12 | 26.1135 | -5.3592e-10 | 26.5135
-6.8212e-12 | 26.5175 | -5.2864e-10 | 26.9295
-1.0004e-11 | 26.9225 | -4.7476e-10 | 27.3336
-2.0009e-11 | 27.3276 | -4.8158e-10 | 27.7426
-1.3642e-11 | 27.7316 | -5.3205e-10 | 28.1566
-1.2278e-11 | 28.1376 | -5.4888e-10 | 28.5616
-1.819e-11 | 28.5436 | -5.5275e-10 | 28.9687
-1.4779e-11 | 28.9477 | -5.6184e-10 | 29.3737
-9.0949e-12 | 29.3527 | -5.5866e-10 | 29.7797
-1.0459e-11 | 29.7587 | -5.4752e-10 | 30.1887
-1.5689e-11 | 30.1637 | -4.1246e-10 | 30.5998
-7.276e-12 | 30.5708 | -5.1728e-10 | 31.0048
-1.3642e-11 | 30.9758 | -5.6207e-10 | 31.4098
-9.7771e-12 | 31.3798 | -5.7116e-10 | 31.8138
-1.3415e-11 | 31.7858 | -4.9363e-10 | 32.2178
-1.3642e-11 | 32.1908 | -5.2069e-10 | 32.6219
-1.1369e-11 | 32.5959 | -5.8253e-10 | 33.0279
-1.7735e-11 | 33.0009 | -4.6293e-10 | 33.4329
-1.4325e-11 | 33.4069 | -4.6998e-10 | 33.8359
-1.4325e-11 | 33.8139 | -5.0932e-10 | 34.24
-9.0949e-12 | 34.22 | -4.9408e-10 | 34.644
-1.0004e-11 | 34.625 | -5.3774e-10 | 35.049
-8.6402e-12 | 35.03 | -5.2364e-10 | 35.458
-9.3223e-12 | 35.435 | -5.3615e-10 | 35.865
-9.0949e-12 | 35.84 | -5.1159e-10 | 36.2721
-1.1369e-11 | 36.2451 | -5.648e-10 | 36.6771
-8.1855e-12 | 36.6501 | -5.5593e-10 | 37.0821
-1.2506e-11 | 37.0541 | -5.432e-10 | 37.4861
-1.1823e-11 | 37.4601 | -5.7662e-10 | 37.8922
-1.1823e-11 | 37.8662 | -5.5275e-10 | 38.2962
-9.5497e-12 | 38.2712 | -3.9836e-10 | 38.7042
-9.0949e-12 | 38.6762 | -5.2819e-10 | 39.1102
-8.6402e-12 | 39.0812 | -5.4388e-10 | 39.5153
-1.0914e-11 | 39.4863 | -5.2751e-10 | 39.9203
-9.5497e-12 | 39.8903 | -4.9749e-10 | 40.3253
-1.2506e-11 | 40.2963 | -6.0345e-10 | 40.7303
-1.387e-11 | 40.7003 | -4.7362e-10 | 41.1454
-9.0949e-12 | 41.1054 | -4.7021e-10 | 41.5504
-8.6402e-12 | 41.5114 | -4.9999e-10 | 41.9554
-1.0687e-11 | 41.9164 | -5.0272e-10 | 42.3754
-9.7771e-12 | 42.3214 | -4.9431e-10 | 42.7794
-8.6402e-12 | 42.7254 | -4.8863e-10 | 43.1835
-7.5033e-12 | 43.1305 | -5.0909e-10 | 43.6125
-2.0236e-11 | 43.5355 | -5.3933e-10 | 44.0165
-1.2506e-11 | 43.9405 | -4.8271e-10 | 44.4225
-1.296e-11 | 44.3465 | -5.0386e-10 | 44.8346
-9.7771e-12 | 44.7516 | -4.7476e-10 | 45.2396
-1.1141e-11 | 45.1566 | -5.2592e-10 | 45.6436
-9.7771e-12 | 45.5626 | -5.0181e-10 | 46.0486
-1.0687e-11 | 45.9676 | -4.602e-10 | 46.4547
-1.296e-11 | 46.3727 | -5.7321e-10 | 46.8607
-9.0949e-12 | 46.7777 | -5.4297e-10 | 47.2657
-9.5497e-12 | 47.1827 | -5.2887e-10 | 47.6717
-9.3223e-12 | 47.5877 | -4.5566e-10 | 48.0787
-1.4779e-11 | 47.9927 | -4.6066e-10 | 48.4838
-8.4128e-12 | 48.3998 | -6.1482e-10 | 48.8888
-1.0914e-11 | 48.8038 | -5.7707e-10 | 49.2948
-1.0459e-11 | 49.2098 | -4.7748e-10 | 49.6998
-1.0914e-11 | 49.6158 | -5.2978e-10 | 50.1049
-1.2051e-11 | 50.0199 | -4.4656e-10 | 50.5099
-1.5007e-11 | 50.4269 | -5.4615e-10 | 50.9139
-1.4552e-11 | 50.8319 | -5.1114e-10 | 51.3189
-1.4325e-11 | 51.2379 | -5.557e-10 | 51.723
-1.1369e-11 | 51.6429 | -4.7976e-10 | 52.128
-9.5497e-12 | 52.048 | -5.582e-10 | 52.532
-7.276e-12 | 52.451 | -5.3888e-10 | 52.937
-1.1141e-11 | 52.856 | -4.7908e-10 | 53.3411
-9.5497e-12 | 53.261 | -5.2114e-10 | 53.7461
-1.4552e-11 | 53.6671 | -4.9408e-10 | 54.1501
-2.2737e-11 | 54.0711 | -4.7726e-10 | 54.5571
-7.0486e-12 | 54.4771 | -5.2569e-10 | 54.9631
-1.2051e-11 | 54.8821 | -5.9026e-10 | 55.3682
-1.1369e-11 | 55.2862 | -5.4752e-10 | 55.7732
-1.0459e-11 | 55.6922 | -5.8799e-10 | 56.1772
-1.1369e-11 | 56.0972 | -5.2046e-10 | 56.5822
-1.2278e-11 | 56.5072 | -4.6452e-10 | 56.9873
-1.2051e-11 | 56.9143 | -5.1796e-10 | 57.3923
-1.0004e-11 | 57.3203 | -5.2478e-10 | 57.7973
-1.3188e-11 | 57.7253 | -4.7976e-10 | 58.2013
-1.3642e-11 | 58.1313 | -4.7021e-10 | 58.6063
-1.0004e-11 | 58.5354 | -4.9022e-10 | 59.0104
-1.1141e-11 | 58.9424 | -4.743e-10 | 59.4154
-5.6843e-12 | 59.3484 | -5.648e-10 | 59.8194
-1.0914e-11 | 59.7524 | -5.5661e-10 | 60.2254
-9.0949e-12 | 60.1574 | -4.6361e-10 | 60.6305
-1.7735e-11 | 60.5615 | -5.7344e-10 | 61.0345
-9.0949e-12 | 60.9675 | -4.8976e-10 | 61.4385
-7.0486e-12 | 61.3715 | -5.4706e-10 | 61.8425
-7.9581e-12 | 61.7755 | -5.7526e-10 | 62.2476
-7.5033e-12 | 62.1806 | -5.0568e-10 | 62.6526
-9.3223e-12 | 62.5856 | -4.9408e-10 | 63.0576
-1.0687e-11 | 62.9896 | -4.9636e-10 | 63.4636
-1.0687e-11 | 63.3946 | -5.1068e-10 | 63.8676
-1.0914e-11 | 63.7996 | -5.6684e-10 | 64.2727
-1.0687e-11 | 64.2047 | -4.2701e-10 | 64.6777
-1.296e-11 | 64.6097 | -4.7885e-10 | 65.0827
-9.0949e-12 | 65.0157 | -5.2387e-10 | 65.4878
-1.2506e-11 | 65.4207 | -4.527e-10 | 65.8928
-1.0459e-11 | 65.8258 | -5.0841e-10 | 66.2978
-1.7735e-11 | 66.2298 | -5.2341e-10 | 66.7038
-1.2278e-11 | 66.6358 | -4.7135e-10 | 67.1088
-9.7771e-12 | 67.0408 | -4.2633e-10 | 67.5129
-9.0949e-12 | 67.4459 | -5.8662e-10 | 67.9169
-1.0459e-11 | 67.8509 | -6.13e-10 | 68.3209
-1.0459e-11 | 68.2579 | -5.2478e-10 | 68.7249
-8.4128e-12 | 68.6629 | -4.3906e-10 | 69.129
-1.3642e-11 | 69.0679 | -5.0977e-10 | 69.534
-1.2278e-11 | 69.473 | -4.8431e-10 | 69.937
-1.4325e-11 | 69.878 | -5.6752e-10 | 70.342
-1.296e-11 | 70.283 | -5.914e-10 | 70.7451
-7.0486e-12 | 70.689 | -5.2182e-10 | 71.1511
-9.5497e-12 | 71.0941 | -5.266e-10 | 71.5551
-1.6371e-11 | 71.5001 | -5.2501e-10 | 71.9601
-2.2055e-11 | 71.9051 | -5.1023e-10 | 72.3651
-1.0687e-11 | 72.3091 | -5.4524e-10 | 72.7692
-1.0459e-11 | 72.7152 | -4.9272e-10 | 73.1742
-1.5007e-11 | 73.1212 | -5.8731e-10 | 73.5782
-1.0687e-11 | 73.5272 | -4.9249e-10 | 73.9832
-1.1141e-11 | 73.9332 | -5.7071e-10 | 74.3872
-1.1369e-11 | 74.3372 | -4.2814e-10 | 74.7923
-1.387e-11 | 74.7433 | -5.7366e-10 | 75.1973
-9.7771e-12 | 75.1503 | -5.2341e-10 | 75.6013
-2.0236e-11 | 75.5553 | -5.2887e-10 | 76.0063
-1.0687e-11 | 75.9603 | -4.8976e-10 | 76.4124
-8.4128e-12 | 76.3654 | -5.4138e-10 | 76.8174
-1.0687e-11 | 76.7714 | -5.3683e-10 | 77.2224
-9.3223e-12 | 77.1764 | -5.1637e-10 | 77.6294
-1.0004e-11 | 77.5804 | -4.8772e-10 | 78.0345
-1.1369e-11 | 77.9855 | -5.4456e-10 | 78.4395
-8.1855e-12 | 78.3905 | -4.9317e-10 | 78.8435
-1.7508e-11 | 78.7955 | -5.4956e-10 | 79.2485
-8.4128e-12 | 79.1995 | -5.4956e-10 | 79.6596
-1.296e-11 | 79.6046 | -5.2978e-10 | 80.0646
-1.0687e-11 | 80.0106 | -5.1159e-10 | 80.4686
-8.1855e-12 | 80.4156 | -5.5775e-10 | 80.8736
-1.0687e-11 | 80.8206 | -5.316e-10 | 81.2776
-1.2506e-11 | 81.2267 | -4.9363e-10 | 81.6827
-8.4128e-12 | 81.6307 | -5.5343e-10 | 82.0867
-8.6402e-12 | 82.0477 | -4.8135e-10 | 82.5087
-1.3642e-11 | 82.4527 | -5.0295e-10 | 82.9147
-8.1855e-12 | 82.8577 | -4.6839e-10 | 83.3198
-1.5234e-11 | 83.2618 | -4.411e-10 | 83.7248
-1.296e-11 | 83.6678 | -5.2046e-10 | 84.1308
|  | -5.4774e-10 | 84.5358
|  | -4.9386e-10 | 84.9419
|  | -5.4501e-10 | 85.3469
|  | -5.8162e-10 | 85.7519
|  | -5.3683e-10 | 86.1579
|  | -5.5093e-10 | 86.5619
|  | -5.5843e-10 | 86.967
|  | -5.3365e-10 | 87.371
|  | -6.1073e-10 | 87.776
|  | -5.2955e-10 | 88.182
|  | -4.7953e-10 | 88.5881
|  | -5.4501e-10 | 88.9931
|  | -5.4501e-10 | 89.3971
|  | -4.8612e-10 | 89.8021
|  | -5.4251e-10 | 90.2072
|  | -4.9545e-10 | 90.6122
|  | -5.2455e-10 | 91.0162
|  | -5.457e-10 | 91.4212
|  | -5.6684e-10 | 91.8252
|  | -5.4615e-10 | 92.2303
|  | -5.6616e-10 | 92.6353
|  | -5.4683e-10 | 93.0403
|  | -5.2091e-10 | 93.4453
|  | -5.2887e-10 | 93.8504
|  | -4.3246e-10 | 94.2544

## 700uA
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
-1.8069e-11 | 4.9387e-12 | -7.4341e-10 | 6.0247e-11
GEMs On |  | GEMs On
-1.3642e-11 | 1.5051 | -7.5102e-10 | 0.31602
-1.2051e-11 | 2.2781 | -8.3674e-10 | 1.0001
-1.387e-11 | 2.6852 | -8.1013e-10 | 1.4071
-9.7771e-12 | 3.0902 | -7.4328e-10 | 1.8121
-1.9327e-11 | 3.4942 | -8.0649e-10 | 2.2181
-1.1141e-11 | 3.9002 | -7.2873e-10 | 2.6212
-1.1141e-11 | 4.3052 | -7.3851e-10 | 3.0262
-8.6402e-12 | 4.7123 | -8.4719e-10 | 3.4302
-1.0687e-11 | 5.1173 | -6.4642e-10 | 3.8352
-8.6402e-12 | 5.5233 | -7.3828e-10 | 4.2402
-1.2051e-11 | 5.9293 | -7.1509e-10 | 4.6443
-8.4128e-12 | 6.3344 | -7.2441e-10 | 5.0493
-1.1596e-11 | 6.7394 | -7.592e-10 | 5.4553
-1.0687e-11 | 7.1454 | -8.3241e-10 | 5.8593
-1.4097e-11 | 7.5514 | -8.5902e-10 | 6.2634
-9.7771e-12 | 7.9575 | -7.8626e-10 | 6.6674
-1.1141e-11 | 8.3625 | -6.4529e-10 | 7.0714
-1.3642e-11 | 8.7695 | -7.6898e-10 | 7.4764
-1.3415e-11 | 9.1745 | -6.9872e-10 | 7.8825
-1.4779e-11 | 9.5805 | -6.6757e-10 | 8.2885
-1.2733e-11 | 9.9866 | -7.3578e-10 | 8.6925
-9.5497e-12 | 10.3926 | -7.6443e-10 | 9.0985
-8.6402e-12 | 10.7976 | -8.5788e-10 | 9.5035
-1.5461e-11 | 11.2036 | -7.6534e-10 | 9.9086
-2.4329e-11 | 11.6077 | -6.7666e-10 | 10.3126
-1.0687e-11 | 12.0147 | -7.6557e-10 | 10.7186
-1.0687e-11 | 12.4197 | -6.8212e-10 | 11.1236
-1.2051e-11 | 12.8257 | -7.5534e-10 | 11.5277
-1.1596e-11 | 13.2398 | -6.6711e-10 | 11.9327
-1.5007e-11 | 13.6448 | -7.2737e-10 | 12.3377
-1.5234e-11 | 14.0508 | -6.6325e-10 | 12.7427
-1.4552e-11 | 14.4558 | -8.1627e-10 | 13.1478
-1.9327e-11 | 14.8608 | -7.6761e-10 | 13.5528
-1.6144e-11 | 15.2669 | -7.4851e-10 | 13.9578
-1.0232e-11 | 15.6729 | -8.1286e-10 | 14.3618
-1.387e-11 | 16.0789 | -7.6966e-10 | 14.7658
-1.0914e-11 | 16.4859 | -8.0627e-10 | 15.1719
-1.4552e-11 | 16.891 | -8.0922e-10 | 15.5779
-2.1146e-11 | 17.299 | -7.3419e-10 | 15.9849
-1.5916e-11 | 17.705 | -7.6011e-10 | 16.3889
-2.3419e-11 | 18.11 | -7.5215e-10 | 16.794
-1.3642e-11 | 18.5161 | -7.2919e-10 | 17.199
-1.6144e-11 | 18.9211 | -6.4369e-10 | 17.604
-9.7771e-12 | 19.3271 | -7.9353e-10 | 18.008
-1.387e-11 | 19.7331 | -6.5484e-10 | 18.413
-1.3188e-11 | 20.1381 | -6.9099e-10 | 18.8181
-1.4097e-11 | 20.5452 | -6.8985e-10 | 19.2231
-1.5234e-11 | 20.9822 | -7.069e-10 | 19.6271
-1.6144e-11 | 21.4152 | -8.3855e-10 | 20.0322
-1.5007e-11 | 21.8212 | -7.1918e-10 | 20.4362
-1.4779e-11 | 22.2433 | -6.8417e-10 | 20.8402
-1.7963e-11 | 22.6593 | -7.6307e-10 | 21.2452
-1.3642e-11 | 23.0653 | -7.06e-10 | 21.6512
-2.16e-11 | 23.4723 | -7.8353e-10 | 22.0553
-1.387e-11 | 23.8784 | -7.9217e-10 | 22.4613
-1.728e-11 | 24.2834 | -7.8671e-10 | 22.8663
-1.5007e-11 | 24.6984 | -6.6416e-10 | 23.2713
-1.728e-11 | 25.1044 | -6.9213e-10 | 23.6774
-1.3642e-11 | 25.5125 | -7.4988e-10 | 24.0814
-1.8872e-11 | 25.9175 | -7.0668e-10 | 24.4874
-1.4552e-11 | 26.3225 | -8.3173e-10 | 24.8914
-1.5916e-11 | 26.7295 | -8.1286e-10 | 25.2965
-1.4097e-11 | 27.1346 | -6.8962e-10 | 25.7015
-1.5234e-11 | 27.5406 | -7.8899e-10 | 26.1075
-1.2733e-11 | 27.9486 | -8.5652e-10 | 26.5125
-1.387e-11 | 28.3546 | -6.7143e-10 | 26.9175
-1.728e-11 | 28.7596 | -6.1914e-10 | 27.3226
-1.4779e-11 | 29.1687 | -7.8376e-10 | 27.7276
-2.2283e-11 | 29.5737 | -7.1122e-10 | 28.1326
-1.7735e-11 | 29.9797 | -8.356e-10 | 28.5366
-1.2506e-11 | 30.3867 | -6.7166e-10 | 28.9417
-1.5007e-11 | 30.7918 | -7.9194e-10 | 29.3467
-3.1605e-11 | 31.1958 | -7.4874e-10 | 29.7517
-1.387e-11 | 31.6018 | -7.3237e-10 | 30.1567
-1.4097e-11 | 32.0068 | -7.8217e-10 | 30.5607
-1.4097e-11 | 32.4119 | -6.5916e-10 | 30.9648
-1.4552e-11 | 32.8189 | -6.8326e-10 | 31.3678
-1.6371e-11 | 33.2249 | -6.8599e-10 | 31.7728
-1.8417e-11 | 33.6309 | -8.074e-10 | 32.1778
-1.387e-11 | 34.0429 | -8.8835e-10 | 32.5829
-1.5461e-11 | 34.448 | -6.778e-10 | 32.9879
-2.0464e-11 | 34.854 | -7.2305e-10 | 33.3929
-2.1373e-11 | 35.262 | -7.6147e-10 | 33.7969
-1.387e-11 | 35.667 | -7.5011e-10 | 34.202
-1.5461e-11 | 36.0731 | -8.5038e-10 | 34.606
-1.2051e-11 | 36.4801 | -7.5147e-10 | 35.011
-1.9327e-11 | 36.8861 | -8.1923e-10 | 35.416
-1.8417e-11 | 37.2921 | -7.4738e-10 | 35.821
-1.6144e-11 | 37.6982 | -6.271e-10 | 36.2261
-1.4552e-11 | 38.1032 | -7.592e-10 | 36.6301
-1.4097e-11 | 38.5082 | -6.08e-10 | 37.0371
-1.4779e-11 | 38.9142 | -7.617e-10 | 37.4421
-1.728e-11 | 39.3192 | -8.1059e-10 | 37.8472
-1.5007e-11 | 39.7243 | -7.0759e-10 | 38.2512
-1.5916e-11 | 40.1303 | -8.1764e-10 | 38.6572
-1.7735e-11 | 40.5433 | -8.6243e-10 | 39.0632
-1.4552e-11 | 40.9563 | -6.9622e-10 | 39.4683
-1.387e-11 | 41.3614 | -7.8489e-10 | 39.8723
-1.9327e-11 | 41.7684 | -8.1036e-10 | 40.2783
-1.8645e-11 | 42.1734 | -6.9781e-10 | 40.6833
-2.1146e-11 | 42.5984 | -6.7644e-10 | 41.0883
-2.9331e-11 | 43.1525 | -8.3719e-10 | 41.4924
-2.2965e-11 | 43.6245 | -7.5374e-10 | 41.8974
-1.8872e-11 | 44.0285 | -7.1577e-10 | 42.3024
-1.5461e-11 | 44.4335 | -7.2214e-10 | 42.7064
-1.7053e-11 | 44.8556 | -7.2941e-10 | 43.1105
-1.9327e-11 | 45.2656 | -7.2532e-10 | 43.5155
-1.6371e-11 | 45.6726 | -7.517e-10 | 43.9195
-1.387e-11 | 46.0806 | -7.7739e-10 | 44.3245
-1.5916e-11 | 46.4857 | -6.869e-10 | 44.7296
-1.5234e-11 | 46.8927 | -7.3805e-10 | 45.1346
-1.4552e-11 | 47.2987 | -6.8212e-10 | 45.5386
-1.5007e-11 | 47.7037 | -7.1418e-10 | 45.9436
-1.4552e-11 | 48.1097 | -6.4347e-10 | 46.3486
-1.728e-11 | 48.5208 | -6.9213e-10 | 46.7517
-1.9327e-11 | 48.9268 | -6.6234e-10 | 47.1577
-1.7053e-11 | 49.3318 | -8.265e-10 | 47.5627
-2.16e-11 | 49.7368 | -7.733e-10 | 47.9687
-1.8417e-11 | 50.1429 | -6.703e-10 | 48.3728
-2.0236e-11 | 50.5479 | -8.3196e-10 | 48.7778
-1.5461e-11 | 50.9529 | -7.8808e-10 | 49.1828
-1.8645e-11 | 51.3589 | -8.2673e-10 | 49.5888
-1.8417e-11 | 51.765 | -6.5779e-10 | 49.9949
-1.728e-11 | 52.17 | -6.6939e-10 | 50.3989
-1.387e-11 | 52.575 | -8.1195e-10 | 50.8049
-2.16e-11 | 52.981 | -8.2423e-10 | 51.2099
-1.8645e-11 | 53.3871 | -6.6029e-10 | 51.615
-1.4552e-11 | 53.7931 | -7.1327e-10 | 52.02
-1.8417e-11 | 54.1991 | -7.6488e-10 | 52.425
-2.2283e-11 | 54.6061 | -8.0036e-10 | 52.831
-1.6144e-11 | 55.0112 | -7.3578e-10 | 53.236
-2.6603e-11 | 55.4162 | -7.0941e-10 | 53.6401
-1.9782e-11 | 55.8232 | -7.6284e-10 | 54.0441
-1.7053e-11 | 56.2302 | -7.7216e-10 | 54.4481
-2.0236e-11 | 56.6342 | -8.074e-10 | 54.8521
-2.0236e-11 | 57.0403 | -7.858e-10 | 55.2592
-1.7963e-11 | 57.4463 | -5.9322e-10 | 55.6642
-2.2737e-11 | 57.8523 | -7.4988e-10 | 56.0682
-2.0464e-11 | 58.2583 | -7.2987e-10 | 56.4732
-1.8645e-11 | 58.6644 | -8.0399e-10 | 56.8772
-1.7963e-11 | 59.0704 | -8.3764e-10 | 57.2803
-2.0009e-11 | 59.4784 | -7.3624e-10 | 57.6843
-2.3647e-11 | 59.8834 | -7.5806e-10 | 58.0903
-2.16e-11 | 60.2895 | -7.3965e-10 | 58.4954
-1.7963e-11 | 60.7015 | -7.4442e-10 | 58.8994
-1.7735e-11 | 61.1075 | -8.4242e-10 | 59.3054
-3.115e-11 | 61.5145 | -7.0304e-10 | 59.7104
-1.8872e-11 | 61.9225 | -7.1418e-10 | 60.1164
-1.9099e-11 | 62.3316 | -8.4265e-10 | 60.5205
-2.2965e-11 | 62.7386 | -8.0081e-10 | 60.9245
-1.7508e-11 | 63.1446 | -8.4378e-10 | 61.3295
-2.2283e-11 | 63.5836 | -7.4829e-10 | 61.7345
-1.6144e-11 | 64.0047 | -7.6329e-10 | 62.1395
-1.8645e-11 | 64.4097 | -6.396e-10 | 62.5456
-1.9099e-11 | 64.8467 | -6.8826e-10 | 62.9506
-2.0236e-11 | 65.2727 | -7.7671e-10 | 63.3556
-2.16e-11 | 65.6778 | -7.8967e-10 | 63.7597
-1.8872e-11 | 66.0898 | -8.0149e-10 | 64.1657
-2.16e-11 | 66.5088 | -6.8349e-10 | 64.5697
-2.0236e-11 | 66.9138 | -7.1555e-10 | 64.9747
-1.6598e-11 | 67.3208 | -6.9531e-10 | 65.3787
-2.16e-11 | 67.7329 | -6.9963e-10 | 65.7838
-1.8872e-11 | 68.1399 | -7.376e-10 | 66.1888
-1.9327e-11 | 68.5449 | -7.3942e-10 | 66.5928
-1.8872e-11 | 68.9509 | -6.653e-10 | 66.9978
-1.9099e-11 | 69.357 | -7.3169e-10 | 67.4029
-2.16e-11 | 69.762 | -7.235e-10 | 67.8079
-2.3419e-11 | 70.167 | -7.1009e-10 | 68.2129
-1.8417e-11 | 70.578 | -8.1309e-10 | 68.6179
-1.9782e-11 | 71.1581 | -7.2964e-10 | 69.0239
-2.16e-11 | 71.5661 | -7.3464e-10 | 69.428
-2.7285e-11 | 71.9731 | -7.0304e-10 | 69.832
-1.9327e-11 | 72.3831 | -7.2441e-10 | 70.238
-2.2737e-11 | 72.9302 | -6.7053e-10 | 70.643
-2.251e-11 | 73.3532 | -6.6029e-10 | 71.0471
-2.1146e-11 | 73.7592 | -7.6147e-10 | 71.4541
-2.6375e-11 | 74.1642 | -8.5015e-10 | 71.8581
-2.0009e-11 | 74.5693 | -7.8239e-10 | 72.2641
-2.2283e-11 | 74.9743 | -7.1645e-10 | 72.6702
-2.4784e-11 | 75.3803 | -7.758e-10 | 73.0752
-2.9331e-11 | 75.7863 | -7.9831e-10 | 73.4802
-2.3874e-11 | 76.1914 | -7.8512e-10 | 73.8852
-1.8872e-11 | 76.5984 | -6.8553e-10 | 74.2892
-2.5466e-11 | 77.0054 | -6.0277e-10 | 74.6943
-1.9099e-11 | 77.4104 | -6.9349e-10 | 75.1003
-1.5461e-11 | 77.8154 | -6.9645e-10 | 75.5053
-2.3647e-11 | 78.2215 | -8.547e-10 | 75.9093
-1.9782e-11 | 78.6275 | -7.0918e-10 | 76.3134
-3.2742e-11 | 79.0335 | -7.3283e-10 | 76.7174
-2.0236e-11 | 79.4385 | -7.6329e-10 | 77.1214
-2.2283e-11 | 79.8426 | -6.4392e-10 | 77.5274
-2.1373e-11 | 80.2486 | -6.4506e-10 | 77.9315
-2.0918e-11 | 80.6536 | -6.5938e-10 | 78.3365
-2.16e-11 | 81.0616 | -7.2146e-10 | 78.7425
-2.1146e-11 | 81.4667 | -7.8899e-10 | 79.1485
-1.9782e-11 | 81.8717 | -7.0304e-10 | 79.5525
-2.1146e-11 | 82.2777 | -7.6261e-10 | 79.9566
-2.6375e-11 | 82.6817 | -7.899e-10 | 80.3626
-2.3874e-11 | 83.0858 | -6.5552e-10 | 80.7676
-2.8649e-11 | 83.4908 | -8.0263e-10 | 81.1726
-3.7062e-11 | 83.8958 | -8.049e-10 | 81.5777
-2.3647e-11 | 84.3018 | -8.0149e-10 | 81.9937
-2.3874e-11 | 84.7088 | -7.6375e-10 | 82.3997
-2.2965e-11 | 85.1619 | -7.3965e-10 | 82.8037
-2.1828e-11 | 85.5679 | -6.9144e-10 | 83.2088
-3.0695e-11 | 85.9729 | -7.0213e-10 | 83.6138
-2.251e-11 | 86.3779 | -6.9781e-10 | 84.0188
-2.6148e-11 | 86.783 | -7.7193e-10 | 84.4238
-2.1146e-11 | 87.187 | -7.4328e-10 | 84.8278
-2.1146e-11 | 87.592 | -7.1827e-10 | 85.2329
-3.0695e-11 | 87.998 | -7.4715e-10 | 85.6379
-2.3647e-11 | 88.4041
-2.3647e-11 | 88.8111
-1.9782e-11 | 89.2161

## 710uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 720uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 730uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 740uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 750uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 760uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 770uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 780uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 790uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 800uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On

## 810uA (hidden)
Source Off |  | Source On
Average | STDEV | Average | STDEV
(uA) | (uA) | (uA) | (uA)
#DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
GEMs On |  | GEMs On